## Modèle budget
Incertitude
C = P * m / Vmère
Concentration mère | mg/L
Source | Estimation | ±a | loi | k | u | coefficient de sensibilité expression | valeur | contribution
Masse Q lecture | 0 | 5e-05 | type B | 1.7321 | 2.8868e-05 | P/V | 0.1 | 8.3333e-12
Masse Q balance (mg) | 25 |  | type B | 2 | 0 | P/V | 0.1 | 0
P (%) Pureté Quinine | 0.99 | 0.01 | type B | 1.7321 | 0.0057735 | m/V | 250 | 2.0833
V mère Fiole jaugée (L) | 0.1 | 0.0001 | type B | 2.4495 | 4.0825e-05 | (-P*m / V^2) | -0.25 | 1.0417e-10
Concentration mère (mg/L) | 247.5 | [5 empty] | u de Cm (mg/L) | 2.0833
[7 empty] | pourcentage incertitude | 0.0084175
Cfille = Cmère * Vmère / Vfille
Concentration fille | mg/L
Source | Estimation | ±a | loi | k | u | coefficient de sensibilité expression | valeur | contribution
Concentration mère | 247.5 | X | type B | X | 2.0833 | Vm/Vf | 0.05 | 0.010851
Vfille fiole jaugée fille | 0.02 | 4e-05 | type B | 2.4495 | 1.633e-05 | Cm/Vf | 12375 | 0.040838
Vmère pipette graduée 1ml | 0.001 | 1e-05 | type B | 1.7321 | 5.7735e-06 | (Cm*Vm)/Vf^2 | 618.75 | 1.2762e-05
Concentratio fille (mg/L) | 12.375 | [5 empty] | u de Cf (mg/L) | 0.051701
0.0041779
Coefficient de dilution | N.A
Source | Estimation | ±a | loi | k | u | coefficient de sensibilité expression | valeur | contribution
V Fiole jaugée | 0.02 | 4e-05 | type B | 2.4495 | 1.633e-05 | 1/Vfiole | 50 | 6.6667e-07
V Pipette graduée | 0.01 | 0.0001 | type B | 1.7321 | 5.7735e-05 | (-Vpipette/Vfiole) | -0.5 | 8.3333e-10
Coefficient de dilution | 0.5 | [5 empty] | u du coefficient de dilution | 6.675e-07
Intensité (u.a)
Source | Estimation | ±a | loi | k | u | coefficient de sensibilité expression | valeur | contribution
Répétabilité (20 mesures) | 0 | x | normale | x | 2.35 | 1 | 1 | 5.5225
Spectrofluomètre lecture | 0 | 0.005 | type B | 1.7321 | 0.0028868 | 1 | 1 | 8.3333e-06
Incertitude de l'appareil | 4816.352 |  | type b | 2 |  | 1 | 1 | 0
Intensité solution inconnue | 4816.352 | [5 empty] | U de I | 5.5225
[7 empty] | % | 0.0011466

## Concentrations filles
Concentration fille | mg/L
Fille 0 - 0 mg/L | Estimation | ±a | loi | k | u | coefficient de sensibilité expression | valeur | contribution
Acide Sulfurique | 1 mol | N/A | N/A | N/A | N/A | N/A | N/A | N/A
[7 empty] | U_tot | 0,0 mg/L
Solution à doser - X mg/L | Estimation | ±a | loi | k | u | coefficient de sensibilité expression | valeur | contribution
Vpipette graduée | 0.001 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Acide Sulfurique | 1 mol | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Vsolution inconnue | [6 empty] | U_tot | 0,0 mg/L
Fille 1 - 2,05 mg/L | Estimation | ±a | loi | k | u | coefficient de sensibilité expression | valeur | contribution
Concentration mère | 41 | 0.8 | type B | 1.7321 | 0.46188 | Vm/Vf | 0.05 | 0.00053333
Vfille fiole jaugée fille | 0.02 | 4e-05 | type B | 2.4495 | 1.633e-05 | Cm/Vf | 2050 | 0.0011207
Vmère pipette graduée | 0.001 | 1e-05 | type B | 1.7321 | 5.7735e-06 | (Cm*Vm)/Vf^2 | 102.5 | 3.5021e-07
Concentratio fille (mg/L) | 2.05 | [6 empty] | 0.0016544
[7 empty] | U_tot | 0,0017 mg/L
Fille 2 - 4,1 mg/L | Estimation | ±a | loi | k | u | coefficient de sensibilité expression | valeur | contribution
Concentration mère | 41 | 0.8 | type B | 1.7321 | 0.46188 | Vm/Vf | 0.1 | 0.0021333
Vfille fiole jaugée fille | 0.02 | 4e-05 | type B | 2.4495 | 1.633e-05 | Cm/Vf | 2050 | 0.0011207
Vmère pipette graduée | 0.002 | 1e-05 | type B | 1.7321 | 5.7735e-06 | (Cm*Vm)/Vf^2 | 205 | 1.4008e-06
Concentratio fille (mg/L) | 4.1 | [6 empty] | 0.0032554
[7 empty] | U_tot | 0,0032 mg/L
Fille 3 - 6,15 mg/L | Estimation | ±a | loi | k | u | coefficient de sensibilité expression | valeur | contribution
Concentration mère | 41 | 0.8 | type B | 1.7321 | 0.46188 | Vm/Vf | 0.15 | 0.0048
Vfille fiole jaugée fille | 0.02 | 4e-05 | type B | 2.4495 | 1.633e-05 | Cm/Vf | 2050 | 0.0011207
Vmère pipette graduée | 0.003 | 1e-05 | type B | 1.7321 | 5.7735e-06 | (Cm*Vm)/Vf^2 | 307.5 | 3.1519e-06
Concentratio fille (mg/L) | 6.15 | [6 empty] | 0.0059238
[7 empty] | U_tot | 0,0060 mg/L
Fille 4 - 8,2 mg/L | Estimation | ±a | loi | k | u | coefficient de sensibilité expression | valeur | contribution
Concentration mère | 41 | 0.8 | type B | 1.7321 | 0.46188 | Vm/Vf | 0.2 | 0.0085333
Vfille fiole jaugée fille | 0.02 | 4e-05 | type B | 2.4495 | 1.633e-05 | Cm/Vf | 2050 | 0.0011207
Vmère pipette graduée | 0.004 | 1e-05 | type B | 1.7321 | 5.7735e-06 | (Cm*Vm)/Vf^2 | 410 | 5.6033e-06
Concentratio fille (mg/L) | 8.2 | [6 empty] | 0.0096596
[7 empty] | U_tot | 0,010 mg/L
Fille 5 - 10,25 mg/L | Estimation | ±a | loi | k | u | coefficient de sensibilité expression | valeur | contribution
Concentration mère | 41 | 0.8 | type B | 1.7321 | 0.46188 | Vm/Vf | 0.25 | 0.013333
Vfille fiole jaugée fille | 0.02 | 4e-05 | type B | 2.4495 | 1.633e-05 | Cm/Vf | 2050 | 0.0011207
Vmère pipette graduée | 0.005 | 1e-05 | type B | 1.7321 | 5.7735e-06 | (Cm*Vm)/Vf^2 | 512.5 | 8.7552e-06
Concentratio fille (mg/L) | 10.25 | [6 empty] | 0.014463
[7 empty] | U_tot | 0,014 mg/L

## Intensité
Intensité | N/A
expression "a" répétabilité | σx/√n
Intensité (u.a) - blanc 0,00 mg/L
Source | Estimation | ±a | loi | k | u | coefficient de sensibilité expression | valeur | contribution
Répétabilité (20 mesures) | 14.26 | 0.053 | normale | 2 | 0.0265 | 1 | 1 | 0.00070225
Spectrofluomètre lecture | 0 | 0.005 | type B | 1.7321 | 0.0028868 | 1 | 1 | 8.3333e-06
Incertitude de l'appareil | N/A | - | type B | 1.7321 | - | 1 | 1 | -
[7 empty] | U_tot | 0.00071058
Intensité (u.a) - fille 2,05 mg/L
Source | Estimation | ±a | loi | k | u | coefficient de sensibilité expression | valeur | contribution
Répétabilité (20 mesures) | 1746.61 | 0.91 | normale | 2 | 0.455 | 1 | 1 | 0.20703
Spectrofluomètre lecture | 0 | 0.005 | type B | 1.7321 | 0.0028868 | 1 | 1 | 8.3333e-06
Incertitude de l'appareil | N/A | - | type B | 1.7321 | - | 1 | 1 | -
[7 empty] | U_tot | 0.20703
Intensité (u.a) - fille 4,1 mg/L
Source | Estimation | ±a | loi | k | u | coefficient de sensibilité expression | valeur | contribution
Répétabilité (20 mesures) | 3345.7 | 1.14 | normale | 2 | 0.57 | 1 | 1 | 0.3249
Spectrofluomètre lecture | 0 | 0.005 | type B | 1.7321 | 0.0028868 | 1 | 1 | 8.3333e-06
Incertitude de l'appareil | N/A | - | type B | 1.7321 | - | 1 | 1 | -
[7 empty] | U_tot | 0.32491
Intensité (u.a) - fille 6,15 mg/L
Source | Estimation | ±a | loi | k | u | coefficient de sensibilité expression | valeur | contribution
Répétabilité (20 mesures) | 4972.89 | 1.77 | normale | 2 | 0.885 | 1 | 1 | 0.78323
Spectrofluomètre lecture | 0 | 0.005 | type B | 1.7321 | 0.0028868 | 1 | 1 | 8.3333e-06
Incertitude de l'appareil | N/A | - | type B | 1.7321 | - | 1 | 1 | -
[7 empty] | U_tot | 0.78323
Intensité (u.a) - fille 8,2 mg/L
Source | Estimation | ±a | loi | k | u | coefficient de sensibilité expression | valeur | contribution
Répétabilité (20 mesures) | 6376.29 | 1.99 | normale | 2 | 0.995 | 1 | 1 | 0.99003
Spectrofluomètre lecture | 0 | 0.005 | type B | 1.7321 | 0.0028868 | 1 | 1 | 8.3333e-06
Incertitude de l'appareil | N/A | - | type B | 1.7321 | - | 1 | 1 | -
[7 empty] | U_tot | 0.99003
Intensité (u.a) - fille 10,25 mg/L
Source | Estimation | ±a | loi | k | u | coefficient de sensibilité expression | valeur | contribution
Répétabilité (20 mesures) | 7791.72 | 2.43 | normale | 2 | 1.215 | 1 | 1 | 1.4762
Spectrofluomètre lecture | 0 | 0.005 | type B | 1.7321 | 0.0028868 | 1 | 1 | 8.3333e-06
Incertitude de l'appareil | N/A | - | type B | 1.7321 | - | 1 | 1 | -
[7 empty] | U_tot | 1.4762
Intensité (u.a) - Solution à doser mg/L
Source | Estimation | ±a | loi | k | u | coefficient de sensibilité expression | valeur | contribution
Répétabilité (20 mesures) | 5045.78 | 1.62 | normale | 2 | 0.81 | 1 | 1 | 0.6561
Spectrofluomètre lecture | 0 | 0.005 | type B | 1.7321 | 0.0028868 | 1 | 1 | 8.3333e-06
Incertitude de l'appareil | N/A | - | type B | 1.7321 | - | 1 | 1 | -
[7 empty] | U_tot | 0.65611

## R&R
Operateur | Concentration | Intensite
A | 10.25 | 7831.82
A | 10.25 | 7857.58
A | 10.25 | 7850.19
A | 10.25 | 7850.49
A | 10.25 | 7854.29
A | 10.25 | 7848.89
A | 10.25 | 7855.91
A | 10.25 | 7848.33
A | 10.25 | 7832.71
A | 10.25 | 7836.08
A | 10.25 | 7831.1
A | 10.25 | 7840.61
A | 10.25 | 7826.26
A | 10.25 | 7838.73
A | 10.25 | 7829.17
A | 10.25 | 7846.86
A | 10.25 | 7844.29
A | 10.25 | 7841.97
A | 10.25 | 7853.72
A | 10.25 | 7843.69
A | 8.2 | 6433.2
A | 8.2 | 6436.07
A | 8.2 | 6433.2
A | 8.2 | 6421.94
A | 8.2 | 6430.08
A | 8.2 | 6437.64
A | 8.2 | 6457.02
A | 8.2 | 6435.81
A | 8.2 | 6440.28
A | 8.2 | 6434.94
A | 8.2 | 6420.34
A | 8.2 | 6446.96
A | 8.2 | 6419.71
A | 8.2 | 6441.5
A | 8.2 | 6433.79
A | 8.2 | 6432.87
A | 8.2 | 6431.27
A | 8.2 | 6437.21
A | 8.2 | 6434.82
A | 8.2 | 6430.08
A | 6.15 | 5209.59
A | 6.15 | 5215.23
A | 6.15 | 5217.76
A | 6.15 | 5214.75
A | 6.15 | 5221.24
A | 6.15 | 5218.19
A | 6.15 | 5219.29
A | 6.15 | 5210.15
A | 6.15 | 5224.66
A | 6.15 | 5215.31
A | 6.15 | 5208.1
A | 6.15 | 5216.36
A | 6.15 | 5210.33
A | 6.15 | 5215.85
A | 6.15 | 5220.14
A | 6.15 | 5216.85
A | 6.15 | 5215.03
A | 6.15 | 5212.43
A | 6.15 | 5225.7
A | 6.15 | 5207.92
A | 4.1 | 3496.5
A | 4.1 | 3496.79
A | 4.1 | 3490.43
A | 4.1 | 3492.79
A | 4.1 | 3486.62
A | 4.1 | 3495.44
A | 4.1 | 3492.62
A | 4.1 | 3495.99
A | 4.1 | 3498.06
A | 4.1 | 3494.22
A | 4.1 | 3497.52
A | 4.1 | 3496.73
A | 4.1 | 3487.45
A | 4.1 | 3490.67
A | 4.1 | 3496.07
A | 4.1 | 3490.48
A | 4.1 | 3497.15
A | 4.1 | 3490.65
A | 4.1 | 3497.23
A | 4.1 | 3483.94
A | 2.05 | 1770.64
A | 2.05 | 1780.37
A | 2.05 | 1776.38
A | 2.05 | 1774.64
A | 2.05 | 1774.38
A | 2.05 | 1772.74
A | 2.05 | 1775.41
A | 2.05 | 1777.1
A | 2.05 | 1781.29
A | 2.05 | 1779.7
A | 2.05 | 1781.94
A | 2.05 | 1776.97
A | 2.05 | 1776.81
A | 2.05 | 1772.77
A | 2.05 | 1773.05
A | 2.05 | 1774.78
A | 2.05 | 1777.76
A | 2.05 | 1777.1
A | 2.05 | 1779.81
A | 2.05 | 1774.02
A | 0 | 14.3677
A | 0 | 14.3678
A | 0 | 14.3679
A | 0 | 14.368
A | 0 | 14.3681
A | 0 | 14.3682
A | 0 | 14.3683
A | 0 | 14.3684
A | 0 | 14.3685
A | 0 | 14.3686
A | 0 | 14.3687
A | 0 | 14.3688
A | 0 | 14.3689
A | 0 | 14.369
A | 0 | 14.3691
A | 0 | 14.3692
A | 0 | 14.3693
A | 0 | 14.3694
A | 0 | 14.3695
A | 0 | 14.3696
A | NA | 5125.11
A | NA | 5114.67
A | NA | 5121.14
A | NA | 5124.89
A | NA | 5122.12
A | NA | 5122.22
A | NA | 5118.68
A | NA | 5125.03
A | NA | 5120.29
A | NA | 5121.53
A | NA | 5119.7
A | NA | 5124.84
A | NA | 5142.66
A | NA | 5124.56
A | NA | 5117.3
A | NA | 5121.31
A | NA | 5138.41
A | NA | 5113.15
A | NA | 5126.64
A | NA | 5110.68
B | 10.25 | 7788.35
B | 10.25 | 7803.64
B | 10.25 | 7803.38
B | 10.25 | 7803.31
B | 10.25 | 7805.27
B | 10.25 | 7812.55
B | 10.25 | 7790.86
B | 10.25 | 7789.51
B | 10.25 | 7787.71
B | 10.25 | 7795.41
B | 10.25 | 7790.39
B | 10.25 | 7775.85
B | 10.25 | 7793.93
B | 10.25 | 7789.55
B | 10.25 | 7794.61
B | 10.25 | 7777
B | 10.25 | 7766.29
B | 10.25 | 7785.68
B | 10.25 | 7790.67
B | 10.25 | 7790.41
B | 8.2 | 6382.54
B | 8.2 | 6372.9
B | 8.2 | 6389.92
B | 8.2 | 6383.13
B | 8.2 | 6388.14
B | 8.2 | 6380.53
B | 8.2 | 6367.95
B | 8.2 | 6385.51
B | 8.2 | 6384.13
B | 8.2 | 6376.57
B | 8.2 | 6368.8
B | 8.2 | 6360.52
B | 8.2 | 6369.69
B | 8.2 | 6361.06
B | 8.2 | 6373.47
B | 8.2 | 6364.65
B | 8.2 | 6384.27
B | 8.2 | 6374.12
B | 8.2 | 6374.84
B | 8.2 | 6383.08
B | 6.15 | 4983.02
B | 6.15 | 4978.55
B | 6.15 | 4981.41
B | 6.15 | 4971.63
B | 6.15 | 4982.24
B | 6.15 | 4976.47
B | 6.15 | 4964.47
B | 6.15 | 4973.77
B | 6.15 | 4982.45
B | 6.15 | 4977.82
B | 6.15 | 4980.73
B | 6.15 | 4969.28
B | 6.15 | 4972.45
B | 6.15 | 4974.11
B | 6.15 | 4969.42
B | 6.15 | 4976.01
B | 6.15 | 4958.73
B | 6.15 | 4961.49
B | 6.15 | 4959.05
B | 6.15 | 4964.67
B | 4.1 | 3341.11
B | 4.1 | 3343.71
B | 4.1 | 3347.16
B | 4.1 | 3346.14
B | 4.1 | 3350.77
B | 4.1 | 3343.14
B | 4.1 | 3351.99
B | 4.1 | 3347.77
B | 4.1 | 3348.99
B | 4.1 | 3349.23
B | 4.1 | 3345.69
B | 4.1 | 3337.16
B | 4.1 | 3349.14
B | 4.1 | 3347.93
B | 4.1 | 3338.38
B | 4.1 | 3347.07
B | 4.1 | 3336.27
B | 4.1 | 3339.82
B | 4.1 | 3355.74
B | 4.1 | 3346.69
B | 2.05 | 1745.02
B | 2.05 | 1751.29
B | 2.05 | 1747.36
B | 2.05 | 1749.94
B | 2.05 | 1751.49
B | 2.05 | 1746.93
B | 2.05 | 1749.68
B | 2.05 | 1743.07
B | 2.05 | 1754.44
B | 2.05 | 1747.61
B | 2.05 | 1740.81
B | 2.05 | 1742.47
B | 2.05 | 1747.62
B | 2.05 | 1752.78
B | 2.05 | 1744.45
B | 2.05 | 1743.4
B | 2.05 | 1743.53
B | 2.05 | 1747.45
B | 2.05 | 1740.06
B | 2.05 | 1742.82
B | 0 | 14.3677
B | 0 | 14.3678
B | 0 | 14.3679
B | 0 | 14.368
B | 0 | 14.3681
B | 0 | 14.3682
B | 0 | 14.3683
B | 0 | 14.3684
B | 0 | 14.3685
B | 0 | 14.3686
B | 0 | 14.3687
B | 0 | 14.3688
B | 0 | 14.3689
B | 0 | 14.369
B | 0 | 14.3691
B | 0 | 14.3692
B | 0 | 14.3693
B | 0 | 14.3694
B | 0 | 14.3695
B | 0 | 14.3696
B | NA | 5056.54
B | NA | 5043.04
B | NA | 5035.03
B | NA | 5046.81
B | NA | 5047.73
B | NA | 5054.06
B | NA | 5061.54
B | NA | 5045.5
B | NA | 5050.54
B | NA | 5048.91
B | NA | 5044.97
B | NA | 5045.02
B | NA | 5039.12
B | NA | 5034.38
B | NA | 5039.15
B | NA | 5047.41
B | NA | 5034.79
B | NA | 5048.71
B | NA | 5050.97
B | NA | 5041.31
C | 10.25 | 7743.33
C | 10.25 | 7740.09
C | 10.25 | 7742.39
C | 10.25 | 7720.41
C | 10.25 | 7744.49
C | 10.25 | 7745.83
C | 10.25 | 7744.71
C | 10.25 | 7736.41
C | 10.25 | 7759.68
C | 10.25 | 7752.92
C | 10.25 | 7733.9
C | 10.25 | 7715.73
C | 10.25 | 7746.72
C | 10.25 | 7739.96
C | 10.25 | 7735.96
C | 10.25 | 7734.04
C | 10.25 | 7732.42
C | 10.25 | 7743.76
C | 10.25 | 7747.68
C | 10.25 | 7737.52
C | 8.2 | 6343.24
C | 8.2 | 6332.34
C | 8.2 | 6336.4
C | 8.2 | 6314.5
C | 8.2 | 6351.34
C | 8.2 | 6361.37
C | 8.2 | 6353.99
C | 8.2 | 6359.59
C | 8.2 | 6366.07
C | 8.2 | 6339.34
C | 8.2 | 6353.02
C | 8.2 | 6357.32
C | 8.2 | 6350.82
C | 8.2 | 6346.43
C | 8.2 | 6345.57
C | 8.2 | 6361.27
C | 8.2 | 6351.96
C | 8.2 | 6365
C | 8.2 | 6348.01
C | 8.2 | 6363.72
C | 6.15 | 5004.54
C | 6.15 | 5002.9
C | 6.15 | 4997
C | 6.15 | 4989.56
C | 6.15 | 4991.94
C | 6.15 | 5004.8
C | 6.15 | 4999.05
C | 6.15 | 5002.2
C | 6.15 | 5004.63
C | 6.15 | 4996.62
C | 6.15 | 4999.69
C | 6.15 | 5001.7
C | 6.15 | 5000.13
C | 6.15 | 4991.21
C | 6.15 | 4990.25
C | 6.15 | 5004
C | 6.15 | 5000.62
C | 6.15 | 5000.61
C | 6.15 | 4993.85
C | 6.15 | 4993.93
C | 4.1 | 3319.03
C | 4.1 | 3313.81
C | 4.1 | 3325.38
C | 4.1 | 3309.38
C | 4.1 | 3311.33
C | 4.1 | 3310.99
C | 4.1 | 3322.18
C | 4.1 | 3315.32
C | 4.1 | 3314.11
C | 4.1 | 3315.29
C | 4.1 | 3314.18
C | 4.1 | 3314.26
C | 4.1 | 3314.34
C | 4.1 | 3310.02
C | 4.1 | 3318.18
C | 4.1 | 3320.41
C | 4.1 | 3309.76
C | 4.1 | 3317.94
C | 4.1 | 3314.18
C | 4.1 | 3314.56
C | 2.05 | 1747.28
C | 2.05 | 1746.27
C | 2.05 | 1749.18
C | 2.05 | 1746.16
C | 2.05 | 1746.55
C | 2.05 | 1754.39
C | 2.05 | 1751.98
C | 2.05 | 1744.4
C | 2.05 | 1744.95
C | 2.05 | 1746.32
C | 2.05 | 1745.56
C | 2.05 | 1747.41
C | 2.05 | 1746.64
C | 2.05 | 1742.35
C | 2.05 | 1751.58
C | 2.05 | 1747.44
C | 2.05 | 1742.62
C | 2.05 | 1748.05
C | 2.05 | 1742.21
C | 2.05 | 1748.07
C | 0 | 14.3677
C | 0 | 14.3678
C | 0 | 14.3679
C | 0 | 14.368
C | 0 | 14.3681
C | 0 | 14.3682
C | 0 | 14.3683
C | 0 | 14.3684
C | 0 | 14.3685
C | 0 | 14.3686
C | 0 | 14.3687
C | 0 | 14.3688
C | 0 | 14.3689
C | 0 | 14.369
C | 0 | 14.3691
C | 0 | 14.3692
C | 0 | 14.3693
C | 0 | 14.3694
C | 0 | 14.3695
C | 0 | 14.3696
C | NA | 4993.08
C | NA | 4995.59
C | NA | 5001.89
C | NA | 4992.87
C | NA | 5000.28
C | NA | 4996.64
C | NA | 4995.38
C | NA | 4992.23
C | NA | 5003.1
C | NA | 4985.26
C | NA | 4989.11
C | NA | 4989.17
C | NA | 4994.21
C | NA | 4981.9
C | NA | 4988.43
C | NA | 4989.99
C | NA | 4987.43
C | NA | 4992.23
C | NA | 4992.91
C | NA | 4984.49
C | NA | 4992.31

## View Data
Concentration | Intensite1 | Intensite2 | Intensite3 | Intensite4 | Intensite5 | Intensite6 | Intensite7 | Intensite8 | Intensite9 | Intensite10 | Intensite11 | Intensite12 | Intensite13 | Intensite14 | Intensite15 | Intensite16 | Intensite17 | Intensite18 | Intensite19 | Intensite20
10.25 | 7404.98 | 7279.91 | 7397.95 | 7323.73 | 7207.57 | 7378.42 | 7278.12 | 7373.56 | 7260.22 | 7340.2 | 7320.47 | 7371.01 | 7326.62 | 7304.43 | 7296.81 | 7428.41 | 7329.7 | 7303.09 | 7298.02 | 7316.51
10.25 | 7406.21 | 7254.29 | 7329 | 7388.47 | 7294.72 | 7303.86 | 7322.15 | 7299.29 | 7230.72 | 7347.65 | 7332.5 | 7332.75 | 7298.16 | 7468.67 | 7385.61 | 7257.06 | 7301.26 | 7316.76 | 7340.68 | 7247.55
10.25 | 7421.97 | 7319.59 | 7245.76 | 7341.79 | 7329.13 | 7366.24 | 7346.66 | 7327.17 | 7357.97 | 7304.71 | 7353.4 | 7236.34 | 7319.5 | 7312.55 | 7375.46 | 7254.56 | 7414.73 | 7326.67 | 7240.18 | 7342.78
10.25 | 7319.38 | 7385.51 | 7221.4 | 7332.61 | 7313.69 | 7344.16 | 7311.38 | 7345.56 | 7326.46 | 7363.82 | 7349.92 | 7282.4 | 7267.91 | 7386.55 | 7313.21 | 7315.24 | 7291.98 | 7349.89 | 7325.41 | 7268.96
10.25 | 7360.08 | 7327.74 | 7393.52 | 7450.71 | 7277.91 | 7324.18 | 7262.96 | 7350.73 | 7238.82 | 7321.85 | 7259.26 | 7372.3 | 7414.04 | 7248.84 | 7343.45 | 7265.12 | 7254.55 | 7349.98 | 7294.56 | 7416.4
10.25 | 7202.07 | 7357.18 | 7305.07 | 7397.4 | 7353.16 | 7295.7 | 7266.37 | 7314.03 | 7439.52 | 7267.76 | 7315.1 | 7352.07 | 7351.94 | 7331.64 | 7327.91 | 7242.62 | 7354.85 | 7348.08 | 7316.88 | 7323.09
10.25 | 7301.17 | 7379.79 | 7292.98 | 7330.35 | 7261.5 | 7361.86 | 7390.75 | 7357.42 | 7446.16 | 7345.48 | 7266.1 | 7397.42 | 7347.54 | 7353.66 | 7364.18 | 7282.85 | 7377.8 | 7290.73 | 7297.91 | 7234.88
10.25 | 7364.43 | 7354.05 | 7213.37 | 7416.11 | 7443.86 | 7274.11 | 7332.4 | 7302.64 | 7228.38 | 7303.11 | 7358.53 | 7260.92 | 7347.03 | 7402.43 | 7261.79 | 7378.81 | 7244.6 | 7311.05 | 7358.91 | 7221.06
10.25 | 7331.77 | 7425.96 | 7413.38 | 7287.24 | 7313.7 | 7288.9 | 7302.92 | 7302.9 | 7265.81 | 7266.36 | 7253.19 | 7367.88 | 7307.51 | 7309.43 | 7285.38 | 7319.37 | 7335.99 | 7283.35 | 7335.39 | 7322.98
10.25 | 7337.14 | 7239.94 | 7329.24 | 7305.5 | 7318.32 | 7373.89 | 7458.51 | 7316.71 | 7366.49 | 7286.3 | 7397.81 | 7363.15 | 7285.78 | 7335.4 | 7413.45 | 7389.06 | 7321.81 | 7351.13 | 7332.5 | 7266.28
10.25 | 7229.94 | 7327.73 | 7345.23 | 7260.21 | 7375.86 | 7325.1 | 7371.48 | 7346.35 | 7323.03 | 7323.52 | 7366.62 | 7352.91 | 7302.67 | 7330.84 | 7222.68 | 7303.01 | 7303.55 | 7286.29 | 7387.68 | 7379.76
10.25 | 7361.54 | 7211.47 | 7318.02 | 7327.85 | 7350.14 | 7334.91 | 7277.02 | 7346.95 | 7333.95 | 7374.24 | 7362.49 | 7368.11 | 7417.86 | 7299.75 | 7348.73 | 7374.35 | 7285.12 | 7299.8 | 7289.74 | 7259.71
10.25 | 7398.4 | 7357.49 | 7272.73 | 7284.56 | 7356.52 | 7246.08 | 7289.25 | 7390.61 | 7329.18 | 7324.89 | 7323.08 | 7264.97 | 7291.53 | 7304.69 | 7359.81 | 7238 | 7311.13 | 7329.74 | 7398.36 | 7359.48
10.25 | 7319.62 | 7327.44 | 7299.32 | 7304.33 | 7305.28 | 7356.19 | 7344.45 | 7300.62 | 7365.74 | 7262.97 | 7268.1 | 7284.67 | 7316.95 | 7284.78 | 7371.87 | 7426.57 | 7332.2 | 7281.85 | 7296.31 | 7312.52
10.25 | 7271.04 | 7273.99 | 7429.31 | 7415.94 | 7321.81 | 7328.15 | 7372.4 | 7322.88 | 7313.33 | 7227.4 | 7381.48 | 7235.78 | 7402.22 | 7373.73 | 7317.07 | 7339.28 | 7386.98 | 7304.05 | 7305.51 | 7386.49
10.25 | 7283.11 | 7307.33 | 7348.62 | 7317.18 | 7306.48 | 7233.94 | 7361.18 | 7369.89 | 7442.2 | 7259.97 | 7273.53 | 7309.07 | 7283.58 | 7316.5 | 7308.37 | 7221.13 | 7230.96 | 7232.09 | 7305.71 | 7301.11
10.25 | 7250.37 | 7260.76 | 7348.09 | 7318.97 | 7318.17 | 7268.21 | 7270.26 | 7288.18 | 7331.4 | 7335.91 | 7312.03 | 7290.84 | 7250.07 | 7321.54 | 7245.05 | 7333.53 | 7291.43 | 7383.16 | 7341.99 | 7281.62
10.25 | 7302.47 | 7315.04 | 7278.65 | 7330.98 | 7392.1 | 7333.31 | 7424.41 | 7297.71 | 7394.91 | 7333.48 | 7342.28 | 7359.89 | 7405.19 | 7341.44 | 7339.8 | 7337.85 | 7277.39 | 7355.03 | 7392.43 | 7287.51
10.25 | 7496.53 | 7331.16 | 7376.92 | 7281.9 | 7263.23 | 7380.21 | 7324.14 | 7244.74 | 7221.01 | 7271.26 | 7262.36 | 7338.44 | 7361.39 | 7333.4 | 7340.95 | 7297.2 | 7391.73 | 7277.65 | 7390.81 | 7345.03
10.25 | 7284.17 | 7260.82 | 7304.7 | 7311.95 | 7388.12 | 7293.78 | 7277.54 | 7376.42 | 7330.38 | 7321.35 | 7405.17 | 7301.17 | 7430.18 | 7303.59 | 7239.05 | 7354.66 | 7277.09 | 7373.61 | 7313.05 | 7328.73
8.2 | 6299.97 | 6321.65 | 6231.37 | 6296.54 | 6272.54 | 6314.11 | 6232.62 | 6241.66 | 6316.28 | 6216.93 | 6257.69 | 6373.93 | 6296.2 | 6229.06 | 6303.53 | 6237.65 | 6217.08 | 6341.52 | 6320.07 | 6284.3
8.2 | 6197.05 | 6205.84 | 6204.56 | 6261.94 | 6314.45 | 6257.87 | 6270.38 | 6331.14 | 6313.16 | 6280.77 | 6249.11 | 6285.31 | 6293.36 | 6305.38 | 6233.92 | 6254.69 | 6282.55 | 6359.96 | 6198.93 | 6280.71
8.2 | 6225.24 | 6281.6 | 6328.05 | 6346.68 | 6259.26 | 6285.55 | 6315.29 | 6228.55 | 6258.42 | 6295.21 | 6224.44 | 6302.98 | 6276.06 | 6262.71 | 6247.35 | 6314.65 | 6258.28 | 6316.9 | 6286.68 | 6351.51
8.2 | 6258.59 | 6306.11 | 6306.48 | 6263.44 | 6305.08 | 6278.33 | 6316.63 | 6280.34 | 6331.98 | 6381.48 | 6352.15 | 6393.89 | 6330.91 | 6305.65 | 6229.14 | 6257.46 | 6304.06 | 6340.57 | 6328.78 | 6289.44
8.2 | 6338.05 | 6303.28 | 6296.11 | 6289.07 | 6249.5 | 6280.14 | 6334.46 | 6355.66 | 6259.26 | 6338.37 | 6301.5 | 6244.89 | 6246.58 | 6246.46 | 6249.51 | 6280.49 | 6288.87 | 6209.66 | 6231.05 | 6272.24
8.2 | 6311.54 | 6290.59 | 6285.3 | 6308.36 | 6291.41 | 6295.46 | 6290.13 | 6307.68 | 6249.28 | 6332.9 | 6276.43 | 6303.64 | 6287.01 | 6263 | 6326.8 | 6323.63 | 6202.77 | 6305.9 | 6274.4 | 6208.64
8.2 | 6322 | 6215.38 | 6323.52 | 6314.11 | 6278.61 | 6233.55 | 6291.99 | 6357.02 | 6317.45 | 6171.54 | 6325.36 | 6176.61 | 6264.28 | 6325.4 | 6312.31 | 6304.57 | 6270.56 | 6300.3 | 6338.82 | 6288.74
8.2 | 6241.94 | 6254.27 | 6353.91 | 6260.87 | 6270.35 | 6366.63 | 6316.73 | 6298.38 | 6258.22 | 6314.09 | 6401.6 | 6305.85 | 6286.75 | 6297.08 | 6243.34 | 6271.98 | 6254.49 | 6294.33 | 6297.38 | 6284.12
8.2 | 6277.06 | 6333.62 | 6256.21 | 6250.55 | 6264.64 | 6255.5 | 6344.03 | 6269.14 | 6302.03 | 6299.2 | 6373.47 | 6351.08 | 6303.01 | 6203.05 | 6253.17 | 6276.57 | 6289.08 | 6278.5 | 6330.62 | 6230.13
8.2 | 6307.97 | 6219.73 | 6275.37 | 6260.88 | 6297.95 | 6317.39 | 6292.26 | 6279.53 | 6289.42 | 6256.6 | 6305.39 | 6258.17 | 6179.99 | 6272.44 | 6207.67 | 6305.5 | 6320.84 | 6223.48 | 6228.62 | 6218.74
8.2 | 6342.11 | 6300.32 | 6263.77 | 6224.13 | 6360.86 | 6252.07 | 6213.91 | 6259.11 | 6252.33 | 6343.32 | 6271.76 | 6306.46 | 6239.46 | 6312.59 | 6272.85 | 6214.76 | 6231.02 | 6285.92 | 6362.81 | 6327.52
8.2 | 6321.06 | 6260.77 | 6374.76 | 6275.45 | 6321.87 | 6164.88 | 6327.69 | 6269.68 | 6267.07 | 6185.01 | 6255.64 | 6249.21 | 6275.08 | 6281.63 | 6340.61 | 6282.23 | 6243.48 | 6243.66 | 6315.74 | 6319.36
8.2 | 6388.22 | 6360.49 | 6312.79 | 6322.35 | 6204.08 | 6187.78 | 6255.92 | 6288.12 | 6159.24 | 6297.39 | 6234.34 | 6305.86 | 6241.89 | 6291.08 | 6279.68 | 6230.76 | 6250.05 | 6300.53 | 6308.58 | 6317.42
8.2 | 6317.53 | 6321.32 | 6278.4 | 6260.61 | 6299.16 | 6300.71 | 6161.97 | 6238.61 | 6237.81 | 6320.47 | 6330.88 | 6329.56 | 6301.49 | 6311.5 | 6298.21 | 6318.99 | 6323.64 | 6292.09 | 6258.05 | 6372.63
8.2 | 6309.16 | 6249.5 | 6261.75 | 6313.95 | 6314.63 | 6344.07 | 6276.04 | 6250.78 | 6307.6 | 6221.09 | 6327.11 | 6198.21 | 6347.31 | 6249.52 | 6296.57 | 6355.96 | 6270.38 | 6293.38 | 6222.28 | 6290.34
8.2 | 6223.17 | 6339.89 | 6276.72 | 6221.98 | 6320.6 | 6298.64 | 6204.95 | 6331.97 | 6253.09 | 6343.24 | 6332.41 | 6321.05 | 6305.46 | 6262.27 | 6277.19 | 6198.97 | 6256.37 | 6287.3 | 6252.97 | 6274.06
8.2 | 6291.9 | 6319.89 | 6206.6 | 6193.93 | 6357.28 | 6305.05 | 6318.6 | 6199.01 | 6203.88 | 6231.96 | 6262.21 | 6268 | 6321.71 | 6275.54 | 6354.05 | 6267.02 | 6304.24 | 6374.73 | 6298.33 | 6336.78
8.2 | 6282.99 | 6288.71 | 6266.82 | 6292.6 | 6383.03 | 6286.84 | 6298.41 | 6378.18 | 6262.13 | 6345.94 | 6311 | 6278.64 | 6351.18 | 6238.55 | 6402.73 | 6258.36 | 6329.8 | 6343.14 | 6279.55 | 6245.95
8.2 | 6268.34 | 6283.48 | 6260.04 | 6309.5 | 6209.11 | 6233.86 | 6353.95 | 6412.71 | 6213.79 | 6329.48 | 6276.02 | 6284.89 | 6276.55 | 6315.76 | 6275.5 | 6311.94 | 6336.2 | 6273.81 | 6265.37 | 6328.93
8.2 | 6191.06 | 6285.96 | 6253.97 | 6203.74 | 6271.04 | 6252.02 | 6284.87 | 6310.86 | 6308.1 | 6300.72 | 6329.85 | 6335.07 | 6328.81 | 6281.31 | 6258.66 | 6309.83 | 6279.9 | 6257.95 | 6307.61 | 6356.87
6.15 | 4984.8 | 4891.06 | 4932.48 | 4947.58 | 4892.8 | 4931.46 | 4847.92 | 4939.3 | 4910.28 | 4976.47 | 4924.2 | 4981.59 | 4913.66 | 4936.74 | 4898.85 | 4853.85 | 4861.18 | 4924.3 | 4920.02 | 4921.04
6.15 | 4940.06 | 4865.58 | 4957.11 | 4907.92 | 4889.71 | 4934.89 | 4912.63 | 4912.02 | 4925.91 | 4908.95 | 4958.49 | 4889.66 | 4880.92 | 4891.77 | 4873.55 | 4911.11 | 4923.87 | 4919.07 | 4875.26 | 4890.47
6.15 | 4941.16 | 4977.04 | 4936.66 | 4927.21 | 4927.04 | 4910.88 | 4933.45 | 4921.42 | 4874.47 | 4939.36 | 4839.27 | 4883.14 | 4893.8 | 4948.6 | 4927.71 | 4833.96 | 4938.36 | 4814.67 | 4913.93 | 4866.65
6.15 | 4892.97 | 4892.87 | 4928.12 | 4890.43 | 4923.77 | 4900.92 | 4909.87 | 4908.68 | 4962.34 | 4981.46 | 4899.37 | 4879.91 | 4905.07 | 4894.9 | 4858.3 | 4935.57 | 4892.51 | 4916.46 | 4924.3 | 4891.65
6.15 | 4869.61 | 4903 | 4979.88 | 4918.93 | 4932.03 | 4898.73 | 4860.64 | 4881.3 | 4868.1 | 4839.69 | 4918.62 | 4881.01 | 4879.31 | 4945.41 | 4903.69 | 4911.79 | 4864.75 | 5005.4 | 4933.64 | 4938.27
6.15 | 4855.4 | 4900.62 | 4879.45 | 4878.4 | 4898.51 | 5000.71 | 4863.45 | 4936.96 | 4927.18 | 4951.95 | 4803.39 | 4859.85 | 4948.88 | 4846.63 | 4905.26 | 4874.94 | 4862.97 | 4907.65 | 4935.67 | 4877.9
6.15 | 4956.87 | 4869.2 | 4917.47 | 4920.09 | 4955.87 | 4949.85 | 4849.31 | 4858.55 | 4926.71 | 4930.5 | 4932.02 | 4941.74 | 4990.93 | 4894.9 | 4960.32 | 4953.95 | 4938.33 | 4902.12 | 4847.21 | 4891.37
6.15 | 4913.54 | 4845.79 | 4929.02 | 4884.12 | 4918.23 | 4972.83 | 4903.2 | 4877.2 | 4874.12 | 4951.64 | 4890.89 | 4920.81 | 4934.2 | 4975.62 | 4871.66 | 4858.65 | 4816.88 | 4838.99 | 4984.14 | 4919.87
6.15 | 4882.42 | 4929.04 | 4947.13 | 4921.09 | 4920 | 4969.8 | 4980.69 | 4886.03 | 4879.93 | 4923.43 | 4906.09 | 4875.09 | 4903.56 | 4902.2 | 4877.15 | 4904.62 | 4852.49 | 4910.98 | 4840.64 | 4934.34
6.15 | 4947.48 | 4932.03 | 4943.42 | 4919.4 | 4946.84 | 4940.39 | 4943.1 | 4907.96 | 4897.42 | 4894.39 | 4810.28 | 4952.38 | 4862.26 | 4905.85 | 4920.98 | 4986.5 | 4935.8 | 4936.41 | 4890.13 | 4916.62
6.15 | 4844.37 | 4916.83 | 4946.88 | 4871.15 | 4884.75 | 4838.82 | 4882.29 | 4937.46 | 4887.55 | 4943.76 | 4925.14 | 4966.92 | 4894.82 | 4909.02 | 4985.24 | 4861.23 | 4906.27 | 4938.83 | 4924.06 | 4946.57
6.15 | 4835.6 | 4857.49 | 4927.93 | 4938.59 | 4885.44 | 4845.32 | 4914.89 | 4928.36 | 4920.06 | 4912.29 | 4989.87 | 4896.32 | 4903.18 | 4936.88 | 4937.68 | 4897.7 | 4985.03 | 4877.4 | 4868.38 | 4917.71
6.15 | 4910.21 | 4947.98 | 4984.29 | 4904.27 | 4959.14 | 4930.89 | 4875.56 | 4894.03 | 4939.93 | 4922.35 | 4873.36 | 4878.58 | 4964.37 | 4898.87 | 4919.58 | 4934.11 | 4910.8 | 4971.83 | 4955.8 | 4852.79
6.15 | 4923.41 | 4876.34 | 4823.9 | 4925.76 | 4915.89 | 4948.19 | 4917.25 | 4980.53 | 4903.86 | 4934.85 | 4953.81 | 4936.06 | 4919.86 | 4943.38 | 4925.32 | 4875.44 | 4895.55 | 4909.01 | 4952.56 | 4878.41
6.15 | 4904.1 | 4959.38 | 4879.62 | 4926.12 | 4869.89 | 4958.55 | 4957.36 | 4943.34 | 4884.08 | 4993.87 | 4892.07 | 4901.4 | 4891.25 | 4910.23 | 4880.69 | 4855.82 | 4984.5 | 4919.43 | 4979.21 | 4896.55
6.15 | 4922.29 | 4942.66 | 4866.18 | 4943.86 | 4924.92 | 4926.28 | 4952.11 | 4910.61 | 4790.24 | 4916.98 | 4984.82 | 4882.18 | 4907.69 | 4955.14 | 4971.62 | 4891.61 | 4910.35 | 4885.9 | 4892.53 | 4875.57
6.15 | 4902.97 | 4938.85 | 4908.13 | 4959.33 | 4907.59 | 4928.39 | 4891.05 | 4891.37 | 4895.38 | 4889.31 | 4847 | 4946.24 | 4909.56 | 4947.05 | 4991.63 | 4882.4 | 4873.91 | 4951.37 | 4968.29 | 4859.68
6.15 | 4916.95 | 4916.11 | 4909.38 | 4922.14 | 4902.66 | 4879.31 | 4834.69 | 4911.69 | 4888.67 | 4888.56 | 4921.81 | 4937.91 | 4967.7 | 4930.91 | 5022.88 | 4904.85 | 4953.17 | 4829.38 | 4858.24 | 4884.86
6.15 | 4916.68 | 4894.44 | 4921.1 | 4945.26 | 4916.62 | 4986.15 | 4937.33 | 4869.56 | 4930.05 | 4870.62 | 4887.13 | 4897.18 | 4910.71 | 4922.1 | 4951.96 | 4933.45 | 4925.92 | 4900.07 | 4905.74 | 4929.51
6.15 | 4876.39 | 4945.24 | 4914.54 | 4935.75 | 4870.52 | 4914.48 | 4952.21 | 4874.65 | 4921.11 | 4889.71 | 4973.28 | 4904.9 | 4923.32 | 4907.95 | 4936.52 | 4927.15 | 4964.11 | 4908.44 | 4934.35 | 4902.41
4.1 | 3347.49 | 3332.32 | 3322.92 | 3292.55 | 3349.73 | 3291.73 | 3328.41 | 3265 | 3333.55 | 3316.87 | 3348.99 | 3323.19 | 3373.94 | 3314.59 | 3355.83 | 3301.7 | 3299.68 | 3332.52 | 3339.17 | 3332.96
4.1 | 3298.5 | 3302.1 | 3378.99 | 3317.41 | 3326.76 | 3346.03 | 3260.05 | 3325.23 | 3249.5 | 3305.84 | 3352.04 | 3313.16 | 3286.59 | 3386.43 | 3347.1 | 3307.24 | 3279.72 | 3315.75 | 3249.74 | 3451.4
4.1 | 3361.48 | 3325.34 | 3350.98 | 3319.11 | 3265.8 | 3293.13 | 3393.69 | 3382.15 | 3336.99 | 3324.75 | 3343.4 | 3349.66 | 3295.01 | 3303.38 | 3306.59 | 3281.91 | 3252.41 | 3290.54 | 3338.4 | 3344.21
4.1 | 3319.54 | 3302.09 | 3311.76 | 3319.66 | 3328.94 | 3326.63 | 3271.8 | 3380.23 | 3340.01 | 3299.55 | 3294.65 | 3285.89 | 3308.49 | 3326.81 | 3362.77 | 3379.48 | 3345.49 | 3349.18 | 3361.86 | 3280.83
4.1 | 3297.86 | 3331.51 | 3289.19 | 3316.49 | 3341.7 | 3295.17 | 3379.68 | 3287.5 | 3356.13 | 3288.96 | 3313.56 | 3302.7 | 3351.76 | 3340.74 | 3353.3 | 3316.29 | 3300.61 | 3320.62 | 3251.27 | 3341.38
4.1 | 3320.09 | 3338.51 | 3348.61 | 3346.04 | 3255.33 | 3367.68 | 3366.9 | 3301.26 | 3248.69 | 3388.33 | 3384.17 | 3259.95 | 3308.68 | 3281.75 | 3308.03 | 3285.36 | 3254.23 | 3371.02 | 3331.79 | 3367.16
4.1 | 3317.71 | 3289.23 | 3352.06 | 3272.63 | 3298.02 | 3341.59 | 3338.42 | 3349.72 | 3268.52 | 3324.87 | 3334.44 | 3337.85 | 3282.96 | 3324.11 | 3333.25 | 3353.79 | 3355.14 | 3353.16 | 3300.09 | 3303.18
4.1 | 3349.64 | 3323.94 | 3377.07 | 3313.52 | 3354.15 | 3304.4 | 3376.54 | 3330.95 | 3341.58 | 3334.77 | 3337.71 | 3304.65 | 3323.19 | 3323.53 | 3310.11 | 3382.66 | 3241.09 | 3319.99 | 3341.04 | 3293.67
4.1 | 3367.04 | 3360.12 | 3348.55 | 3365.01 | 3355.68 | 3300.72 | 3292.02 | 3333.71 | 3349.34 | 3308.44 | 3346.13 | 3337.46 | 3309.99 | 3311.55 | 3341.71 | 3310.74 | 3339.35 | 3316.56 | 3329.47 | 3337.84
4.1 | 3379.08 | 3302.05 | 3320.41 | 3323.16 | 3347.42 | 3368.4 | 3290.47 | 3248.31 | 3342.81 | 3279.47 | 3347.32 | 3296.7 | 3327.52 | 3359.37 | 3347.09 | 3262.68 | 3314.67 | 3326.83 | 3321.88 | 3334.62
4.1 | 3348.57 | 3308.45 | 3306.87 | 3439.97 | 3344.14 | 3360.1 | 3422.87 | 3367.62 | 3276.71 | 3285.68 | 3317.85 | 3261.55 | 3299.56 | 3300.94 | 3284.59 | 3359.11 | 3327.28 | 3261.89 | 3366.64 | 3299.24
4.1 | 3339.8 | 3363.94 | 3293.81 | 3343.36 | 3360.45 | 3325.41 | 3273.97 | 3331.03 | 3310.76 | 3354.14 | 3330.33 | 3271.72 | 3336.77 | 3302.15 | 3314.89 | 3301.96 | 3314.64 | 3368.76 | 3285.35 | 3329.95
4.1 | 3329.01 | 3359.08 | 3321.17 | 3339.29 | 3328.58 | 3263.62 | 3278.08 | 3333 | 3350.42 | 3321.79 | 3362.06 | 3304.61 | 3260.53 | 3276.89 | 3310.85 | 3300.14 | 3315.62 | 3331.06 | 3307.73 | 3204.33
4.1 | 3321.83 | 3344.27 | 3338.04 | 3311.96 | 3303.63 | 3332.91 | 3405.79 | 3244.39 | 3319.28 | 3341.97 | 3364.8 | 3400.43 | 3337.82 | 3320.79 | 3303.19 | 3344.64 | 3355.25 | 3314.07 | 3305.36 | 3341.65
4.1 | 3309.39 | 3299.56 | 3368.78 | 3406.04 | 3361.51 | 3354.71 | 3258.85 | 3287.66 | 3311.91 | 3372.58 | 3315.01 | 3258.32 | 3284.32 | 3270.58 | 3337.08 | 3282.85 | 3332.19 | 3259.46 | 3307.21 | 3399.22
4.1 | 3291.21 | 3310.33 | 3344.56 | 3321.51 | 3360.23 | 3286.04 | 3372.09 | 3332.67 | 3338.39 | 3318.25 | 3351.62 | 3346.62 | 3272.72 | 3343.8 | 3306.71 | 3300.67 | 3340.9 | 3365.2 | 3342.24 | 3344.34
4.1 | 3350.73 | 3336.29 | 3262.8 | 3367.1 | 3377.47 | 3371.83 | 3286.72 | 3330.91 | 3309.59 | 3294.79 | 3389.91 | 3302.65 | 3290.37 | 3262.27 | 3316.81 | 3350.13 | 3306.48 | 3318.82 | 3291.94 | 3236.18
4.1 | 3373.66 | 3353.89 | 3313.92 | 3357.93 | 3278.75 | 3356.26 | 3353.51 | 3360.33 | 3276.35 | 3377.91 | 3289.14 | 3326.89 | 3327.4 | 3352.29 | 3299.04 | 3282.58 | 3240.57 | 3264.46 | 3330.23 | 3295.56
4.1 | 3337.86 | 3356.9 | 3274.72 | 3300.1 | 3395.07 | 3302.97 | 3309.92 | 3280.71 | 3373.73 | 3320.88 | 3258.58 | 3325.38 | 3239.04 | 3274.53 | 3361.06 | 3316.97 | 3383.35 | 3334.75 | 3329.44 | 3304.75
4.1 | 3279.58 | 3371.71 | 3321.92 | 3305.09 | 3340.89 | 3273.98 | 3286.5 | 3328.46 | 3288.96 | 3355.69 | 3307.07 | 3366.94 | 3296.58 | 3320.99 | 3371.35 | 3355 | 3311.36 | 3366.27 | 3330.99 | 3298.64
2.05 | 1697.32 | 1744.37 | 1722.1 | 1738.09 | 1803.05 | 1786.46 | 1735.3 | 1726.53 | 1781.08 | 1793.58 | 1743.64 | 1719.97 | 1798.41 | 1681.29 | 1777.85 | 1803.6 | 1785.34 | 1737.39 | 1707.5 | 1745.04
2.05 | 1689.5 | 1728.03 | 1732.95 | 1690.16 | 1749.88 | 1726.53 | 1781.92 | 1723.26 | 1714.1 | 1804.78 | 1752.35 | 1756.25 | 1714.9 | 1745.26 | 1770.15 | 1748.6 | 1690.82 | 1756.86 | 1781.47 | 1769.95
2.05 | 1760.33 | 1735.62 | 1719.17 | 1767.63 | 1749.91 | 1695.34 | 1703.24 | 1811.78 | 1731.85 | 1681.53 | 1726.73 | 1758.72 | 1783.79 | 1772.93 | 1743.51 | 1716.32 | 1717.36 | 1746.42 | 1686.8 | 1722
2.05 | 1787.84 | 1734.39 | 1753.86 | 1737.81 | 1777.27 | 1795.08 | 1748.72 | 1732.14 | 1743.5 | 1736.38 | 1713.14 | 1737.37 | 1763.73 | 1789.91 | 1703.92 | 1721.46 | 1769.74 | 1719.66 | 1741.63 | 1733.02
2.05 | 1698.93 | 1726.7 | 1779.07 | 1670.49 | 1733.63 | 1720.89 | 1772 | 1733.92 | 1715.64 | 1737.99 | 1761.4 | 1754.52 | 1699.33 | 1702.37 | 1757.68 | 1753.38 | 1783.38 | 1735.58 | 1712.61 | 1718.79
2.05 | 1717.32 | 1688.71 | 1756.56 | 1799.32 | 1727.84 | 1707.03 | 1744.83 | 1740.31 | 1769.31 | 1736.21 | 1734.96 | 1768.05 | 1719.26 | 1742.65 | 1761.11 | 1804.45 | 1764.97 | 1758.37 | 1763.42 | 1725.22
2.05 | 1752.64 | 1744.58 | 1711.98 | 1744.1 | 1757.86 | 1727.5 | 1654.47 | 1704.97 | 1762.97 | 1703.56 | 1768.9 | 1717.69 | 1736.21 | 1761.25 | 1725.18 | 1670.77 | 1687.57 | 1791.4 | 1724.54 | 1710.47
2.05 | 1781.71 | 1738.01 | 1705.79 | 1733.79 | 1740.53 | 1748.49 | 1765.86 | 1762.1 | 1723.72 | 1721.33 | 1765.22 | 1725.6 | 1729.68 | 1765.76 | 1766.06 | 1764.3 | 1752.36 | 1718.13 | 1715.62 | 1697.68
2.05 | 1679.73 | 1743.25 | 1721.55 | 1769.53 | 1787.22 | 1737.92 | 1723.46 | 1770.57 | 1724.7 | 1799.3 | 1738.8 | 1766.29 | 1703.56 | 1733.82 | 1728.06 | 1712.4 | 1716.18 | 1745.27 | 1754.19 | 1770.47
2.05 | 1730.74 | 1757.43 | 1744.51 | 1753.95 | 1791.86 | 1742.23 | 1732.02 | 1761.27 | 1710.5 | 1741.02 | 1702.36 | 1720.9 | 1722.88 | 1725.81 | 1758.78 | 1732.65 | 1727.83 | 1785.31 | 1710.5 | 1784.18
2.05 | 1715.06 | 1696.75 | 1767.66 | 1762.67 | 1695.89 | 1715.31 | 1812.19 | 1730.2 | 1762.64 | 1728.66 | 1747.32 | 1726.82 | 1707.76 | 1720.94 | 1721.11 | 1736.51 | 1748.48 | 1784.66 | 1748.05 | 1767.31
2.05 | 1698.96 | 1706.79 | 1740.85 | 1749.22 | 1722.79 | 1756.05 | 1769.33 | 1790.67 | 1748.82 | 1768.82 | 1776.32 | 1753.63 | 1746.32 | 1747.99 | 1727.75 | 1733.1 | 1759.57 | 1709.92 | 1744.83 | 1764.44
2.05 | 1720.52 | 1745.07 | 1773.86 | 1735.31 | 1719.78 | 1733.88 | 1729.43 | 1736.58 | 1742.56 | 1697.54 | 1761.83 | 1731.35 | 1712.91 | 1729.96 | 1794.85 | 1774.36 | 1711.39 | 1689.55 | 1719.45 | 1793.91
2.05 | 1795.96 | 1731.27 | 1722.27 | 1731.38 | 1790.67 | 1737.35 | 1764.45 | 1701.04 | 1786.31 | 1753.81 | 1702.36 | 1823.13 | 1750.63 | 1714.96 | 1767.2 | 1663.39 | 1717.06 | 1776.93 | 1747.07 | 1749.52
2.05 | 1769.43 | 1767.76 | 1764.17 | 1717.65 | 1704.29 | 1737.22 | 1750.95 | 1729.4 | 1766.53 | 1743.44 | 1771.44 | 1724.76 | 1776.82 | 1761.62 | 1787.01 | 1737.77 | 1750.49 | 1712.64 | 1760.09 | 1758.49
2.05 | 1745.22 | 1752.7 | 1704.41 | 1689.85 | 1792.32 | 1717.35 | 1765.47 | 1768.46 | 1739.11 | 1779.79 | 1645.98 | 1750.28 | 1718.55 | 1744.2 | 1659.01 | 1701.75 | 1704.18 | 1759.71 | 1783.9 | 1766.91
2.05 | 1715.67 | 1734.93 | 1729.88 | 1781.79 | 1724.12 | 1753.33 | 1780.88 | 1687.52 | 1723.44 | 1723.91 | 1756.21 | 1707.33 | 1799.08 | 1749.67 | 1709.68 | 1755.15 | 1722.01 | 1742.78 | 1748.37 | 1726.84
2.05 | 1750 | 1764.58 | 1685.2 | 1744.8 | 1705.88 | 1759.58 | 1755.87 | 1710.01 | 1709.71 | 1778.05 | 1787.94 | 1756.48 | 1760.44 | 1737.59 | 1746.68 | 1779.88 | 1745.36 | 1719.9 | 1718.38 | 1752.77
2.05 | 1724.58 | 1754.55 | 1746 | 1715.46 | 1749.62 | 1732.02 | 1743.41 | 1818.37 | 1727.76 | 1752.16 | 1734.82 | 1719.66 | 1795.74 | 1778.09 | 1763.46 | 1747.53 | 1766.05 | 1795.97 | 1745.24 | 1732.48
2.05 | 1744.81 | 1689.85 | 1716.94 | 1774.79 | 1731.62 | 1775.84 | 1740.31 | 1795.85 | 1796.73 | 1728.58 | 1741.5 | 1756.3 | 1754.78 | 1721.73 | 1743.94 | 1800.33 | 1713.78 | 1668.05 | 1760.65 | 1731.06
0 | 120.75 | 124.55 | 111.22 | 109.72 | 113.64 | 117.6 | 111.22 | 115.38 | 112.23 | 116.98 | 118.36 | 116.19 | 120.28 | 116.96 | 116.73 | 115.81 | 119.03 | 117.7 | 111.87 | 126.82
0 | 117.21 | 113.89 | 126.83 | 111.18 | 119 | 110.29 | 111.26 | 125.3 | 123.51 | 122.84 | 124.94 | 110.49 | 129.65 | 119.5 | 124.45 | 113.28 | 127.13 | 118.88 | 112.03 | 113.15
0 | 120 | 122.68 | 127.1 | 117.09 | 124.1 | 122.88 | 120.03 | 115.91 | 109.35 | 110.41 | 111.13 | 115.67 | 116.19 | 111.08 | 116.82 | 122.15 | 114.54 | 122.45 | 124.63 | 120.39
0 | 121.37 | 115.65 | 116.77 | 118.56 | 107.02 | 109.79 | 115.96 | 116.89 | 112.43 | 118.85 | 110.73 | 117.27 | 118.45 | 118.34 | 109.11 | 120.72 | 122.55 | 120.1 | 116.75 | 114.24
0 | 122.33 | 112.04 | 110.73 | 120.77 | 121.86 | 119.97 | 101.77 | 116.37 | 116.57 | 122.15 | 112.43 | 122.36 | 120.29 | 120.77 | 114.34 | 115.46 | 126.42 | 117.18 | 127.49 | 110.56
0 | 128.6 | 114.15 | 123.85 | 111 | 116.35 | 118.95 | 111.53 | 108.47 | 112.92 | 126.66 | 116.71 | 129.88 | 113.65 | 117.21 | 124.15 | 122.95 | 122.73 | 115.18 | 121.34 | 118.39
0 | 117.58 | 119.82 | 128.64 | 120.72 | 113.36 | 119.21 | 105.66 | 108.44 | 113.26 | 120.25 | 117.02 | 124.25 | 116.25 | 117.27 | 106.6 | 111.65 | 116.05 | 115.03 | 109.72 | 111.66
0 | 121.3 | 120.51 | 125.38 | 116.76 | 118.84 | 116.04 | 119.21 | 116 | 125.49 | 112.67 | 113.11 | 121.57 | 126.81 | 113.83 | 125.1 | 124.37 | 116.6 | 110.69 | 117.91 | 111.48
0 | 116.96 | 123.44 | 119 | 111.2 | 116.88 | 111.92 | 116.21 | 111.29 | 118.41 | 114.98 | 115.46 | 114.51 | 117.51 | 117.61 | 120.93 | 122.64 | 119.29 | 116.41 | 120.79 | 121.42
0 | 118.4 | 108.85 | 117.08 | 114.39 | 117.99 | 108.15 | 119.28 | 115.87 | 128.52 | 119.68 | 115.45 | 109.36 | 111.42 | 119.28 | 117.39 | 117.86 | 106.94 | 120.23 | 110.07 | 121.12
0 | 118.87 | 119.96 | 109.76 | 120.28 | 118.08 | 112.62 | 119.47 | 119.5 | 115.47 | 117.24 | 121.58 | 112.75 | 110.67 | 119.5 | 120.51 | 120.01 | 115.81 | 113.94 | 122.45 | 120.75
0 | 111.88 | 115.99 | 123.26 | 109.64 | 122.57 | 112.47 | 119.11 | 117.79 | 128.64 | 110.48 | 122.45 | 120.45 | 119.61 | 117.64 | 120.42 | 111.92 | 121.84 | 117.2 | 115.92 | 114.94
0 | 120.34 | 116.81 | 121.04 | 115.35 | 117.75 | 114.77 | 126.48 | 113.02 | 119.56 | 116.75 | 121.29 | 120.23 | 122.97 | 110.24 | 117.62 | 123.06 | 118.08 | 119.5 | 121.69 | 115.99
0 | 117.69 | 106.7 | 117.53 | 111.5 | 115.41 | 119.29 | 119.01 | 119.38 | 114.43 | 119.08 | 122.17 | 117.95 | 122.35 | 123.12 | 110.63 | 113.27 | 114.61 | 124.05 | 120.44 | 117.07
0 | 120.43 | 118.65 | 112.14 | 114.4 | 125.73 | 114.49 | 111.32 | 119.51 | 122.94 | 116.67 | 131.84 | 114.36 | 120.06 | 121.89 | 111.44 | 118.32 | 116.29 | 123.8 | 118.3 | 108.05
0 | 120.24 | 115.94 | 117.69 | 123.97 | 116.69 | 105.46 | 122.33 | 114.06 | 120.19 | 113.82 | 122.68 | 119.18 | 124.09 | 112.19 | 122.45 | 119.1 | 118.69 | 103.52 | 113.98 | 114.45
0 | 121.5 | 120.04 | 120.22 | 126.02 | 119.55 | 116.64 | 119.52 | 122.86 | 131.14 | 114.96 | 109.34 | 121.67 | 109.32 | 114.6 | 114.81 | 116.7 | 116.94 | 125 | 120.29 | 112.65
0 | 118.47 | 120.39 | 116.12 | 114.21 | 117.41 | 121.79 | 118.72 | 118.79 | 113.12 | 120.43 | 117.22 | 124.94 | 112.12 | 121.24 | 127.8 | 121.44 | 114.25 | 122.51 | 113.67 | 132.21
0 | 112.45 | 121.21 | 119.88 | 109.72 | 121.57 | 105.09 | 122.81 | 120.38 | 130.38 | 113.99 | 114.11 | 119.59 | 107.64 | 121.19 | 115.6 | 116.62 | 119.59 | 122.27 | 122.57 | 105.11
0 | 126.35 | 123.49 | 125.17 | 128.7 | 120.66 | 123.2 | 119.34 | 114.51 | 119.92 | 117.51 | 117.22 | 116.6 | 116.33 | 110.39 | 103.32 | 113.73 | 122.71 | 115.72 | 122.05 | 119.16
NA | 4914.46 | 4906.99 | 4877.82 | 4879.96 | 4863.19 | 4877.25 | 4912.34 | 4982.26 | 4819.89 | 4845.76 | 4826.95 | 4910.02 | 4884.95 | 4881.56 | 4882.07 | 4803.59 | 4867.44 | 4853.92 | 4848.3 | 4921.55
NA | 4887.74 | 4894.58 | 4905.94 | 4829.65 | 4908.39 | 4860.54 | 4823.03 | 4885.49 | 4880.23 | 4855.93 | 4897.61 | 4911.46 | 4887.36 | 4934.65 | 4910.58 | 4924.56 | 4962.7 | 4930.21 | 4895.67 | 4930.23
NA | 4908.91 | 4972.45 | 4890.25 | 4822.11 | 4899.3 | 4917.51 | 4878.78 | 4882.78 | 4901.25 | 4820.03 | 4823.51 | 4856.05 | 4991.12 | 4843.49 | 4897.19 | 4837.02 | 4858.42 | 4940.02 | 4919.87 | 4927.9
NA | 4847.83 | 4885.33 | 4925.43 | 4891.05 | 4894.71 | 4888.94 | 4876.97 | 4885.19 | 4903.5 | 4869.33 | 4924.64 | 4886.48 | 4883.41 | 4950.57 | 4832.43 | 4829.65 | 4876.75 | 4850.15 | 4810.93 | 4755.95
NA | 4878.65 | 4859.85 | 4963.67 | 4824.61 | 4889.27 | 4802.76 | 4846.94 | 4883.08 | 4874.38 | 4743.84 | 4817.52 | 4896.23 | 4940.43 | 4802.44 | 4869.88 | 4932.54 | 4984.52 | 4805.32 | 4899.85 | 4923.41
NA | 4881.74 | 4922.87 | 4934.57 | 4935.84 | 4943.76 | 4888.8 | 4935.9 | 4942.24 | 4918.2 | 4905.01 | 4939.05 | 4897.79 | 4846.49 | 4841.98 | 4853.04 | 4940.35 | 4913.91 | 4901.57 | 4846.15 | 4850.47
NA | 4830.44 | 4999.62 | 4898.47 | 4816.2 | 4958.62 | 4911.71 | 4896.99 | 4970 | 4943.95 | 4783.78 | 4857.75 | 4883.39 | 4903.99 | 4909.6 | 4890.46 | 4835.95 | 4927.15 | 4904.49 | 4818.57 | 4853.39
NA | 4891.17 | 4919.46 | 4834.41 | 4879.91 | 4843.81 | 4889.48 | 4892.09 | 4904.58 | 4899.49 | 4868.35 | 4925.13 | 4886.72 | 4924.48 | 4774.94 | 4977.2 | 4882.83 | 4818.34 | 4881.95 | 4847.48 | 4921.97
NA | 4845.7 | 4938.87 | 4830.98 | 4901.17 | 4964.68 | 4931.99 | 4983.12 | 4798.85 | 4841.81 | 4880.35 | 4858.06 | 4928.14 | 5009.77 | 4851.83 | 4827.52 | 4907.76 | 4857.69 | 4894.6 | 4812.4 | 4989.95
NA | 4945.94 | 4923.04 | 4996.66 | 4902.98 | 4914.4 | 4853.99 | 4899.82 | 4867.83 | 4902.37 | 4886.97 | 4822.24 | 4823.87 | 4850.93 | 4801.52 | 4948.35 | 4966.73 | 4821.69 | 4868.78 | 4907.89 | 4886.85
NA | 4831.15 | 4843.22 | 4889.77 | 4892.41 | 4874.59 | 4833.37 | 4894.58 | 4904.39 | 4917.28 | 4853.44 | 4908.07 | 4848.31 | 4844.01 | 4867.45 | 4912.28 | 4909.24 | 4846.83 | 4828.5 | 4838.69 | 4937.7
NA | 4790 | 4782.07 | 4926.26 | 4816.27 | 4806.88 | 4889.49 | 4877.4 | 4843.88 | 4891.82 | 4896.39 | 4958.56 | 4855.79 | 4954.82 | 4892 | 4807.46 | 4871.58 | 4989.46 | 4842.9 | 4904.45 | 4957.38
NA | 4877.14 | 4835.61 | 4956.13 | 4916.55 | 4991.87 | 4867.83 | 4899.02 | 4816.06 | 4880.93 | 4892.91 | 4925.01 | 4851.41 | 4931.58 | 4905.83 | 4839.12 | 4953.39 | 4863.23 | 4936.19 | 4854.56 | 4859.42
NA | 4962.04 | 4885.03 | 4781.46 | 4915.49 | 4902.4 | 4831.95 | 4900.92 | 4926.43 | 4820.81 | 4896.82 | 4859.53 | 4818.52 | 4894.87 | 4870.33 | 4883.19 | 4927.27 | 4843.14 | 4936.86 | 4802.97 | 4818.36
NA | 4847.4 | 4905.84 | 4928.79 | 4899.94 | 4940.03 | 4899.53 | 4844.8 | 4830.9 | 4868.44 | 4904.14 | 4857.8 | 4870.72 | 4845.22 | 4836.82 | 4875.39 | 4855.69 | 4900.63 | 4883.22 | 4858.35 | 4762.06
NA | 4964.61 | 5003.88 | 4902.13 | 4850.66 | 4817.74 | 4919 | 4877.91 | 4769.39 | 4945.42 | 4931.55 | 4828.44 | 4859.03 | 4919.68 | 4892.7 | 4883.19 | 4879.05 | 5017.29 | 4894.72 | 4915.81 | 4742.42
NA | 4855.7 | 4863.56 | 4842.66 | 4916.45 | 4876.15 | 4875.03 | 4855.99 | 4883.26 | 4907.5 | 4863.35 | 4938.69 | 4859.7 | 4972.49 | 4829.37 | 4857.99 | 4842.27 | 4858.01 | 4920.68 | 4770.7 | 4800.98
NA | 4855.88 | 4863.95 | 4955.13 | 4964.02 | 5003.83 | 4881.38 | 4855.32 | 4933.37 | 4878.01 | 4857.58 | 4916.13 | 4912.18 | 4906.55 | 4893.98 | 4769.95 | 4883.39 | 4866.67 | 4920.74 | 4798.78 | 4877.5
NA | 4982.89 | 4924.71 | 4886.71 | 4921.17 | 4850.8 | 4782.75 | 4897.59 | 4900.02 | 4840.96 | 4913.57 | 4888.09 | 4911.75 | 4901.14 | 4819.2 | 4908.52 | 4875.88 | 4916.8 | 4851.39 | 4761.47 | 4841.38
NA | 4874.13 | 4864.45 | 4963.6 | 4915.91 | 4805.26 | 4885.11 | 4927.99 | 4833.83 | 4834.16 | 4864.82 | 4885.1 | 4919.92 | 4841.37 | 4869.74 | 4854.07 | 4860.93 | 4928.3 | 4882.71 | 4899.29 | 4864.85

## Data
Date | Operateur | Concentration | Intensite | Colonne1
2024-11-18 | A | 10.25 | 7367.83
2024-11-18 | A | 10.25 | 7263.95
2024-11-18 | A | 10.25 | 7249.46
2024-11-18 | A | 10.25 | 7308.53
2024-11-18 | A | 10.25 | 7400.25
2024-11-18 | A | 10.25 | 7330.93
2024-11-18 | A | 10.25 | 7385.26
2024-11-18 | A | 10.25 | 7333.84
2024-11-18 | A | 10.25 | 7278.31
2024-11-18 | A | 10.25 | 7318.61
2024-11-18 | A | 10.25 | 7300.5
2024-11-18 | A | 10.25 | 7284.41
2024-11-18 | A | 10.25 | 7338.53
2024-11-18 | A | 10.25 | 7381.66
2024-11-18 | A | 10.25 | 7432
2024-11-18 | A | 10.25 | 7251.22
2024-11-18 | A | 10.25 | 7374.35
2024-11-18 | A | 10.25 | 7327.14
2024-11-18 | A | 10.25 | 7333.99
2024-11-18 | A | 10.25 | 7267.45
2024-11-18 | A | 8.2 | 6389.8
2024-11-18 | A | 8.2 | 6358.86
2024-11-18 | A | 8.2 | 6273.9
2024-11-18 | A | 8.2 | 6225.39
2024-11-18 | A | 8.2 | 6268.5
2024-11-18 | A | 8.2 | 6273.57
2024-11-18 | A | 8.2 | 6245.52
2024-11-18 | A | 8.2 | 6189.69
2024-11-18 | A | 8.2 | 6290.13
2024-11-18 | A | 8.2 | 6301.78
2024-11-18 | A | 8.2 | 6300.06
2024-11-18 | A | 8.2 | 6365.03
2024-11-18 | A | 8.2 | 6334.79
2024-11-18 | A | 8.2 | 6340.71
2024-11-18 | A | 8.2 | 6327.63
2024-11-18 | A | 8.2 | 6262.26
2024-11-18 | A | 8.2 | 6273.18
2024-11-18 | A | 8.2 | 6298.54
2024-11-18 | A | 8.2 | 6240.91
2024-11-18 | A | 8.2 | 6195
2024-11-18 | A | 6.15 | 5016.79
2024-11-18 | A | 6.15 | 4915.47
2024-11-18 | A | 6.15 | 4929.34
2024-11-18 | A | 6.15 | 4868.7
2024-11-18 | A | 6.15 | 4924.22
2024-11-18 | A | 6.15 | 4872.78
2024-11-18 | A | 6.15 | 5035.98
2024-11-18 | A | 6.15 | 4891.19
2024-11-18 | A | 6.15 | 4875.23
2024-11-18 | A | 6.15 | 4892.68
2024-11-18 | A | 6.15 | 4846.29
2024-11-18 | A | 6.15 | 4889.43
2024-11-18 | A | 6.15 | 4891.3
2024-11-18 | A | 6.15 | 4932.95
2024-11-18 | A | 6.15 | 4857.66
2024-11-18 | A | 6.15 | 4977.3
2024-11-18 | A | 6.15 | 4871.78
2024-11-18 | A | 6.15 | 4955.57
2024-11-18 | A | 6.15 | 4989.38
2024-11-18 | A | 6.15 | 4927.89
2024-11-18 | A | 4.1 | 3341.95
2024-11-18 | A | 4.1 | 3369.45
2024-11-18 | A | 4.1 | 3325.12
2024-11-18 | A | 4.1 | 3344.93
2024-11-18 | A | 4.1 | 3302.1
2024-11-18 | A | 4.1 | 3295.22
2024-11-18 | A | 4.1 | 3318.66
2024-11-18 | A | 4.1 | 3375.76
2024-11-18 | A | 4.1 | 3294.91
2024-11-18 | A | 4.1 | 3343.78
2024-11-18 | A | 4.1 | 3303.86
2024-11-18 | A | 4.1 | 3322.21
2024-11-18 | A | 4.1 | 3338.89
2024-11-18 | A | 4.1 | 3347.81
2024-11-18 | A | 4.1 | 3346.01
2024-11-18 | A | 4.1 | 3306.53
2024-11-18 | A | 4.1 | 3316.64
2024-11-18 | A | 4.1 | 3233.33
2024-11-18 | A | 4.1 | 3314.18
2024-11-18 | A | 4.1 | 3297.4
2024-11-18 | A | 2.05 | 1731.23
2024-11-18 | A | 2.05 | 1716.01
2024-11-18 | A | 2.05 | 1742.52
2024-11-18 | A | 2.05 | 1748.07
2024-11-18 | A | 2.05 | 1731.08
2024-11-18 | A | 2.05 | 1750.37
2024-11-18 | A | 2.05 | 1768.19
2024-11-18 | A | 2.05 | 1763.23
2024-11-18 | A | 2.05 | 1742.57
2024-11-18 | A | 2.05 | 1753.18
2024-11-18 | A | 2.05 | 1725.24
2024-11-18 | A | 2.05 | 1724.66
2024-11-18 | A | 2.05 | 1758.39
2024-11-18 | A | 2.05 | 1755.95
2024-11-18 | A | 2.05 | 1764.43
2024-11-18 | A | 2.05 | 1752.76
2024-11-18 | A | 2.05 | 1762.23
2024-11-18 | A | 2.05 | 1746.97
2024-11-18 | A | 2.05 | 1700.05
2024-11-18 | A | 2.05 | 1755.62
2024-11-18 | A | 0 | 119.54
2024-11-18 | A | 0 | 122.07
2024-11-18 | A | 0 | 130.4
2024-11-18 | A | 0 | 119.75
2024-11-18 | A | 0 | 120.38
2024-11-18 | A | 0 | 116.79
2024-11-18 | A | 0 | 125.75
2024-11-18 | A | 0 | 115.65
2024-11-18 | A | 0 | 124.07
2024-11-18 | A | 0 | 122.1
2024-11-18 | A | 0 | 120.37
2024-11-18 | A | 0 | 113.32
2024-11-18 | A | 0 | 123.96
2024-11-18 | A | 0 | 128.88
2024-11-18 | A | 0 | 112.67
2024-11-18 | A | 0 | 118.41
2024-11-18 | A | 0 | 128.81
2024-11-18 | A | 0 | 116.96
2024-11-18 | A | 0 | 114.2
2024-11-18 | A | 0 | 115.71
2024-11-18 | A | NA | 4914.46 | 5651.629
2024-11-18 | A | NA | 4887.74 | 5620.901
2024-11-18 | A | NA | 4908.91 | 5645.2465
2024-11-18 | A | NA | 4847.83 | 5575.0045
2024-11-18 | A | NA | 4878.65 | 5610.4475
2024-11-18 | A | NA | 4881.74 | 5614.001
2024-11-18 | A | NA | 4830.44 | 5555.006
2024-11-18 | A | NA | 4891.17 | 5624.8455
2024-11-18 | A | NA | 4845.7 | 5572.555
2024-11-18 | A | NA | 4945.94 | 5687.831
2024-11-18 | A | NA | 4831.15 | 5555.8225
2024-11-18 | A | NA | 4790 | 5508.5
2024-11-18 | A | NA | 4877.14 | 5608.711
2024-11-18 | A | NA | 4962.04 | 5706.346
2024-11-18 | A | NA | 4847.4 | 5574.51
2024-11-18 | A | NA | 4964.61 | 5709.3015
2024-11-18 | A | NA | 4855.7 | 5584.055
2024-11-18 | A | NA | 4855.88 | 5584.262
2024-11-18 | A | NA | 4982.89 | 5730.3235
2024-11-18 | A | NA | 4874.13 | 5605.2495
2024-11-19 | A | 10.25 | 7260.36
2024-11-19 | A | 10.25 | 7342.97
2024-11-19 | A | 10.25 | 7271.95
2024-11-19 | A | 10.25 | 7294.51
2024-11-19 | A | 10.25 | 7287.44
2024-11-19 | A | 10.25 | 7216.3
2024-11-19 | A | 10.25 | 7310.03
2024-11-19 | A | 10.25 | 7341.1
2024-11-19 | A | 10.25 | 7323.89
2024-11-19 | A | 10.25 | 7321.04
2024-11-19 | A | 10.25 | 7185.19
2024-11-19 | A | 10.25 | 7339.93
2024-11-19 | A | 10.25 | 7370.18
2024-11-19 | A | 10.25 | 7298.95
2024-11-19 | A | 10.25 | 7423.35
2024-11-19 | A | 10.25 | 7360.47
2024-11-19 | A | 10.25 | 7388.15
2024-11-19 | A | 10.25 | 7330.75
2024-11-19 | A | 10.25 | 7436.16
2024-11-19 | A | 10.25 | 7313.06
2024-11-19 | A | 8.2 | 6287.88
2024-11-19 | A | 8.2 | 6314.25
2024-11-19 | A | 8.2 | 6292.95
2024-11-19 | A | 8.2 | 6313.39
2024-11-19 | A | 8.2 | 6330.57
2024-11-19 | A | 8.2 | 6298.52
2024-11-19 | A | 8.2 | 6342.48
2024-11-19 | A | 8.2 | 6263.19
2024-11-19 | A | 8.2 | 6345.01
2024-11-19 | A | 8.2 | 6248.07
2024-11-19 | A | 8.2 | 6302.61
2024-11-19 | A | 8.2 | 6195.8
2024-11-19 | A | 8.2 | 6318.09
2024-11-19 | A | 8.2 | 6321
2024-11-19 | A | 8.2 | 6260.85
2024-11-19 | A | 8.2 | 6346.42
2024-11-19 | A | 8.2 | 6279.29
2024-11-19 | A | 8.2 | 6254.38
2024-11-19 | A | 8.2 | 6266.57
2024-11-19 | A | 8.2 | 6329.93
2024-11-19 | A | 6.15 | 4893.57
2024-11-19 | A | 6.15 | 4850.37
2024-11-19 | A | 6.15 | 4846.6
2024-11-19 | A | 6.15 | 4875.39
2024-11-19 | A | 6.15 | 4940.5
2024-11-19 | A | 6.15 | 4949.39
2024-11-19 | A | 6.15 | 4919.4
2024-11-19 | A | 6.15 | 4938.63
2024-11-19 | A | 6.15 | 4993.53
2024-11-19 | A | 6.15 | 4919.56
2024-11-19 | A | 6.15 | 4918.41
2024-11-19 | A | 6.15 | 4932.38
2024-11-19 | A | 6.15 | 4869.3
2024-11-19 | A | 6.15 | 4846.47
2024-11-19 | A | 6.15 | 4922.04
2024-11-19 | A | 6.15 | 4893.81
2024-11-19 | A | 6.15 | 4938.3
2024-11-19 | A | 6.15 | 4898.63
2024-11-19 | A | 6.15 | 4884.72
2024-11-19 | A | 6.15 | 4904.94
2024-11-19 | A | 4.1 | 3292.84
2024-11-19 | A | 4.1 | 3312.31
2024-11-19 | A | 4.1 | 3344.23
2024-11-19 | A | 4.1 | 3268.2
2024-11-19 | A | 4.1 | 3331.8
2024-11-19 | A | 4.1 | 3267.42
2024-11-19 | A | 4.1 | 3311.62
2024-11-19 | A | 4.1 | 3325.36
2024-11-19 | A | 4.1 | 3335.03
2024-11-19 | A | 4.1 | 3298.19
2024-11-19 | A | 4.1 | 3340.48
2024-11-19 | A | 4.1 | 3300.2
2024-11-19 | A | 4.1 | 3330.29
2024-11-19 | A | 4.1 | 3324.69
2024-11-19 | A | 4.1 | 3335.86
2024-11-19 | A | 4.1 | 3320.66
2024-11-19 | A | 4.1 | 3274.43
2024-11-19 | A | 4.1 | 3295.78
2024-11-19 | A | 4.1 | 3349.02
2024-11-19 | A | 4.1 | 3294.33
2024-11-19 | A | 2.05 | 1710.06
2024-11-19 | A | 2.05 | 1730.14
2024-11-19 | A | 2.05 | 1702.6
2024-11-19 | A | 2.05 | 1757.64
2024-11-19 | A | 2.05 | 1760.56
2024-11-19 | A | 2.05 | 1739.23
2024-11-19 | A | 2.05 | 1727.96
2024-11-19 | A | 2.05 | 1729.78
2024-11-19 | A | 2.05 | 1775.75
2024-11-19 | A | 2.05 | 1752.52
2024-11-19 | A | 2.05 | 1765.65
2024-11-19 | A | 2.05 | 1712.01
2024-11-19 | A | 2.05 | 1753.61
2024-11-19 | A | 2.05 | 1723.56
2024-11-19 | A | 2.05 | 1760.8
2024-11-19 | A | 2.05 | 1737.28
2024-11-19 | A | 2.05 | 1787.86
2024-11-19 | A | 2.05 | 1739.03
2024-11-19 | A | 2.05 | 1723.48
2024-11-19 | A | 2.05 | 1735.7
2024-11-19 | A | 0 | 117.29
2024-11-19 | A | 0 | 115.94
2024-11-19 | A | 0 | 114.04
2024-11-19 | A | 0 | 121.04
2024-11-19 | A | 0 | 128.71
2024-11-19 | A | 0 | 118.85
2024-11-19 | A | 0 | 115.78
2024-11-19 | A | 0 | 112.66
2024-11-19 | A | 0 | 117.55
2024-11-19 | A | 0 | 110.67
2024-11-19 | A | 0 | 116.94
2024-11-19 | A | 0 | 112.19
2024-11-19 | A | 0 | 119.77
2024-11-19 | A | 0 | 112.86
2024-11-19 | A | 0 | 124.53
2024-11-19 | A | 0 | 119.21
2024-11-19 | A | 0 | 119.71
2024-11-19 | A | 0 | 112.39
2024-11-19 | A | 0 | 120.03
2024-11-19 | A | 0 | 110.72
2024-11-19 | A | NA | 4906.99
2024-11-19 | A | NA | 4894.58
2024-11-19 | A | NA | 4972.45
2024-11-19 | A | NA | 4885.33
2024-11-19 | A | NA | 4859.85
2024-11-19 | A | NA | 4922.87
2024-11-19 | A | NA | 4999.62
2024-11-19 | A | NA | 4919.46
2024-11-19 | A | NA | 4938.87
2024-11-19 | A | NA | 4923.04
2024-11-19 | A | NA | 4843.22
2024-11-19 | A | NA | 4782.07
2024-11-19 | A | NA | 4835.61
2024-11-19 | A | NA | 4885.03
2024-11-19 | A | NA | 4905.84
2024-11-19 | A | NA | 5003.88
2024-11-19 | A | NA | 4863.56
2024-11-19 | A | NA | 4863.95
2024-11-19 | A | NA | 4924.71
2024-11-19 | A | NA | 4864.45
2024-11-20 | A | 10.25 | 7296.68
2024-11-20 | A | 10.25 | 7386.84
2024-11-20 | A | 10.25 | 7467.84
2024-11-20 | A | 10.25 | 7276.16
2024-11-20 | A | 10.25 | 7282.41
2024-11-20 | A | 10.25 | 7254.4
2024-11-20 | A | 10.25 | 7377.56
2024-11-20 | A | 10.25 | 7269.14
2024-11-20 | A | 10.25 | 7467.23
2024-11-20 | A | 10.25 | 7313.41
2024-11-20 | A | 10.25 | 7385.84
2024-11-20 | A | 10.25 | 7290.59
2024-11-20 | A | 10.25 | 7327.62
2024-11-20 | A | 10.25 | 7307.8
2024-11-20 | A | 10.25 | 7365.14
2024-11-20 | A | 10.25 | 7349.67
2024-11-20 | A | 10.25 | 7346.62
2024-11-20 | A | 10.25 | 7349.22
2024-11-20 | A | 10.25 | 7389.99
2024-11-20 | A | 10.25 | 7239.14
2024-11-20 | A | 8.2 | 6294.24
2024-11-20 | A | 8.2 | 6228.27
2024-11-20 | A | 8.2 | 6249.02
2024-11-20 | A | 8.2 | 6296.21
2024-11-20 | A | 8.2 | 6240.62
2024-11-20 | A | 8.2 | 6250.12
2024-11-20 | A | 8.2 | 6282.56
2024-11-20 | A | 8.2 | 6312.84
2024-11-20 | A | 8.2 | 6214.89
2024-11-20 | A | 8.2 | 6222.88
2024-11-20 | A | 8.2 | 6292.15
2024-11-20 | A | 8.2 | 6231.05
2024-11-20 | A | 8.2 | 6314.05
2024-11-20 | A | 8.2 | 6369.35
2024-11-20 | A | 8.2 | 6294.95
2024-11-20 | A | 8.2 | 6244.1
2024-11-20 | A | 8.2 | 6267.9
2024-11-20 | A | 8.2 | 6298.35
2024-11-20 | A | 8.2 | 6315.89
2024-11-20 | A | 8.2 | 6268.51
2024-11-20 | A | 6.15 | 4833.39
2024-11-20 | A | 6.15 | 4944.48
2024-11-20 | A | 6.15 | 4855.81
2024-11-20 | A | 6.15 | 4887.72
2024-11-20 | A | 6.15 | 4909.73
2024-11-20 | A | 6.15 | 4957.51
2024-11-20 | A | 6.15 | 4867.71
2024-11-20 | A | 6.15 | 4890.4
2024-11-20 | A | 6.15 | 4968.9
2024-11-20 | A | 6.15 | 4857.58
2024-11-20 | A | 6.15 | 4956.14
2024-11-20 | A | 6.15 | 4998.3
2024-11-20 | A | 6.15 | 4941.65
2024-11-20 | A | 6.15 | 4927.03
2024-11-20 | A | 6.15 | 4881.15
2024-11-20 | A | 6.15 | 4874.16
2024-11-20 | A | 6.15 | 4841.47
2024-11-20 | A | 6.15 | 4883.16
2024-11-20 | A | 6.15 | 4991.35
2024-11-20 | A | 6.15 | 4951.05
2024-11-20 | A | 4.1 | 3335.59
2024-11-20 | A | 4.1 | 3317.48
2024-11-20 | A | 4.1 | 3328.9
2024-11-20 | A | 4.1 | 3410.92
2024-11-20 | A | 4.1 | 3362.37
2024-11-20 | A | 4.1 | 3333.9
2024-11-20 | A | 4.1 | 3290.27
2024-11-20 | A | 4.1 | 3300.26
2024-11-20 | A | 4.1 | 3323.96
2024-11-20 | A | 4.1 | 3301.43
2024-11-20 | A | 4.1 | 3324.74
2024-11-20 | A | 4.1 | 3342.14
2024-11-20 | A | 4.1 | 3340.47
2024-11-20 | A | 4.1 | 3290.15
2024-11-20 | A | 4.1 | 3302.06
2024-11-20 | A | 4.1 | 3338.4
2024-11-20 | A | 4.1 | 3315.67
2024-11-20 | A | 4.1 | 3303.28
2024-11-20 | A | 4.1 | 3257.61
2024-11-20 | A | 4.1 | 3289.92
2024-11-20 | A | 2.05 | 1752.99
2024-11-20 | A | 2.05 | 1745.42
2024-11-20 | A | 2.05 | 1758.9
2024-11-20 | A | 2.05 | 1722.66
2024-11-20 | A | 2.05 | 1763.42
2024-11-20 | A | 2.05 | 1707.73
2024-11-20 | A | 2.05 | 1732.57
2024-11-20 | A | 2.05 | 1716.02
2024-11-20 | A | 2.05 | 1728.45
2024-11-20 | A | 2.05 | 1766.69
2024-11-20 | A | 2.05 | 1713.35
2024-11-20 | A | 2.05 | 1729.1
2024-11-20 | A | 2.05 | 1769.6
2024-11-20 | A | 2.05 | 1762.35
2024-11-20 | A | 2.05 | 1738.1
2024-11-20 | A | 2.05 | 1731.04
2024-11-20 | A | 2.05 | 1790.72
2024-11-20 | A | 2.05 | 1774.48
2024-11-20 | A | 2.05 | 1729.6
2024-11-20 | A | 2.05 | 1716.7
2024-11-20 | A | 0 | 121.04
2024-11-20 | A | 0 | 121.12
2024-11-20 | A | 0 | 119.79
2024-11-20 | A | 0 | 118.5
2024-11-20 | A | 0 | 113.02
2024-11-20 | A | 0 | 113.78
2024-11-20 | A | 0 | 114.33
2024-11-20 | A | 0 | 116.78
2024-11-20 | A | 0 | 116.72
2024-11-20 | A | 0 | 121.77
2024-11-20 | A | 0 | 121.15
2024-11-20 | A | 0 | 115.66
2024-11-20 | A | 0 | 121.5
2024-11-20 | A | 0 | 121.43
2024-11-20 | A | 0 | 118.67
2024-11-20 | A | 0 | 126.22
2024-11-20 | A | 0 | 113.42
2024-11-20 | A | 0 | 120.84
2024-11-20 | A | 0 | 113.7
2024-11-20 | A | 0 | 119.54
2024-11-20 | A | NA | 4877.82
2024-11-20 | A | NA | 4905.94
2024-11-20 | A | NA | 4890.25
2024-11-20 | A | NA | 4925.43
2024-11-20 | A | NA | 4963.67
2024-11-20 | A | NA | 4934.57
2024-11-20 | A | NA | 4898.47
2024-11-20 | A | NA | 4834.41
2024-11-20 | A | NA | 4830.98
2024-11-20 | A | NA | 4996.66
2024-11-20 | A | NA | 4889.77
2024-11-20 | A | NA | 4926.26
2024-11-20 | A | NA | 4956.13
2024-11-20 | A | NA | 4781.46
2024-11-20 | A | NA | 4928.79
2024-11-20 | A | NA | 4902.13
2024-11-20 | A | NA | 4842.66
2024-11-20 | A | NA | 4955.13
2024-11-20 | A | NA | 4886.71
2024-11-20 | A | NA | 4963.6
2024-11-21 | A | 10.25 | 7283.39
2024-11-21 | A | 10.25 | 7319.79
2024-11-21 | A | 10.25 | 7268.78
2024-11-21 | A | 10.25 | 7359.06
2024-11-21 | A | 10.25 | 7333.1
2024-11-21 | A | 10.25 | 7349.73
2024-11-21 | A | 10.25 | 7258.02
2024-11-21 | A | 10.25 | 7338.11
2024-11-21 | A | 10.25 | 7344.12
2024-11-21 | A | 10.25 | 7337.17
2024-11-21 | A | 10.25 | 7260.97
2024-11-21 | A | 10.25 | 7421.4
2024-11-21 | A | 10.25 | 7307.77
2024-11-21 | A | 10.25 | 7399.19
2024-11-21 | A | 10.25 | 7253.94
2024-11-21 | A | 10.25 | 7235.07
2024-11-21 | A | 10.25 | 7389.93
2024-11-21 | A | 10.25 | 7422.77
2024-11-21 | A | 10.25 | 7362.71
2024-11-21 | A | 10.25 | 7263.59
2024-11-21 | A | 8.2 | 6224.12
2024-11-21 | A | 8.2 | 6243.23
2024-11-21 | A | 8.2 | 6298.12
2024-11-21 | A | 8.2 | 6309.09
2024-11-21 | A | 8.2 | 6237.03
2024-11-21 | A | 8.2 | 6227.93
2024-11-21 | A | 8.2 | 6315.43
2024-11-21 | A | 8.2 | 6288.01
2024-11-21 | A | 8.2 | 6216.61
2024-11-21 | A | 8.2 | 6282.93
2024-11-21 | A | 8.2 | 6317.71
2024-11-21 | A | 8.2 | 6321.91
2024-11-21 | A | 8.2 | 6361.26
2024-11-21 | A | 8.2 | 6280.1
2024-11-21 | A | 8.2 | 6299.99
2024-11-21 | A | 8.2 | 6147.24
2024-11-21 | A | 8.2 | 6334.19
2024-11-21 | A | 8.2 | 6328.35
2024-11-21 | A | 8.2 | 6257.51
2024-11-21 | A | 8.2 | 6261.85
2024-11-21 | A | 6.15 | 4923.07
2024-11-21 | A | 6.15 | 4906.29
2024-11-21 | A | 6.15 | 4933.93
2024-11-21 | A | 6.15 | 4907.76
2024-11-21 | A | 6.15 | 4914.8
2024-11-21 | A | 6.15 | 4869.38
2024-11-21 | A | 6.15 | 4902.83
2024-11-21 | A | 6.15 | 4903.01
2024-11-21 | A | 6.15 | 4905.06
2024-11-21 | A | 6.15 | 4908.53
2024-11-21 | A | 6.15 | 4925.63
2024-11-21 | A | 6.15 | 4961.71
2024-11-21 | A | 6.15 | 4925.48
2024-11-21 | A | 6.15 | 4896.59
2024-11-21 | A | 6.15 | 4982.95
2024-11-21 | A | 6.15 | 4948.48
2024-11-21 | A | 6.15 | 4930.23
2024-11-21 | A | 6.15 | 4907.92
2024-11-21 | A | 6.15 | 4873.34
2024-11-21 | A | 6.15 | 4929.79
2024-11-21 | A | 4.1 | 3360.73
2024-11-21 | A | 4.1 | 3351.99
2024-11-21 | A | 4.1 | 3356.36
2024-11-21 | A | 4.1 | 3305.95
2024-11-21 | A | 4.1 | 3289.39
2024-11-21 | A | 4.1 | 3374.49
2024-11-21 | A | 4.1 | 3275.4
2024-11-21 | A | 4.1 | 3317.94
2024-11-21 | A | 4.1 | 3323.4
2024-11-21 | A | 4.1 | 3275.01
2024-11-21 | A | 4.1 | 3377.43
2024-11-21 | A | 4.1 | 3264.84
2024-11-21 | A | 4.1 | 3323.27
2024-11-21 | A | 4.1 | 3317.29
2024-11-21 | A | 4.1 | 3278.54
2024-11-21 | A | 4.1 | 3388.98
2024-11-21 | A | 4.1 | 3331.69
2024-11-21 | A | 4.1 | 3314.77
2024-11-21 | A | 4.1 | 3317.67
2024-11-21 | A | 4.1 | 3357.49
2024-11-21 | A | 2.05 | 1711.53
2024-11-21 | A | 2.05 | 1699.42
2024-11-21 | A | 2.05 | 1717.94
2024-11-21 | A | 2.05 | 1706.42
2024-11-21 | A | 2.05 | 1723.83
2024-11-21 | A | 2.05 | 1740.29
2024-11-21 | A | 2.05 | 1721.61
2024-11-21 | A | 2.05 | 1764.22
2024-11-21 | A | 2.05 | 1773.49
2024-11-21 | A | 2.05 | 1782.08
2024-11-21 | A | 2.05 | 1736.54
2024-11-21 | A | 2.05 | 1743.03
2024-11-21 | A | 2.05 | 1789.61
2024-11-21 | A | 2.05 | 1738.82
2024-11-21 | A | 2.05 | 1774.88
2024-11-21 | A | 2.05 | 1762.35
2024-11-21 | A | 2.05 | 1693.59
2024-11-21 | A | 2.05 | 1755.74
2024-11-21 | A | 2.05 | 1732.3
2024-11-21 | A | 2.05 | 1774.02
2024-11-21 | A | 0 | 122.01
2024-11-21 | A | 0 | 116.21
2024-11-21 | A | 0 | 116.36
2024-11-21 | A | 0 | 116.73
2024-11-21 | A | 0 | 119.6
2024-11-21 | A | 0 | 115.57
2024-11-21 | A | 0 | 120.07
2024-11-21 | A | 0 | 117.5
2024-11-21 | A | 0 | 121.25
2024-11-21 | A | 0 | 126.51
2024-11-21 | A | 0 | 118.03
2024-11-21 | A | 0 | 115.85
2024-11-21 | A | 0 | 113.21
2024-11-21 | A | 0 | 118.43
2024-11-21 | A | 0 | 116.69
2024-11-21 | A | 0 | 113.3
2024-11-21 | A | 0 | 120.57
2024-11-21 | A | 0 | 110.72
2024-11-21 | A | 0 | 115.26
2024-11-21 | A | 0 | 119.51
2024-11-21 | A | NA | 4879.96
2024-11-21 | A | NA | 4829.65
2024-11-21 | A | NA | 4822.11
2024-11-21 | A | NA | 4891.05
2024-11-21 | A | NA | 4824.61
2024-11-21 | A | NA | 4935.84
2024-11-21 | A | NA | 4816.2
2024-11-21 | A | NA | 4879.91
2024-11-21 | A | NA | 4901.17
2024-11-21 | A | NA | 4902.98
2024-11-21 | A | NA | 4892.41
2024-11-21 | A | NA | 4816.27
2024-11-21 | A | NA | 4916.55
2024-11-21 | A | NA | 4915.49
2024-11-21 | A | NA | 4899.94
2024-11-21 | A | NA | 4850.66
2024-11-21 | A | NA | 4916.45
2024-11-21 | A | NA | 4964.02
2024-11-21 | A | NA | 4921.17
2024-11-21 | A | NA | 4915.91
2024-11-22 | A | 10.25 | 7287.97
2024-11-22 | A | 10.25 | 7441.28
2024-11-22 | A | 10.25 | 7304.32
2024-11-22 | A | 10.25 | 7318.78
2024-11-22 | A | 10.25 | 7368.99
2024-11-22 | A | 10.25 | 7384.35
2024-11-22 | A | 10.25 | 7332.89
2024-11-22 | A | 10.25 | 7331.8
2024-11-22 | A | 10.25 | 7200.54
2024-11-22 | A | 10.25 | 7396.95
2024-11-22 | A | 10.25 | 7215.41
2024-11-22 | A | 10.25 | 7341.38
2024-11-22 | A | 10.25 | 7339.51
2024-11-22 | A | 10.25 | 7342.66
2024-11-22 | A | 10.25 | 7281.71
2024-11-22 | A | 10.25 | 7233.59
2024-11-22 | A | 10.25 | 7339.94
2024-11-22 | A | 10.25 | 7320.66
2024-11-22 | A | 10.25 | 7367.88
2024-11-22 | A | 10.25 | 7351.81
2024-11-22 | A | 8.2 | 6313.98
2024-11-22 | A | 8.2 | 6292.86
2024-11-22 | A | 8.2 | 6274.07
2024-11-22 | A | 8.2 | 6325.05
2024-11-22 | A | 8.2 | 6257.35
2024-11-22 | A | 8.2 | 6225.19
2024-11-22 | A | 8.2 | 6330.52
2024-11-22 | A | 8.2 | 6273.08
2024-11-22 | A | 8.2 | 6272.54
2024-11-22 | A | 8.2 | 6268.81
2024-11-22 | A | 8.2 | 6361.44
2024-11-22 | A | 8.2 | 6344.47
2024-11-22 | A | 8.2 | 6312.69
2024-11-22 | A | 8.2 | 6301.75
2024-11-22 | A | 8.2 | 6227.56
2024-11-22 | A | 8.2 | 6291.41
2024-11-22 | A | 8.2 | 6252.41
2024-11-22 | A | 8.2 | 6339.11
2024-11-22 | A | 8.2 | 6301.11
2024-11-22 | A | 8.2 | 6340.36
2024-11-22 | A | 6.15 | 4931.64
2024-11-22 | A | 6.15 | 4921.58
2024-11-22 | A | 6.15 | 4938.21
2024-11-22 | A | 6.15 | 4933.76
2024-11-22 | A | 6.15 | 4940.13
2024-11-22 | A | 6.15 | 4883.44
2024-11-22 | A | 6.15 | 4930.84
2024-11-22 | A | 6.15 | 4954.49
2024-11-22 | A | 6.15 | 4939.44
2024-11-22 | A | 6.15 | 4852.74
2024-11-22 | A | 6.15 | 4997.28
2024-11-22 | A | 6.15 | 4933.19
2024-11-22 | A | 6.15 | 4901.48
2024-11-22 | A | 6.15 | 4868.89
2024-11-22 | A | 6.15 | 4948.25
2024-11-22 | A | 6.15 | 4873.84
2024-11-22 | A | 6.15 | 4887.68
2024-11-22 | A | 6.15 | 4941.94
2024-11-22 | A | 6.15 | 4868.23
2024-11-22 | A | 6.15 | 4881.32
2024-11-22 | A | 4.1 | 3329.52
2024-11-22 | A | 4.1 | 3356.68
2024-11-22 | A | 4.1 | 3328.53
2024-11-22 | A | 4.1 | 3384
2024-11-22 | A | 4.1 | 3272.4
2024-11-22 | A | 4.1 | 3376.89
2024-11-22 | A | 4.1 | 3362.71
2024-11-22 | A | 4.1 | 3296.87
2024-11-22 | A | 4.1 | 3370.14
2024-11-22 | A | 4.1 | 3379.96
2024-11-22 | A | 4.1 | 3348.5
2024-11-22 | A | 4.1 | 3325.46
2024-11-22 | A | 4.1 | 3283.19
2024-11-22 | A | 4.1 | 3314.78
2024-11-22 | A | 4.1 | 3403.94
2024-11-22 | A | 4.1 | 3409.27
2024-11-22 | A | 4.1 | 3270.09
2024-11-22 | A | 4.1 | 3303.85
2024-11-22 | A | 4.1 | 3408.57
2024-11-22 | A | 4.1 | 3296.97
2024-11-22 | A | 2.05 | 1732.41
2024-11-22 | A | 2.05 | 1747.46
2024-11-22 | A | 2.05 | 1716.26
2024-11-22 | A | 2.05 | 1744.52
2024-11-22 | A | 2.05 | 1744.93
2024-11-22 | A | 2.05 | 1713.5
2024-11-22 | A | 2.05 | 1705.23
2024-11-22 | A | 2.05 | 1758.2
2024-11-22 | A | 2.05 | 1763.01
2024-11-22 | A | 2.05 | 1745.15
2024-11-22 | A | 2.05 | 1740.67
2024-11-22 | A | 2.05 | 1766.91
2024-11-22 | A | 2.05 | 1760.22
2024-11-22 | A | 2.05 | 1685.29
2024-11-22 | A | 2.05 | 1777.13
2024-11-22 | A | 2.05 | 1748.23
2024-11-22 | A | 2.05 | 1794.91
2024-11-22 | A | 2.05 | 1759
2024-11-22 | A | 2.05 | 1726.85
2024-11-22 | A | 2.05 | 1724.61
2024-11-22 | A | 0 | 116.34
2024-11-22 | A | 0 | 119.7
2024-11-22 | A | 0 | 117.51
2024-11-22 | A | 0 | 113.7
2024-11-22 | A | 0 | 115.28
2024-11-22 | A | 0 | 117.87
2024-11-22 | A | 0 | 113
2024-11-22 | A | 0 | 120.51
2024-11-22 | A | 0 | 116.63
2024-11-22 | A | 0 | 130.37
2024-11-22 | A | 0 | 121.31
2024-11-22 | A | 0 | 120.38
2024-11-22 | A | 0 | 114.95
2024-11-22 | A | 0 | 111.28
2024-11-22 | A | 0 | 114.82
2024-11-22 | A | 0 | 118.36
2024-11-22 | A | 0 | 119.48
2024-11-22 | A | 0 | 121.07
2024-11-22 | A | 0 | 115.9
2024-11-22 | A | 0 | 122.34
2024-11-22 | A | NA | 4863.19
2024-11-22 | A | NA | 4908.39
2024-11-22 | A | NA | 4899.3
2024-11-22 | A | NA | 4894.71
2024-11-22 | A | NA | 4889.27
2024-11-22 | A | NA | 4943.76
2024-11-22 | A | NA | 4958.62
2024-11-22 | A | NA | 4843.81
2024-11-22 | A | NA | 4964.68
2024-11-22 | A | NA | 4914.4
2024-11-22 | A | NA | 4874.59
2024-11-22 | A | NA | 4806.88
2024-11-22 | A | NA | 4991.87
2024-11-22 | A | NA | 4902.4
2024-11-22 | A | NA | 4940.03
2024-11-22 | A | NA | 4817.74
2024-11-22 | A | NA | 4876.15
2024-11-22 | A | NA | 5003.83
2024-11-22 | A | NA | 4850.8
2024-11-22 | A | NA | 4805.26
2024-11-24 | A | 10.25 | 7296.15
2024-11-24 | A | 10.25 | 7327.25
2024-11-24 | A | 10.25 | 7305.73
2024-11-24 | A | 10.25 | 7414.78
2024-11-24 | A | 10.25 | 7305.81
2024-11-24 | A | 10.25 | 7341.38
2024-11-24 | A | 10.25 | 7368.15
2024-11-24 | A | 10.25 | 7383.34
2024-11-24 | A | 10.25 | 7231.62
2024-11-24 | A | 10.25 | 7250.25
2024-11-24 | A | 10.25 | 7365.22
2024-11-24 | A | 10.25 | 7247.66
2024-11-24 | A | 10.25 | 7340.15
2024-11-24 | A | 10.25 | 7301.25
2024-11-24 | A | 10.25 | 7350.78
2024-11-24 | A | 10.25 | 7331.91
2024-11-24 | A | 10.25 | 7314.41
2024-11-24 | A | 10.25 | 7345.29
2024-11-24 | A | 10.25 | 7372.13
2024-11-24 | A | 10.25 | 7306.83
2024-11-24 | A | 8.2 | 6316.88
2024-11-24 | A | 8.2 | 6334.99
2024-11-24 | A | 8.2 | 6297.27
2024-11-24 | A | 8.2 | 6344.57
2024-11-24 | A | 8.2 | 6395.48
2024-11-24 | A | 8.2 | 6353.79
2024-11-24 | A | 8.2 | 6302.38
2024-11-24 | A | 8.2 | 6317.54
2024-11-24 | A | 8.2 | 6198.23
2024-11-24 | A | 8.2 | 6316.87
2024-11-24 | A | 8.2 | 6237.89
2024-11-24 | A | 8.2 | 6300.84
2024-11-24 | A | 8.2 | 6271.01
2024-11-24 | A | 8.2 | 6274.06
2024-11-24 | A | 8.2 | 6360.27
2024-11-24 | A | 8.2 | 6198.37
2024-11-24 | A | 8.2 | 6280.27
2024-11-24 | A | 8.2 | 6190.65
2024-11-24 | A | 8.2 | 6275.93
2024-11-24 | A | 8.2 | 6275.75
2024-11-24 | A | 6.15 | 4941.11
2024-11-24 | A | 6.15 | 4922.94
2024-11-24 | A | 6.15 | 4929.93
2024-11-24 | A | 6.15 | 4932.96
2024-11-24 | A | 6.15 | 4913.12
2024-11-24 | A | 6.15 | 4892.08
2024-11-24 | A | 6.15 | 4912.77
2024-11-24 | A | 6.15 | 4899.51
2024-11-24 | A | 6.15 | 4905.37
2024-11-24 | A | 6.15 | 4940.36
2024-11-24 | A | 6.15 | 4897.3
2024-11-24 | A | 6.15 | 4968.01
2024-11-24 | A | 6.15 | 4850.64
2024-11-24 | A | 6.15 | 4945.91
2024-11-24 | A | 6.15 | 4936.41
2024-11-24 | A | 6.15 | 4962.33
2024-11-24 | A | 6.15 | 4946.79
2024-11-24 | A | 6.15 | 4934.86
2024-11-24 | A | 6.15 | 4876.31
2024-11-24 | A | 6.15 | 4917.85
2024-11-24 | A | 4.1 | 3341.25
2024-11-24 | A | 4.1 | 3301.13
2024-11-24 | A | 4.1 | 3293.27
2024-11-24 | A | 4.1 | 3362.83
2024-11-24 | A | 4.1 | 3344.41
2024-11-24 | A | 4.1 | 3322.88
2024-11-24 | A | 4.1 | 3369.7
2024-11-24 | A | 4.1 | 3289.9
2024-11-24 | A | 4.1 | 3328.75
2024-11-24 | A | 4.1 | 3344.65
2024-11-24 | A | 4.1 | 3344.12
2024-11-24 | A | 4.1 | 3347.96
2024-11-24 | A | 4.1 | 3272.72
2024-11-24 | A | 4.1 | 3388.09
2024-11-24 | A | 4.1 | 3332.51
2024-11-24 | A | 4.1 | 3317.41
2024-11-24 | A | 4.1 | 3336.59
2024-11-24 | A | 4.1 | 3334.41
2024-11-24 | A | 4.1 | 3360.85
2024-11-24 | A | 4.1 | 3315.68
2024-11-24 | A | 2.05 | 1708.83
2024-11-24 | A | 2.05 | 1778.74
2024-11-24 | A | 2.05 | 1725.18
2024-11-24 | A | 2.05 | 1747.52
2024-11-24 | A | 2.05 | 1710.56
2024-11-24 | A | 2.05 | 1753.04
2024-11-24 | A | 2.05 | 1722.64
2024-11-24 | A | 2.05 | 1784.93
2024-11-24 | A | 2.05 | 1729.67
2024-11-24 | A | 2.05 | 1740.84
2024-11-24 | A | 2.05 | 1767.95
2024-11-24 | A | 2.05 | 1723.94
2024-11-24 | A | 2.05 | 1764.11
2024-11-24 | A | 2.05 | 1798.37
2024-11-24 | A | 2.05 | 1742.56
2024-11-24 | A | 2.05 | 1782.69
2024-11-24 | A | 2.05 | 1739.71
2024-11-24 | A | 2.05 | 1772.63
2024-11-24 | A | 2.05 | 1773.19
2024-11-24 | A | 2.05 | 1773.07
2024-11-24 | A | 0 | 121.2
2024-11-24 | A | 0 | 121.97
2024-11-24 | A | 0 | 121.56
2024-11-24 | A | 0 | 122.1
2024-11-24 | A | 0 | 110.5
2024-11-24 | A | 0 | 119.85
2024-11-24 | A | 0 | 113.36
2024-11-24 | A | 0 | 121.55
2024-11-24 | A | 0 | 123.06
2024-11-24 | A | 0 | 113
2024-11-24 | A | 0 | 114.09
2024-11-24 | A | 0 | 126.29
2024-11-24 | A | 0 | 117.46
2024-11-24 | A | 0 | 114.92
2024-11-24 | A | 0 | 120.96
2024-11-24 | A | 0 | 112.78
2024-11-24 | A | 0 | 116.78
2024-11-24 | A | 0 | 114.33
2024-11-24 | A | 0 | 120.84
2024-11-24 | A | 0 | 123.21
2024-11-24 | A | NA | 4877.25
2024-11-24 | A | NA | 4860.54
2024-11-24 | A | NA | 4917.51
2024-11-24 | A | NA | 4888.94
2024-11-24 | A | NA | 4802.76
2024-11-24 | A | NA | 4888.8
2024-11-24 | A | NA | 4911.71
2024-11-24 | A | NA | 4889.48
2024-11-24 | A | NA | 4931.99
2024-11-24 | A | NA | 4853.99
2024-11-24 | A | NA | 4833.37
2024-11-24 | A | NA | 4889.49
2024-11-24 | A | NA | 4867.83
2024-11-24 | A | NA | 4831.95
2024-11-24 | A | NA | 4899.53
2024-11-24 | A | NA | 4919
2024-11-24 | A | NA | 4875.03
2024-11-24 | A | NA | 4881.38
2024-11-24 | A | NA | 4782.75
2024-11-24 | A | NA | 4885.11
2024-11-25 | A | 10.25 | 7328.89
2024-11-25 | A | 10.25 | 7339.99
2024-11-25 | A | 10.25 | 7327.73
2024-11-25 | A | 10.25 | 7422.05
2024-11-25 | A | 10.25 | 7306.46
2024-11-25 | A | 10.25 | 7313.81
2024-11-25 | A | 10.25 | 7343.7
2024-11-25 | A | 10.25 | 7383.08
2024-11-25 | A | 10.25 | 7326.65
2024-11-25 | A | 10.25 | 7453.93
2024-11-25 | A | 10.25 | 7237.41
2024-11-25 | A | 10.25 | 7316.33
2024-11-25 | A | 10.25 | 7331.97
2024-11-25 | A | 10.25 | 7367
2024-11-25 | A | 10.25 | 7355.75
2024-11-25 | A | 10.25 | 7217.07
2024-11-25 | A | 10.25 | 7332.73
2024-11-25 | A | 10.25 | 7284.65
2024-11-25 | A | 10.25 | 7312.64
2024-11-25 | A | 10.25 | 7309.87
2024-11-25 | A | 8.2 | 6268.65
2024-11-25 | A | 8.2 | 6313.33
2024-11-25 | A | 8.2 | 6246.29
2024-11-25 | A | 8.2 | 6318.47
2024-11-25 | A | 8.2 | 6290.45
2024-11-25 | A | 8.2 | 6244.12
2024-11-25 | A | 8.2 | 6277.6
2024-11-25 | A | 8.2 | 6313.86
2024-11-25 | A | 8.2 | 6353.54
2024-11-25 | A | 8.2 | 6344.41
2024-11-25 | A | 8.2 | 6259.08
2024-11-25 | A | 8.2 | 6338.75
2024-11-25 | A | 8.2 | 6287.54
2024-11-25 | A | 8.2 | 6304.07
2024-11-25 | A | 8.2 | 6242.12
2024-11-25 | A | 8.2 | 6290
2024-11-25 | A | 8.2 | 6352.21
2024-11-25 | A | 8.2 | 6280.36
2024-11-25 | A | 8.2 | 6210.35
2024-11-25 | A | 8.2 | 6250.94
2024-11-25 | A | 6.15 | 4846.25
2024-11-25 | A | 6.15 | 4951.33
2024-11-25 | A | 6.15 | 4959.16
2024-11-25 | A | 6.15 | 4932.11
2024-11-25 | A | 6.15 | 4921.33
2024-11-25 | A | 6.15 | 4860.08
2024-11-25 | A | 6.15 | 4931.86
2024-11-25 | A | 6.15 | 4930
2024-11-25 | A | 6.15 | 4911.74
2024-11-25 | A | 6.15 | 4871.43
2024-11-25 | A | 6.15 | 4874.94
2024-11-25 | A | 6.15 | 4858.15
2024-11-25 | A | 6.15 | 4903.64
2024-11-25 | A | 6.15 | 4883.36
2024-11-25 | A | 6.15 | 4946.41
2024-11-25 | A | 6.15 | 4863.48
2024-11-25 | A | 6.15 | 4947.99
2024-11-25 | A | 6.15 | 4931.37
2024-11-25 | A | 6.15 | 4839.52
2024-11-25 | A | 6.15 | 4875.25
2024-11-25 | A | 4.1 | 3333.82
2024-11-25 | A | 4.1 | 3329.63
2024-11-25 | A | 4.1 | 3323.06
2024-11-25 | A | 4.1 | 3370.15
2024-11-25 | A | 4.1 | 3337.12
2024-11-25 | A | 4.1 | 3294.6
2024-11-25 | A | 4.1 | 3342.15
2024-11-25 | A | 4.1 | 3287.49
2024-11-25 | A | 4.1 | 3269.27
2024-11-25 | A | 4.1 | 3291.73
2024-11-25 | A | 4.1 | 3294.98
2024-11-25 | A | 4.1 | 3320.57
2024-11-25 | A | 4.1 | 3404.86
2024-11-25 | A | 4.1 | 3296.89
2024-11-25 | A | 4.1 | 3274.68
2024-11-25 | A | 4.1 | 3268.76
2024-11-25 | A | 4.1 | 3313.96
2024-11-25 | A | 4.1 | 3350.96
2024-11-25 | A | 4.1 | 3323.51
2024-11-25 | A | 4.1 | 3305.39
2024-11-25 | A | 2.05 | 1724.79
2024-11-25 | A | 2.05 | 1737.33
2024-11-25 | A | 2.05 | 1734.09
2024-11-25 | A | 2.05 | 1769.44
2024-11-25 | A | 2.05 | 1674.29
2024-11-25 | A | 2.05 | 1726.27
2024-11-25 | A | 2.05 | 1753.09
2024-11-25 | A | 2.05 | 1746.79
2024-11-25 | A | 2.05 | 1767.28
2024-11-25 | A | 2.05 | 1722.68
2024-11-25 | A | 2.05 | 1799.65
2024-11-25 | A | 2.05 | 1709.16
2024-11-25 | A | 2.05 | 1770.66
2024-11-25 | A | 2.05 | 1753.73
2024-11-25 | A | 2.05 | 1741.55
2024-11-25 | A | 2.05 | 1693.83
2024-11-25 | A | 2.05 | 1750.66
2024-11-25 | A | 2.05 | 1767
2024-11-25 | A | 2.05 | 1745.06
2024-11-25 | A | 2.05 | 1756.99
2024-11-25 | A | 0 | 120.2
2024-11-25 | A | 0 | 119.55
2024-11-25 | A | 0 | 118.85
2024-11-25 | A | 0 | 109.7
2024-11-25 | A | 0 | 118.52
2024-11-25 | A | 0 | 109.32
2024-11-25 | A | 0 | 115.98
2024-11-25 | A | 0 | 114.77
2024-11-25 | A | 0 | 122.38
2024-11-25 | A | 0 | 120.57
2024-11-25 | A | 0 | 112.94
2024-11-25 | A | 0 | 116.8
2024-11-25 | A | 0 | 118.07
2024-11-25 | A | 0 | 115.73
2024-11-25 | A | 0 | 115.19
2024-11-25 | A | 0 | 114.91
2024-11-25 | A | 0 | 121.1
2024-11-25 | A | 0 | 120.66
2024-11-25 | A | 0 | 119.32
2024-11-25 | A | 0 | 107.58
2024-11-25 | A | NA | 4912.34
2024-11-25 | A | NA | 4823.03
2024-11-25 | A | NA | 4878.78
2024-11-25 | A | NA | 4876.97
2024-11-25 | A | NA | 4846.94
2024-11-25 | A | NA | 4935.9
2024-11-25 | A | NA | 4896.99
2024-11-25 | A | NA | 4892.09
2024-11-25 | A | NA | 4983.12
2024-11-25 | A | NA | 4899.82
2024-11-25 | A | NA | 4894.58
2024-11-25 | A | NA | 4877.4
2024-11-25 | A | NA | 4899.02
2024-11-25 | A | NA | 4900.92
2024-11-25 | A | NA | 4844.8
2024-11-25 | A | NA | 4877.91
2024-11-25 | A | NA | 4855.99
2024-11-25 | A | NA | 4855.32
2024-11-25 | A | NA | 4897.59
2024-11-25 | A | NA | 4927.99
2024-11-26 | B | 10.25 | 7307.83
2024-11-26 | B | 10.25 | 7295.06
2024-11-26 | B | 10.25 | 7433.81
2024-11-26 | B | 10.25 | 7301.87
2024-11-26 | B | 10.25 | 7355.29
2024-11-26 | B | 10.25 | 7305.76
2024-11-26 | B | 10.25 | 7335.95
2024-11-26 | B | 10.25 | 7352.24
2024-11-26 | B | 10.25 | 7395.85
2024-11-26 | B | 10.25 | 7377.37
2024-11-26 | B | 10.25 | 7408.29
2024-11-26 | B | 10.25 | 7316.95
2024-11-26 | B | 10.25 | 7325.36
2024-11-26 | B | 10.25 | 7279.03
2024-11-26 | B | 10.25 | 7300.26
2024-11-26 | B | 10.25 | 7292.44
2024-11-26 | B | 10.25 | 7282.32
2024-11-26 | B | 10.25 | 7329.51
2024-11-26 | B | 10.25 | 7284.62
2024-11-26 | B | 10.25 | 7294.98
2024-11-26 | B | 8.2 | 6326.55
2024-11-26 | B | 8.2 | 6228.3
2024-11-26 | B | 8.2 | 6334.19
2024-11-26 | B | 8.2 | 6354.12
2024-11-26 | B | 8.2 | 6308.92
2024-11-26 | B | 8.2 | 6288.78
2024-11-26 | B | 8.2 | 6272.73
2024-11-26 | B | 8.2 | 6303.26
2024-11-26 | B | 8.2 | 6316.98
2024-11-26 | B | 8.2 | 6325.86
2024-11-26 | B | 8.2 | 6189.24
2024-11-26 | B | 8.2 | 6270.76
2024-11-26 | B | 8.2 | 6319.48
2024-11-26 | B | 8.2 | 6204.29
2024-11-26 | B | 8.2 | 6324.61
2024-11-26 | B | 8.2 | 6261.87
2024-11-26 | B | 8.2 | 6302.53
2024-11-26 | B | 8.2 | 6352.53
2024-11-26 | B | 8.2 | 6337.49
2024-11-26 | B | 8.2 | 6251.33
2024-11-26 | B | 6.15 | 4975.41
2024-11-26 | B | 6.15 | 4881.48
2024-11-26 | B | 6.15 | 4899.3
2024-11-26 | B | 6.15 | 4898.9
2024-11-26 | B | 6.15 | 4867
2024-11-26 | B | 6.15 | 4909.86
2024-11-26 | B | 6.15 | 4922.89
2024-11-26 | B | 6.15 | 4953.44
2024-11-26 | B | 6.15 | 4974.6
2024-11-26 | B | 6.15 | 4953.49
2024-11-26 | B | 6.15 | 5005.55
2024-11-26 | B | 6.15 | 4912.56
2024-11-26 | B | 6.15 | 4897.92
2024-11-26 | B | 6.15 | 4865.91
2024-11-26 | B | 6.15 | 4958.9
2024-11-26 | B | 6.15 | 4992.27
2024-11-26 | B | 6.15 | 4902.62
2024-11-26 | B | 6.15 | 4879.57
2024-11-26 | B | 6.15 | 4920.18
2024-11-26 | B | 6.15 | 4948.69
2024-11-26 | B | 4.1 | 3352.68
2024-11-26 | B | 4.1 | 3355.56
2024-11-26 | B | 4.1 | 3378.31
2024-11-26 | B | 4.1 | 3302.09
2024-11-26 | B | 4.1 | 3311
2024-11-26 | B | 4.1 | 3368.79
2024-11-26 | B | 4.1 | 3293.01
2024-11-26 | B | 4.1 | 3350.51
2024-11-26 | B | 4.1 | 3354.82
2024-11-26 | B | 4.1 | 3397.99
2024-11-26 | B | 4.1 | 3346.17
2024-11-26 | B | 4.1 | 3372.76
2024-11-26 | B | 4.1 | 3365.2
2024-11-26 | B | 4.1 | 3287.19
2024-11-26 | B | 4.1 | 3355.24
2024-11-26 | B | 4.1 | 3340.35
2024-11-26 | B | 4.1 | 3340
2024-11-26 | B | 4.1 | 3344.44
2024-11-26 | B | 4.1 | 3359.71
2024-11-26 | B | 4.1 | 3409.79
2024-11-26 | B | 2.05 | 1682.88
2024-11-26 | B | 2.05 | 1767.28
2024-11-26 | B | 2.05 | 1720.5
2024-11-26 | B | 2.05 | 1768.22
2024-11-26 | B | 2.05 | 1728.95
2024-11-26 | B | 2.05 | 1702.88
2024-11-26 | B | 2.05 | 1721.03
2024-11-26 | B | 2.05 | 1726.57
2024-11-26 | B | 2.05 | 1704.51
2024-11-26 | B | 2.05 | 1742.58
2024-11-26 | B | 2.05 | 1740.21
2024-11-26 | B | 2.05 | 1780.13
2024-11-26 | B | 2.05 | 1702.43
2024-11-26 | B | 2.05 | 1691.82
2024-11-26 | B | 2.05 | 1737.53
2024-11-26 | B | 2.05 | 1731.63
2024-11-26 | B | 2.05 | 1720.32
2024-11-26 | B | 2.05 | 1714.45
2024-11-26 | B | 2.05 | 1689.72
2024-11-26 | B | 2.05 | 1718.13
2024-11-26 | B | 0 | 113.63
2024-11-26 | B | 0 | 117.34
2024-11-26 | B | 0 | 120.68
2024-11-26 | B | 0 | 106
2024-11-26 | B | 0 | 114.84
2024-11-26 | B | 0 | 115.98
2024-11-26 | B | 0 | 115.44
2024-11-26 | B | 0 | 117.15
2024-11-26 | B | 0 | 114.79
2024-11-26 | B | 0 | 116.42
2024-11-26 | B | 0 | 115.96
2024-11-26 | B | 0 | 122.12
2024-11-26 | B | 0 | 121.93
2024-11-26 | B | 0 | 113.19
2024-11-26 | B | 0 | 117.83
2024-11-26 | B | 0 | 116.87
2024-11-26 | B | 0 | 115.21
2024-11-26 | B | 0 | 117.68
2024-11-26 | B | 0 | 112.32
2024-11-26 | B | 0 | 118.06
2024-11-26 | B | NA | 4982.26
2024-11-26 | B | NA | 4885.49
2024-11-26 | B | NA | 4882.78
2024-11-26 | B | NA | 4885.19
2024-11-26 | B | NA | 4883.08
2024-11-26 | B | NA | 4942.24
2024-11-26 | B | NA | 4970
2024-11-26 | B | NA | 4904.58
2024-11-26 | B | NA | 4798.85
2024-11-26 | B | NA | 4867.83
2024-11-26 | B | NA | 4904.39
2024-11-26 | B | NA | 4843.88
2024-11-26 | B | NA | 4816.06
2024-11-26 | B | NA | 4926.43
2024-11-26 | B | NA | 4830.9
2024-11-26 | B | NA | 4769.39
2024-11-26 | B | NA | 4883.26
2024-11-26 | B | NA | 4933.37
2024-11-26 | B | NA | 4900.02
2024-11-26 | B | NA | 4833.83
2024-11-27 | B | 10.25 | 7400.57
2024-11-27 | B | 10.25 | 7375.33
2024-11-27 | B | 10.25 | 7246.6
2024-11-27 | B | 10.25 | 7380.15
2024-11-27 | B | 10.25 | 7372.66
2024-11-27 | B | 10.25 | 7318.6
2024-11-27 | B | 10.25 | 7295.58
2024-11-27 | B | 10.25 | 7395.38
2024-11-27 | B | 10.25 | 7405.11
2024-11-27 | B | 10.25 | 7286.62
2024-11-27 | B | 10.25 | 7199.82
2024-11-27 | B | 10.25 | 7231.97
2024-11-27 | B | 10.25 | 7383.44
2024-11-27 | B | 10.25 | 7322.11
2024-11-27 | B | 10.25 | 7344.19
2024-11-27 | B | 10.25 | 7381.38
2024-11-27 | B | 10.25 | 7311.93
2024-11-27 | B | 10.25 | 7297.96
2024-11-27 | B | 10.25 | 7288.86
2024-11-27 | B | 10.25 | 7299.12
2024-11-27 | B | 8.2 | 6342.33
2024-11-27 | B | 8.2 | 6295.97
2024-11-27 | B | 8.2 | 6324.81
2024-11-27 | B | 8.2 | 6292.42
2024-11-27 | B | 8.2 | 6274.12
2024-11-27 | B | 8.2 | 6225.82
2024-11-27 | B | 8.2 | 6308.58
2024-11-27 | B | 8.2 | 6223.96
2024-11-27 | B | 8.2 | 6272.78
2024-11-27 | B | 8.2 | 6319.62
2024-11-27 | B | 8.2 | 6326.83
2024-11-27 | B | 8.2 | 6261.3
2024-11-27 | B | 8.2 | 6290.47
2024-11-27 | B | 8.2 | 6304.86
2024-11-27 | B | 8.2 | 6295.93
2024-11-27 | B | 8.2 | 6276.77
2024-11-27 | B | 8.2 | 6283.25
2024-11-27 | B | 8.2 | 6295.08
2024-11-27 | B | 8.2 | 6258.62
2024-11-27 | B | 8.2 | 6355.29
2024-11-27 | B | 6.15 | 4913.77
2024-11-27 | B | 6.15 | 4915.69
2024-11-27 | B | 6.15 | 4902.26
2024-11-27 | B | 6.15 | 4847.44
2024-11-27 | B | 6.15 | 4966.78
2024-11-27 | B | 6.15 | 4972.51
2024-11-27 | B | 6.15 | 4893.53
2024-11-27 | B | 6.15 | 4970.76
2024-11-27 | B | 6.15 | 4862.5
2024-11-27 | B | 6.15 | 4868.09
2024-11-27 | B | 6.15 | 4930.86
2024-11-27 | B | 6.15 | 4955.55
2024-11-27 | B | 6.15 | 4938.53
2024-11-27 | B | 6.15 | 4965.25
2024-11-27 | B | 6.15 | 4898.6
2024-11-27 | B | 6.15 | 4914.2
2024-11-27 | B | 6.15 | 4900.81
2024-11-27 | B | 6.15 | 4919.34
2024-11-27 | B | 6.15 | 4913.9
2024-11-27 | B | 6.15 | 4911.43
2024-11-27 | B | 4.1 | 3359.65
2024-11-27 | B | 4.1 | 3294.41
2024-11-27 | B | 4.1 | 3364.94
2024-11-27 | B | 4.1 | 3372.8
2024-11-27 | B | 4.1 | 3312.68
2024-11-27 | B | 4.1 | 3341.2
2024-11-27 | B | 4.1 | 3311.65
2024-11-27 | B | 4.1 | 3336.14
2024-11-27 | B | 4.1 | 3274.44
2024-11-27 | B | 4.1 | 3337.04
2024-11-27 | B | 4.1 | 3295.47
2024-11-27 | B | 4.1 | 3335.46
2024-11-27 | B | 4.1 | 3321.24
2024-11-27 | B | 4.1 | 3307.81
2024-11-27 | B | 4.1 | 3269.12
2024-11-27 | B | 4.1 | 3258.79
2024-11-27 | B | 4.1 | 3256.36
2024-11-27 | B | 4.1 | 3362.38
2024-11-27 | B | 4.1 | 3281.19
2024-11-27 | B | 4.1 | 3357.12
2024-11-27 | B | 2.05 | 1778.81
2024-11-27 | B | 2.05 | 1694.72
2024-11-27 | B | 2.05 | 1757.48
2024-11-27 | B | 2.05 | 1738.73
2024-11-27 | B | 2.05 | 1749.04
2024-11-27 | B | 2.05 | 1777.06
2024-11-27 | B | 2.05 | 1770.13
2024-11-27 | B | 2.05 | 1773.69
2024-11-27 | B | 2.05 | 1753.09
2024-11-27 | B | 2.05 | 1702.05
2024-11-27 | B | 2.05 | 1755.71
2024-11-27 | B | 2.05 | 1767.25
2024-11-27 | B | 2.05 | 1742.65
2024-11-27 | B | 2.05 | 1791.93
2024-11-27 | B | 2.05 | 1747.53
2024-11-27 | B | 2.05 | 1767.48
2024-11-27 | B | 2.05 | 1814.49
2024-11-27 | B | 2.05 | 1778.28
2024-11-27 | B | 2.05 | 1755.71
2024-11-27 | B | 2.05 | 1700.17
2024-11-27 | B | 0 | 123.19
2024-11-27 | B | 0 | 113.78
2024-11-27 | B | 0 | 119.83
2024-11-27 | B | 0 | 110.87
2024-11-27 | B | 0 | 108.56
2024-11-27 | B | 0 | 110.28
2024-11-27 | B | 0 | 128.06
2024-11-27 | B | 0 | 113.25
2024-11-27 | B | 0 | 115.09
2024-11-27 | B | 0 | 117.45
2024-11-27 | B | 0 | 110.49
2024-11-27 | B | 0 | 113.58
2024-11-27 | B | 0 | 116.77
2024-11-27 | B | 0 | 116.09
2024-11-27 | B | 0 | 118.82
2024-11-27 | B | 0 | 109.64
2024-11-27 | B | 0 | 120.49
2024-11-27 | B | 0 | 113.4
2024-11-27 | B | 0 | 115.7
2024-11-27 | B | 0 | 119.03
2024-11-27 | B | NA | 4819.89
2024-11-27 | B | NA | 4880.23
2024-11-27 | B | NA | 4901.25
2024-11-27 | B | NA | 4903.5
2024-11-27 | B | NA | 4874.38
2024-11-27 | B | NA | 4918.2
2024-11-27 | B | NA | 4943.95
2024-11-27 | B | NA | 4899.49
2024-11-27 | B | NA | 4841.81
2024-11-27 | B | NA | 4902.37
2024-11-27 | B | NA | 4917.28
2024-11-27 | B | NA | 4891.82
2024-11-27 | B | NA | 4880.93
2024-11-27 | B | NA | 4820.81
2024-11-27 | B | NA | 4868.44
2024-11-27 | B | NA | 4945.42
2024-11-27 | B | NA | 4907.5
2024-11-27 | B | NA | 4878.01
2024-11-27 | B | NA | 4840.96
2024-11-27 | B | NA | 4834.16
2024-11-28 | B | 10.25 | 7251.45
2024-11-28 | B | 10.25 | 7305.85
2024-11-28 | B | 10.25 | 7349.32
2024-11-28 | B | 10.25 | 7268.95
2024-11-28 | B | 10.25 | 7321.36
2024-11-28 | B | 10.25 | 7372.4
2024-11-28 | B | 10.25 | 7362.34
2024-11-28 | B | 10.25 | 7423.95
2024-11-28 | B | 10.25 | 7366.38
2024-11-28 | B | 10.25 | 7259.19
2024-11-28 | B | 10.25 | 7283.01
2024-11-28 | B | 10.25 | 7272.9
2024-11-28 | B | 10.25 | 7341.55
2024-11-28 | B | 10.25 | 7288.72
2024-11-28 | B | 10.25 | 7377.27
2024-11-28 | B | 10.25 | 7444.69
2024-11-28 | B | 10.25 | 7352.53
2024-11-28 | B | 10.25 | 7209.92
2024-11-28 | B | 10.25 | 7296.78
2024-11-28 | B | 10.25 | 7337.89
2024-11-28 | B | 8.2 | 6256.17
2024-11-28 | B | 8.2 | 6361.79
2024-11-28 | B | 8.2 | 6258.54
2024-11-28 | B | 8.2 | 6367.68
2024-11-28 | B | 8.2 | 6333.02
2024-11-28 | B | 8.2 | 6329.34
2024-11-28 | B | 8.2 | 6300.28
2024-11-28 | B | 8.2 | 6258.74
2024-11-28 | B | 8.2 | 6171.54
2024-11-28 | B | 8.2 | 6261.87
2024-11-28 | B | 8.2 | 6226.04
2024-11-28 | B | 8.2 | 6357.73
2024-11-28 | B | 8.2 | 6296.2
2024-11-28 | B | 8.2 | 6275.09
2024-11-28 | B | 8.2 | 6244.88
2024-11-28 | B | 8.2 | 6263.56
2024-11-28 | B | 8.2 | 6324.75
2024-11-28 | B | 8.2 | 6339.37
2024-11-28 | B | 8.2 | 6243.05
2024-11-28 | B | 8.2 | 6239.41
2024-11-28 | B | 6.15 | 4880.91
2024-11-28 | B | 6.15 | 4837.27
2024-11-28 | B | 6.15 | 4996.75
2024-11-28 | B | 6.15 | 4911.31
2024-11-28 | B | 6.15 | 4917.22
2024-11-28 | B | 6.15 | 4900.11
2024-11-28 | B | 6.15 | 4880.85
2024-11-28 | B | 6.15 | 4859.02
2024-11-28 | B | 6.15 | 4852.58
2024-11-28 | B | 6.15 | 4932.3
2024-11-28 | B | 6.15 | 4977.21
2024-11-28 | B | 6.15 | 4859.73
2024-11-28 | B | 6.15 | 4920.26
2024-11-28 | B | 6.15 | 4942.52
2024-11-28 | B | 6.15 | 4959.13
2024-11-28 | B | 6.15 | 4907.4
2024-11-28 | B | 6.15 | 4975.81
2024-11-28 | B | 6.15 | 4961.3
2024-11-28 | B | 6.15 | 4909.35
2024-11-28 | B | 6.15 | 4898.11
2024-11-28 | B | 4.1 | 3301.2
2024-11-28 | B | 4.1 | 3327.55
2024-11-28 | B | 4.1 | 3309.32
2024-11-28 | B | 4.1 | 3307.22
2024-11-28 | B | 4.1 | 3353.36
2024-11-28 | B | 4.1 | 3394.83
2024-11-28 | B | 4.1 | 3303.53
2024-11-28 | B | 4.1 | 3285.25
2024-11-28 | B | 4.1 | 3340.05
2024-11-28 | B | 4.1 | 3357.14
2024-11-28 | B | 4.1 | 3333.98
2024-11-28 | B | 4.1 | 3250.09
2024-11-28 | B | 4.1 | 3352.54
2024-11-28 | B | 4.1 | 3321.95
2024-11-28 | B | 4.1 | 3368.51
2024-11-28 | B | 4.1 | 3333.16
2024-11-28 | B | 4.1 | 3355.83
2024-11-28 | B | 4.1 | 3373.56
2024-11-28 | B | 4.1 | 3340.18
2024-11-28 | B | 4.1 | 3331.21
2024-11-28 | B | 2.05 | 1752.18
2024-11-28 | B | 2.05 | 1775.33
2024-11-28 | B | 2.05 | 1719.81
2024-11-28 | B | 2.05 | 1734.75
2024-11-28 | B | 2.05 | 1746.33
2024-11-28 | B | 2.05 | 1740.84
2024-11-28 | B | 2.05 | 1766.16
2024-11-28 | B | 2.05 | 1703.6
2024-11-28 | B | 2.05 | 1741.61
2024-11-28 | B | 2.05 | 1775.21
2024-11-28 | B | 2.05 | 1724.06
2024-11-28 | B | 2.05 | 1728.55
2024-11-28 | B | 2.05 | 1685.36
2024-11-28 | B | 2.05 | 1703.01
2024-11-28 | B | 2.05 | 1764.05
2024-11-28 | B | 2.05 | 1734.04
2024-11-28 | B | 2.05 | 1783.21
2024-11-28 | B | 2.05 | 1750.13
2024-11-28 | B | 2.05 | 1713.93
2024-11-28 | B | 2.05 | 1752.42
2024-11-28 | B | 0 | 122.3
2024-11-28 | B | 0 | 114.44
2024-11-28 | B | 0 | 124.87
2024-11-28 | B | 0 | 112.24
2024-11-28 | B | 0 | 122.55
2024-11-28 | B | 0 | 115.07
2024-11-28 | B | 0 | 119.06
2024-11-28 | B | 0 | 110.19
2024-11-28 | B | 0 | 120.02
2024-11-28 | B | 0 | 123.02
2024-11-28 | B | 0 | 117.35
2024-11-28 | B | 0 | 121.77
2024-11-28 | B | 0 | 110.75
2024-11-28 | B | 0 | 121.12
2024-11-28 | B | 0 | 125.33
2024-11-28 | B | 0 | 104.78
2024-11-28 | B | 0 | 116.73
2024-11-28 | B | 0 | 122.91
2024-11-28 | B | 0 | 118.89
2024-11-28 | B | 0 | 114.28
2024-11-28 | B | NA | 4845.76
2024-11-28 | B | NA | 4855.93
2024-11-28 | B | NA | 4820.03
2024-11-28 | B | NA | 4869.33
2024-11-28 | B | NA | 4743.84
2024-11-28 | B | NA | 4905.01
2024-11-28 | B | NA | 4783.78
2024-11-28 | B | NA | 4868.35
2024-11-28 | B | NA | 4880.35
2024-11-28 | B | NA | 4886.97
2024-11-28 | B | NA | 4853.44
2024-11-28 | B | NA | 4896.39
2024-11-28 | B | NA | 4892.91
2024-11-28 | B | NA | 4896.82
2024-11-28 | B | NA | 4904.14
2024-11-28 | B | NA | 4931.55
2024-11-28 | B | NA | 4863.35
2024-11-28 | B | NA | 4857.58
2024-11-28 | B | NA | 4913.57
2024-11-28 | B | NA | 4864.82
2024-12-02 | B | 10.25 | 7281.12
2024-12-02 | B | 10.25 | 7257.35
2024-12-02 | B | 10.25 | 7292.07
2024-12-02 | B | 10.25 | 7393.85
2024-12-02 | B | 10.25 | 7264.4
2024-12-02 | B | 10.25 | 7412.75
2024-12-02 | B | 10.25 | 7392.2
2024-12-02 | B | 10.25 | 7278.77
2024-12-02 | B | 10.25 | 7381.21
2024-12-02 | B | 10.25 | 7327.93
2024-12-02 | B | 10.25 | 7374.87
2024-12-02 | B | 10.25 | 7263.38
2024-12-02 | B | 10.25 | 7280.05
2024-12-02 | B | 10.25 | 7203.75
2024-12-02 | B | 10.25 | 7336.07
2024-12-02 | B | 10.25 | 7284.18
2024-12-02 | B | 10.25 | 7375.36
2024-12-02 | B | 10.25 | 7400.42
2024-12-02 | B | 10.25 | 7250.11
2024-12-02 | B | 10.25 | 7262.97
2024-12-02 | B | 8.2 | 6320.07
2024-12-02 | B | 8.2 | 6295.47
2024-12-02 | B | 8.2 | 6268.55
2024-12-02 | B | 8.2 | 6303.12
2024-12-02 | B | 8.2 | 6306.41
2024-12-02 | B | 8.2 | 6292.94
2024-12-02 | B | 8.2 | 6305.68
2024-12-02 | B | 8.2 | 6197.19
2024-12-02 | B | 8.2 | 6332.31
2024-12-02 | B | 8.2 | 6275.18
2024-12-02 | B | 8.2 | 6320.62
2024-12-02 | B | 8.2 | 6256.44
2024-12-02 | B | 8.2 | 6343.24
2024-12-02 | B | 8.2 | 6211
2024-12-02 | B | 8.2 | 6264.76
2024-12-02 | B | 8.2 | 6279.66
2024-12-02 | B | 8.2 | 6288.31
2024-12-02 | B | 8.2 | 6349.48
2024-12-02 | B | 8.2 | 6324.76
2024-12-02 | B | 8.2 | 6276.31
2024-12-02 | B | 6.15 | 4955.74
2024-12-02 | B | 6.15 | 4944.38
2024-12-02 | B | 6.15 | 4939.38
2024-12-02 | B | 6.15 | 4984.09
2024-12-02 | B | 6.15 | 4951.98
2024-12-02 | B | 6.15 | 4957.56
2024-12-02 | B | 6.15 | 4994.74
2024-12-02 | B | 6.15 | 4895.01
2024-12-02 | B | 6.15 | 4947.23
2024-12-02 | B | 6.15 | 4863.34
2024-12-02 | B | 6.15 | 4914.07
2024-12-02 | B | 6.15 | 4858.78
2024-12-02 | B | 6.15 | 4892.21
2024-12-02 | B | 6.15 | 4992.12
2024-12-02 | B | 6.15 | 4930
2024-12-02 | B | 6.15 | 4941.81
2024-12-02 | B | 6.15 | 4928.73
2024-12-02 | B | 6.15 | 4874.41
2024-12-02 | B | 6.15 | 4961.89
2024-12-02 | B | 6.15 | 4865.82
2024-12-02 | B | 4.1 | 3364.36
2024-12-02 | B | 4.1 | 3399.78
2024-12-02 | B | 4.1 | 3312.7
2024-12-02 | B | 4.1 | 3404.71
2024-12-02 | B | 4.1 | 3294.87
2024-12-02 | B | 4.1 | 3299.73
2024-12-02 | B | 4.1 | 3391.65
2024-12-02 | B | 4.1 | 3328.65
2024-12-02 | B | 4.1 | 3314.65
2024-12-02 | B | 4.1 | 3325.18
2024-12-02 | B | 4.1 | 3295
2024-12-02 | B | 4.1 | 3337.61
2024-12-02 | B | 4.1 | 3327.39
2024-12-02 | B | 4.1 | 3295.58
2024-12-02 | B | 4.1 | 3336.3
2024-12-02 | B | 4.1 | 3322.09
2024-12-02 | B | 4.1 | 3275.14
2024-12-02 | B | 4.1 | 3307.14
2024-12-02 | B | 4.1 | 3312.57
2024-12-02 | B | 4.1 | 3285
2024-12-02 | B | 2.05 | 1784.32
2024-12-02 | B | 2.05 | 1767.86
2024-12-02 | B | 2.05 | 1754.07
2024-12-02 | B | 2.05 | 1719.73
2024-12-02 | B | 2.05 | 1772.79
2024-12-02 | B | 2.05 | 1691.26
2024-12-02 | B | 2.05 | 1699.74
2024-12-02 | B | 2.05 | 1750.32
2024-12-02 | B | 2.05 | 1760.31
2024-12-02 | B | 2.05 | 1762.44
2024-12-02 | B | 2.05 | 1732.48
2024-12-02 | B | 2.05 | 1726.22
2024-12-02 | B | 2.05 | 1770.53
2024-12-02 | B | 2.05 | 1788.35
2024-12-02 | B | 2.05 | 1756.44
2024-12-02 | B | 2.05 | 1767.29
2024-12-02 | B | 2.05 | 1722.15
2024-12-02 | B | 2.05 | 1759.52
2024-12-02 | B | 2.05 | 1706.52
2024-12-02 | B | 2.05 | 1772.7
2024-12-02 | B | 0 | 115.83
2024-12-02 | B | 0 | 123.48
2024-12-02 | B | 0 | 125.64
2024-12-02 | B | 0 | 119.19
2024-12-02 | B | 0 | 111.34
2024-12-02 | B | 0 | 118.16
2024-12-02 | B | 0 | 115.36
2024-12-02 | B | 0 | 116.63
2024-12-02 | B | 0 | 121.59
2024-12-02 | B | 0 | 125.45
2024-12-02 | B | 0 | 111.72
2024-12-02 | B | 0 | 123.71
2024-12-02 | B | 0 | 123.43
2024-12-02 | B | 0 | 119.73
2024-12-02 | B | 0 | 111.4
2024-12-02 | B | 0 | 111.89
2024-12-02 | B | 0 | 122.76
2024-12-02 | B | 0 | 114.02
2024-12-02 | B | 0 | 111.88
2024-12-02 | B | 0 | 113.78
2024-12-02 | B | NA | 4826.95
2024-12-02 | B | NA | 4897.61
2024-12-02 | B | NA | 4823.51
2024-12-02 | B | NA | 4924.64
2024-12-02 | B | NA | 4817.52
2024-12-02 | B | NA | 4939.05
2024-12-02 | B | NA | 4857.75
2024-12-02 | B | NA | 4925.13
2024-12-02 | B | NA | 4858.06
2024-12-02 | B | NA | 4822.24
2024-12-02 | B | NA | 4908.07
2024-12-02 | B | NA | 4958.56
2024-12-02 | B | NA | 4925.01
2024-12-02 | B | NA | 4859.53
2024-12-02 | B | NA | 4857.8
2024-12-02 | B | NA | 4828.44
2024-12-02 | B | NA | 4938.69
2024-12-02 | B | NA | 4916.13
2024-12-02 | B | NA | 4888.09
2024-12-02 | B | NA | 4885.1
2024-12-03 | B | 10.25 | 7425.02
2024-12-03 | B | 10.25 | 7218.92
2024-12-03 | B | 10.25 | 7305.41
2024-12-03 | B | 10.25 | 7428.88
2024-12-03 | B | 10.25 | 7345.72
2024-12-03 | B | 10.25 | 7351.73
2024-12-03 | B | 10.25 | 7347.55
2024-12-03 | B | 10.25 | 7323.58
2024-12-03 | B | 10.25 | 7307.95
2024-12-03 | B | 10.25 | 7352.98
2024-12-03 | B | 10.25 | 7311.79
2024-12-03 | B | 10.25 | 7343.39
2024-12-03 | B | 10.25 | 7387.57
2024-12-03 | B | 10.25 | 7253.88
2024-12-03 | B | 10.25 | 7345.21
2024-12-03 | B | 10.25 | 7355.39
2024-12-03 | B | 10.25 | 7209.66
2024-12-03 | B | 10.25 | 7369.01
2024-12-03 | B | 10.25 | 7421.1
2024-12-03 | B | 10.25 | 7248.22
2024-12-03 | B | 8.2 | 6288.53
2024-12-03 | B | 8.2 | 6268.85
2024-12-03 | B | 8.2 | 6245.96
2024-12-03 | B | 8.2 | 6235.36
2024-12-03 | B | 8.2 | 6299.63
2024-12-03 | B | 8.2 | 6234.82
2024-12-03 | B | 8.2 | 6245.27
2024-12-03 | B | 8.2 | 6285.05
2024-12-03 | B | 8.2 | 6336.67
2024-12-03 | B | 8.2 | 6271.51
2024-12-03 | B | 8.2 | 6275.15
2024-12-03 | B | 8.2 | 6271.05
2024-12-03 | B | 8.2 | 6285.73
2024-12-03 | B | 8.2 | 6338.81
2024-12-03 | B | 8.2 | 6276.96
2024-12-03 | B | 8.2 | 6283.24
2024-12-03 | B | 8.2 | 6312.55
2024-12-03 | B | 8.2 | 6150.32
2024-12-03 | B | 8.2 | 6288.6
2024-12-03 | B | 8.2 | 6267.78
2024-12-03 | B | 6.15 | 4966.87
2024-12-03 | B | 6.15 | 4933.21
2024-12-03 | B | 6.15 | 4869.46
2024-12-03 | B | 6.15 | 4914.72
2024-12-03 | B | 6.15 | 4907.49
2024-12-03 | B | 6.15 | 4885.29
2024-12-03 | B | 6.15 | 4860.52
2024-12-03 | B | 6.15 | 4893.45
2024-12-03 | B | 6.15 | 4982.03
2024-12-03 | B | 6.15 | 4908.78
2024-12-03 | B | 6.15 | 4867.67
2024-12-03 | B | 6.15 | 4858.32
2024-12-03 | B | 6.15 | 4926.03
2024-12-03 | B | 6.15 | 4908.67
2024-12-03 | B | 6.15 | 4910.67
2024-12-03 | B | 6.15 | 4920.91
2024-12-03 | B | 6.15 | 4883.18
2024-12-03 | B | 6.15 | 4927.94
2024-12-03 | B | 6.15 | 4951.81
2024-12-03 | B | 6.15 | 4961.63
2024-12-03 | B | 4.1 | 3334.98
2024-12-03 | B | 4.1 | 3251.43
2024-12-03 | B | 4.1 | 3323.96
2024-12-03 | B | 4.1 | 3312.37
2024-12-03 | B | 4.1 | 3348.33
2024-12-03 | B | 4.1 | 3289.14
2024-12-03 | B | 4.1 | 3328.02
2024-12-03 | B | 4.1 | 3354.74
2024-12-03 | B | 4.1 | 3316.37
2024-12-03 | B | 4.1 | 3308.71
2024-12-03 | B | 4.1 | 3377.29
2024-12-03 | B | 4.1 | 3335.66
2024-12-03 | B | 4.1 | 3345.75
2024-12-03 | B | 4.1 | 3349.94
2024-12-03 | B | 4.1 | 3307.01
2024-12-03 | B | 4.1 | 3336.93
2024-12-03 | B | 4.1 | 3306.7
2024-12-03 | B | 4.1 | 3276.65
2024-12-03 | B | 4.1 | 3288.03
2024-12-03 | B | 4.1 | 3336.74
2024-12-03 | B | 2.05 | 1734.92
2024-12-03 | B | 2.05 | 1759.57
2024-12-03 | B | 2.05 | 1759.79
2024-12-03 | B | 2.05 | 1745.17
2024-12-03 | B | 2.05 | 1743.7
2024-12-03 | B | 2.05 | 1725.44
2024-12-03 | B | 2.05 | 1723.67
2024-12-03 | B | 2.05 | 1739.68
2024-12-03 | B | 2.05 | 1722.2
2024-12-03 | B | 2.05 | 1729.86
2024-12-03 | B | 2.05 | 1724.69
2024-12-03 | B | 2.05 | 1766.88
2024-12-03 | B | 2.05 | 1793.53
2024-12-03 | B | 2.05 | 1701.21
2024-12-03 | B | 2.05 | 1709.18
2024-12-03 | B | 2.05 | 1708.07
2024-12-03 | B | 2.05 | 1754.8
2024-12-03 | B | 2.05 | 1718.33
2024-12-03 | B | 2.05 | 1751.94
2024-12-03 | B | 2.05 | 1768.33
2024-12-03 | B | 0 | 120.44
2024-12-03 | B | 0 | 116.68
2024-12-03 | B | 0 | 121.2
2024-12-03 | B | 0 | 123.14
2024-12-03 | B | 0 | 125.94
2024-12-03 | B | 0 | 106.06
2024-12-03 | B | 0 | 126.84
2024-12-03 | B | 0 | 124.71
2024-12-03 | B | 0 | 120.96
2024-12-03 | B | 0 | 119.98
2024-12-03 | B | 0 | 114.91
2024-12-03 | B | 0 | 110.82
2024-12-03 | B | 0 | 121.44
2024-12-03 | B | 0 | 124.67
2024-12-03 | B | 0 | 115.39
2024-12-03 | B | 0 | 113.66
2024-12-03 | B | 0 | 111.11
2024-12-03 | B | 0 | 122.88
2024-12-03 | B | 0 | 120.07
2024-12-03 | B | 0 | 119.5
2024-12-03 | B | NA | 4910.02
2024-12-03 | B | NA | 4911.46
2024-12-03 | B | NA | 4856.05
2024-12-03 | B | NA | 4886.48
2024-12-03 | B | NA | 4896.23
2024-12-03 | B | NA | 4897.79
2024-12-03 | B | NA | 4883.39
2024-12-03 | B | NA | 4886.72
2024-12-03 | B | NA | 4928.14
2024-12-03 | B | NA | 4823.87
2024-12-03 | B | NA | 4848.31
2024-12-03 | B | NA | 4855.79
2024-12-03 | B | NA | 4851.41
2024-12-03 | B | NA | 4818.52
2024-12-03 | B | NA | 4870.72
2024-12-03 | B | NA | 4859.03
2024-12-03 | B | NA | 4859.7
2024-12-03 | B | NA | 4912.18
2024-12-03 | B | NA | 4911.75
2024-12-03 | B | NA | 4919.92
2024-12-04 | B | 10.25 | 7267.46
2024-12-04 | B | 10.25 | 7382.62
2024-12-04 | B | 10.25 | 7308.54
2024-12-04 | B | 10.25 | 7309.87
2024-12-04 | B | 10.25 | 7390.97
2024-12-04 | B | 10.25 | 7272.98
2024-12-04 | B | 10.25 | 7280.12
2024-12-04 | B | 10.25 | 7337.29
2024-12-04 | B | 10.25 | 7312.84
2024-12-04 | B | 10.25 | 7263.19
2024-12-04 | B | 10.25 | 7320.85
2024-12-04 | B | 10.25 | 7293.52
2024-12-04 | B | 10.25 | 7319.2
2024-12-04 | B | 10.25 | 7349.03
2024-12-04 | B | 10.25 | 7285
2024-12-04 | B | 10.25 | 7195.01
2024-12-04 | B | 10.25 | 7347.9
2024-12-04 | B | 10.25 | 7380.93
2024-12-04 | B | 10.25 | 7396.17
2024-12-04 | B | 10.25 | 7342.61
2024-12-04 | B | 8.2 | 6274.21
2024-12-04 | B | 8.2 | 6307.83
2024-12-04 | B | 8.2 | 6296.8
2024-12-04 | B | 8.2 | 6307.53
2024-12-04 | B | 8.2 | 6330.53
2024-12-04 | B | 8.2 | 6342.62
2024-12-04 | B | 8.2 | 6339.11
2024-12-04 | B | 8.2 | 6290.29
2024-12-04 | B | 8.2 | 6182.15
2024-12-04 | B | 8.2 | 6339.79
2024-12-04 | B | 8.2 | 6291.7
2024-12-04 | B | 8.2 | 6264.88
2024-12-04 | B | 8.2 | 6299.53
2024-12-04 | B | 8.2 | 6303.54
2024-12-04 | B | 8.2 | 6281.94
2024-12-04 | B | 8.2 | 6307.65
2024-12-04 | B | 8.2 | 6413.85
2024-12-04 | B | 8.2 | 6321.79
2024-12-04 | B | 8.2 | 6272.75
2024-12-04 | B | 8.2 | 6311.85
2024-12-04 | B | 6.15 | 4892.91
2024-12-04 | B | 6.15 | 4866.52
2024-12-04 | B | 6.15 | 4888.77
2024-12-04 | B | 6.15 | 4924.59
2024-12-04 | B | 6.15 | 4871.21
2024-12-04 | B | 6.15 | 4943.31
2024-12-04 | B | 6.15 | 4973.14
2024-12-04 | B | 6.15 | 4954.31
2024-12-04 | B | 6.15 | 4866.66
2024-12-04 | B | 6.15 | 4880.24
2024-12-04 | B | 6.15 | 4918.02
2024-12-04 | B | 6.15 | 4872.52
2024-12-04 | B | 6.15 | 4913.23
2024-12-04 | B | 6.15 | 4877.6
2024-12-04 | B | 6.15 | 4931.97
2024-12-04 | B | 6.15 | 4941.77
2024-12-04 | B | 6.15 | 4960.37
2024-12-04 | B | 6.15 | 4911.91
2024-12-04 | B | 6.15 | 4878.89
2024-12-04 | B | 6.15 | 4949.43
2024-12-04 | B | 4.1 | 3315.31
2024-12-04 | B | 4.1 | 3286.2
2024-12-04 | B | 4.1 | 3354.83
2024-12-04 | B | 4.1 | 3267.19
2024-12-04 | B | 4.1 | 3330.28
2024-12-04 | B | 4.1 | 3383.78
2024-12-04 | B | 4.1 | 3329.17
2024-12-04 | B | 4.1 | 3388.01
2024-12-04 | B | 4.1 | 3264.24
2024-12-04 | B | 4.1 | 3275.68
2024-12-04 | B | 4.1 | 3284.56
2024-12-04 | B | 4.1 | 3327.45
2024-12-04 | B | 4.1 | 3313.89
2024-12-04 | B | 4.1 | 3361.17
2024-12-04 | B | 4.1 | 3308.3
2024-12-04 | B | 4.1 | 3321.71
2024-12-04 | B | 4.1 | 3326.65
2024-12-04 | B | 4.1 | 3304.02
2024-12-04 | B | 4.1 | 3320
2024-12-04 | B | 4.1 | 3295.18
2024-12-04 | B | 2.05 | 1781.92
2024-12-04 | B | 2.05 | 1780.27
2024-12-04 | B | 2.05 | 1744.78
2024-12-04 | B | 2.05 | 1712.13
2024-12-04 | B | 2.05 | 1786.64
2024-12-04 | B | 2.05 | 1699.83
2024-12-04 | B | 2.05 | 1756.83
2024-12-04 | B | 2.05 | 1751.67
2024-12-04 | B | 2.05 | 1768.66
2024-12-04 | B | 2.05 | 1784.72
2024-12-04 | B | 2.05 | 1748.7
2024-12-04 | B | 2.05 | 1738.17
2024-12-04 | B | 2.05 | 1804.17
2024-12-04 | B | 2.05 | 1768.03
2024-12-04 | B | 2.05 | 1697.23
2024-12-04 | B | 2.05 | 1725.58
2024-12-04 | B | 2.05 | 1728.03
2024-12-04 | B | 2.05 | 1748.47
2024-12-04 | B | 2.05 | 1731.38
2024-12-04 | B | 2.05 | 1698.87
2024-12-04 | B | 0 | 120.36
2024-12-04 | B | 0 | 106.96
2024-12-04 | B | 0 | 128.36
2024-12-04 | B | 0 | 122.3
2024-12-04 | B | 0 | 117.72
2024-12-04 | B | 0 | 123.69
2024-12-04 | B | 0 | 118.65
2024-12-04 | B | 0 | 124.56
2024-12-04 | B | 0 | 128.67
2024-12-04 | B | 0 | 116.21
2024-12-04 | B | 0 | 117.28
2024-12-04 | B | 0 | 114.13
2024-12-04 | B | 0 | 120.81
2024-12-04 | B | 0 | 121.92
2024-12-04 | B | 0 | 125.55
2024-12-04 | B | 0 | 115.14
2024-12-04 | B | 0 | 122.3
2024-12-04 | B | 0 | 125.46
2024-12-04 | B | 0 | 115.85
2024-12-04 | B | 0 | 112.67
2024-12-04 | B | NA | 4884.95
2024-12-04 | B | NA | 4887.36
2024-12-04 | B | NA | 4991.12
2024-12-04 | B | NA | 4883.41
2024-12-04 | B | NA | 4940.43
2024-12-04 | B | NA | 4846.49
2024-12-04 | B | NA | 4903.99
2024-12-04 | B | NA | 4924.48
2024-12-04 | B | NA | 5009.77
2024-12-04 | B | NA | 4850.93
2024-12-04 | B | NA | 4844.01
2024-12-04 | B | NA | 4954.82
2024-12-04 | B | NA | 4931.58
2024-12-04 | B | NA | 4894.87
2024-12-04 | B | NA | 4845.22
2024-12-04 | B | NA | 4919.68
2024-12-04 | B | NA | 4972.49
2024-12-04 | B | NA | 4906.55
2024-12-04 | B | NA | 4901.14
2024-12-04 | B | NA | 4841.37
2024-12-05 | B | 10.25 | 7357.19
2024-12-05 | B | 10.25 | 7322.38
2024-12-05 | B | 10.25 | 7315.72
2024-12-05 | B | 10.25 | 7322.59
2024-12-05 | B | 10.25 | 7273.92
2024-12-05 | B | 10.25 | 7227.35
2024-12-05 | B | 10.25 | 7354.69
2024-12-05 | B | 10.25 | 7311.99
2024-12-05 | B | 10.25 | 7319.87
2024-12-05 | B | 10.25 | 7288.13
2024-12-05 | B | 10.25 | 7318.34
2024-12-05 | B | 10.25 | 7431.29
2024-12-05 | B | 10.25 | 7392.6
2024-12-05 | B | 10.25 | 7378.54
2024-12-05 | B | 10.25 | 7345.92
2024-12-05 | B | 10.25 | 7309.47
2024-12-05 | B | 10.25 | 7292.5
2024-12-05 | B | 10.25 | 7324.26
2024-12-05 | B | 10.25 | 7190.89
2024-12-05 | B | 10.25 | 7278.62
2024-12-05 | B | 8.2 | 6339.11
2024-12-05 | B | 8.2 | 6255.58
2024-12-05 | B | 8.2 | 6346.69
2024-12-05 | B | 8.2 | 6301.49
2024-12-05 | B | 8.2 | 6295.73
2024-12-05 | B | 8.2 | 6185.89
2024-12-05 | B | 8.2 | 6247.49
2024-12-05 | B | 8.2 | 6281.41
2024-12-05 | B | 8.2 | 6310.05
2024-12-05 | B | 8.2 | 6284.47
2024-12-05 | B | 8.2 | 6222.76
2024-12-05 | B | 8.2 | 6169.56
2024-12-05 | B | 8.2 | 6267.8
2024-12-05 | B | 8.2 | 6270.41
2024-12-05 | B | 8.2 | 6283.31
2024-12-05 | B | 8.2 | 6280.23
2024-12-05 | B | 8.2 | 6321.72
2024-12-05 | B | 8.2 | 6295.1
2024-12-05 | B | 8.2 | 6276.87
2024-12-05 | B | 8.2 | 6264.2
2024-12-05 | B | 6.15 | 4869.95
2024-12-05 | B | 6.15 | 4944.95
2024-12-05 | B | 6.15 | 4884.62
2024-12-05 | B | 6.15 | 4863.77
2024-12-05 | B | 6.15 | 4891.96
2024-12-05 | B | 6.15 | 4888.15
2024-12-05 | B | 6.15 | 4913.71
2024-12-05 | B | 6.15 | 4935.41
2024-12-05 | B | 6.15 | 4909.11
2024-12-05 | B | 6.15 | 4893.35
2024-12-05 | B | 6.15 | 4884.29
2024-12-05 | B | 6.15 | 4940.47
2024-12-05 | B | 6.15 | 4950.43
2024-12-05 | B | 6.15 | 4925.47
2024-12-05 | B | 6.15 | 4923.19
2024-12-05 | B | 6.15 | 4941.91
2024-12-05 | B | 6.15 | 4939.64
2024-12-05 | B | 6.15 | 4916.74
2024-12-05 | B | 6.15 | 4893.94
2024-12-05 | B | 6.15 | 4876.17
2024-12-05 | B | 4.1 | 3282.78
2024-12-05 | B | 4.1 | 3301.47
2024-12-05 | B | 4.1 | 3283.01
2024-12-05 | B | 4.1 | 3303.02
2024-12-05 | B | 4.1 | 3300.56
2024-12-05 | B | 4.1 | 3340.9
2024-12-05 | B | 4.1 | 3330.07
2024-12-05 | B | 4.1 | 3350.88
2024-12-05 | B | 4.1 | 3287.36
2024-12-05 | B | 4.1 | 3275.47
2024-12-05 | B | 4.1 | 3304.67
2024-12-05 | B | 4.1 | 3366.58
2024-12-05 | B | 4.1 | 3309.16
2024-12-05 | B | 4.1 | 3346.97
2024-12-05 | B | 4.1 | 3294.22
2024-12-05 | B | 4.1 | 3317.15
2024-12-05 | B | 4.1 | 3325.8
2024-12-05 | B | 4.1 | 3324.14
2024-12-05 | B | 4.1 | 3393.56
2024-12-05 | B | 4.1 | 3296.23
2024-12-05 | B | 2.05 | 1763.3
2024-12-05 | B | 2.05 | 1750.53
2024-12-05 | B | 2.05 | 1716.85
2024-12-05 | B | 2.05 | 1759.43
2024-12-05 | B | 2.05 | 1769.59
2024-12-05 | B | 2.05 | 1753.64
2024-12-05 | B | 2.05 | 1783.86
2024-12-05 | B | 2.05 | 1842.73
2024-12-05 | B | 2.05 | 1752.74
2024-12-05 | B | 2.05 | 1775.37
2024-12-05 | B | 2.05 | 1741.07
2024-12-05 | B | 2.05 | 1752.86
2024-12-05 | B | 2.05 | 1773.68
2024-12-05 | B | 2.05 | 1760.28
2024-12-05 | B | 2.05 | 1759.78
2024-12-05 | B | 2.05 | 1738.44
2024-12-05 | B | 2.05 | 1778.11
2024-12-05 | B | 2.05 | 1760.1
2024-12-05 | B | 2.05 | 1725.4
2024-12-05 | B | 2.05 | 1741.37
2024-12-05 | B | 0 | 110.7
2024-12-05 | B | 0 | 115.1
2024-12-05 | B | 0 | 113.5
2024-12-05 | B | 0 | 122.33
2024-12-05 | B | 0 | 116.16
2024-12-05 | B | 0 | 118.79
2024-12-05 | B | 0 | 115.62
2024-12-05 | B | 0 | 123.11
2024-12-05 | B | 0 | 128.52
2024-12-05 | B | 0 | 123.34
2024-12-05 | B | 0 | 112.89
2024-12-05 | B | 0 | 115.76
2024-12-05 | B | 0 | 116.98
2024-12-05 | B | 0 | 111.3
2024-12-05 | B | 0 | 113.26
2024-12-05 | B | 0 | 118.8
2024-12-05 | B | 0 | 120.26
2024-12-05 | B | 0 | 116.14
2024-12-05 | B | 0 | 115.31
2024-12-05 | B | 0 | 116.55
2024-12-05 | B | NA | 4881.56
2024-12-05 | B | NA | 4934.65
2024-12-05 | B | NA | 4843.49
2024-12-05 | B | NA | 4950.57
2024-12-05 | B | NA | 4802.44
2024-12-05 | B | NA | 4841.98
2024-12-05 | B | NA | 4909.6
2024-12-05 | B | NA | 4774.94
2024-12-05 | B | NA | 4851.83
2024-12-05 | B | NA | 4801.52
2024-12-05 | B | NA | 4867.45
2024-12-05 | B | NA | 4892
2024-12-05 | B | NA | 4905.83
2024-12-05 | B | NA | 4870.33
2024-12-05 | B | NA | 4836.82
2024-12-05 | B | NA | 4892.7
2024-12-05 | B | NA | 4829.37
2024-12-05 | B | NA | 4893.98
2024-12-05 | B | NA | 4819.2
2024-12-05 | B | NA | 4869.74
2024-12-06 | C | 10.25 | 7354.22
2024-12-06 | C | 10.25 | 7373.47
2024-12-06 | C | 10.25 | 7375.29
2024-12-06 | C | 10.25 | 7328.68
2024-12-06 | C | 10.25 | 7345.66
2024-12-06 | C | 10.25 | 7301.03
2024-12-06 | C | 10.25 | 7298.17
2024-12-06 | C | 10.25 | 7342.94
2024-12-06 | C | 10.25 | 7341.59
2024-12-06 | C | 10.25 | 7307.72
2024-12-06 | C | 10.25 | 7364.31
2024-12-06 | C | 10.25 | 7349.15
2024-12-06 | C | 10.25 | 7313.68
2024-12-06 | C | 10.25 | 7347.44
2024-12-06 | C | 10.25 | 7376.58
2024-12-06 | C | 10.25 | 7341.7
2024-12-06 | C | 10.25 | 7268.64
2024-12-06 | C | 10.25 | 7355.51
2024-12-06 | C | 10.25 | 7371.11
2024-12-06 | C | 10.25 | 7436.03
2024-12-06 | C | 8.2 | 6194.79
2024-12-06 | C | 8.2 | 6341.36
2024-12-06 | C | 8.2 | 6263.27
2024-12-06 | C | 8.2 | 6185.02
2024-12-06 | C | 8.2 | 6245
2024-12-06 | C | 8.2 | 6277.91
2024-12-06 | C | 8.2 | 6292.96
2024-12-06 | C | 8.2 | 6277.6
2024-12-06 | C | 8.2 | 6266.56
2024-12-06 | C | 8.2 | 6238.48
2024-12-06 | C | 8.2 | 6337.35
2024-12-06 | C | 8.2 | 6343.12
2024-12-06 | C | 8.2 | 6299.37
2024-12-06 | C | 8.2 | 6273.87
2024-12-06 | C | 8.2 | 6326.63
2024-12-06 | C | 8.2 | 6291.14
2024-12-06 | C | 8.2 | 6222.35
2024-12-06 | C | 8.2 | 6323.2
2024-12-06 | C | 8.2 | 6358.47
2024-12-06 | C | 8.2 | 6303.22
2024-12-06 | C | 6.15 | 4960.88
2024-12-06 | C | 6.15 | 4897.95
2024-12-06 | C | 6.15 | 4884.55
2024-12-06 | C | 6.15 | 4961.09
2024-12-06 | C | 6.15 | 4830
2024-12-06 | C | 6.15 | 4932.56
2024-12-06 | C | 6.15 | 4884.45
2024-12-06 | C | 6.15 | 4829.7
2024-12-06 | C | 6.15 | 4865.66
2024-12-06 | C | 6.15 | 4926.7
2024-12-06 | C | 6.15 | 4821.69
2024-12-06 | C | 6.15 | 4887.23
2024-12-06 | C | 6.15 | 4911.19
2024-12-06 | C | 6.15 | 4931.84
2024-12-06 | C | 6.15 | 4922.52
2024-12-06 | C | 6.15 | 4938.61
2024-12-06 | C | 6.15 | 4885.33
2024-12-06 | C | 6.15 | 4842.88
2024-12-06 | C | 6.15 | 4871.63
2024-12-06 | C | 6.15 | 5016.5
2024-12-06 | C | 4.1 | 3318.53
2024-12-06 | C | 4.1 | 3380.46
2024-12-06 | C | 4.1 | 3366.29
2024-12-06 | C | 4.1 | 3312.78
2024-12-06 | C | 4.1 | 3269.85
2024-12-06 | C | 4.1 | 3350.33
2024-12-06 | C | 4.1 | 3311.88
2024-12-06 | C | 4.1 | 3320.56
2024-12-06 | C | 4.1 | 3388.9
2024-12-06 | C | 4.1 | 3328.5
2024-12-06 | C | 4.1 | 3292.05
2024-12-06 | C | 4.1 | 3297.3
2024-12-06 | C | 4.1 | 3350.39
2024-12-06 | C | 4.1 | 3341.94
2024-12-06 | C | 4.1 | 3276.05
2024-12-06 | C | 4.1 | 3357.68
2024-12-06 | C | 4.1 | 3356.45
2024-12-06 | C | 4.1 | 3347.31
2024-12-06 | C | 4.1 | 3340.72
2024-12-06 | C | 4.1 | 3285.71
2024-12-06 | C | 2.05 | 1756.82
2024-12-06 | C | 2.05 | 1724.12
2024-12-06 | C | 2.05 | 1797.2
2024-12-06 | C | 2.05 | 1744.68
2024-12-06 | C | 2.05 | 1705.05
2024-12-06 | C | 2.05 | 1767.04
2024-12-06 | C | 2.05 | 1765.27
2024-12-06 | C | 2.05 | 1743.47
2024-12-06 | C | 2.05 | 1727.01
2024-12-06 | C | 2.05 | 1758.11
2024-12-06 | C | 2.05 | 1737.18
2024-12-06 | C | 2.05 | 1785.32
2024-12-06 | C | 2.05 | 1722.47
2024-12-06 | C | 2.05 | 1775.14
2024-12-06 | C | 2.05 | 1772.13
2024-12-06 | C | 2.05 | 1721.19
2024-12-06 | C | 2.05 | 1743.89
2024-12-06 | C | 2.05 | 1732.84
2024-12-06 | C | 2.05 | 1729.98
2024-12-06 | C | 2.05 | 1704.15
2024-12-06 | C | 0 | 116.01
2024-12-06 | C | 0 | 121.82
2024-12-06 | C | 0 | 117.43
2024-12-06 | C | 0 | 117.92
2024-12-06 | C | 0 | 111.51
2024-12-06 | C | 0 | 111.91
2024-12-06 | C | 0 | 121.36
2024-12-06 | C | 0 | 120.11
2024-12-06 | C | 0 | 113.38
2024-12-06 | C | 0 | 121.62
2024-12-06 | C | 0 | 107.08
2024-12-06 | C | 0 | 116.62
2024-12-06 | C | 0 | 104.37
2024-12-06 | C | 0 | 122.79
2024-12-06 | C | 0 | 120.94
2024-12-06 | C | 0 | 118.61
2024-12-06 | C | 0 | 127.59
2024-12-06 | C | 0 | 123.17
2024-12-06 | C | 0 | 119.87
2024-12-06 | C | 0 | 124.81
2024-12-06 | C | NA | 4882.07
2024-12-06 | C | NA | 4910.58
2024-12-06 | C | NA | 4897.19
2024-12-06 | C | NA | 4832.43
2024-12-06 | C | NA | 4869.88
2024-12-06 | C | NA | 4853.04
2024-12-06 | C | NA | 4890.46
2024-12-06 | C | NA | 4977.2
2024-12-06 | C | NA | 4827.52
2024-12-06 | C | NA | 4948.35
2024-12-06 | C | NA | 4912.28
2024-12-06 | C | NA | 4807.46
2024-12-06 | C | NA | 4839.12
2024-12-06 | C | NA | 4883.19
2024-12-06 | C | NA | 4875.39
2024-12-06 | C | NA | 4883.19
2024-12-06 | C | NA | 4857.99
2024-12-06 | C | NA | 4769.95
2024-12-06 | C | NA | 4908.52
2024-12-06 | C | NA | 4854.07
2024-12-09 | C | 10.25 | 7398.91
2024-12-09 | C | 10.25 | 7301.53
2024-12-09 | C | 10.25 | 7378.62
2024-12-09 | C | 10.25 | 7327.16
2024-12-09 | C | 10.25 | 7420.69
2024-12-09 | C | 10.25 | 7252.76
2024-12-09 | C | 10.25 | 7380.45
2024-12-09 | C | 10.25 | 7323.68
2024-12-09 | C | 10.25 | 7348.58
2024-12-09 | C | 10.25 | 7345.9
2024-12-09 | C | 10.25 | 7300.95
2024-12-09 | C | 10.25 | 7217.82
2024-12-09 | C | 10.25 | 7399.02
2024-12-09 | C | 10.25 | 7330.57
2024-12-09 | C | 10.25 | 7411.6
2024-12-09 | C | 10.25 | 7256.16
2024-12-09 | C | 10.25 | 7301.27
2024-12-09 | C | 10.25 | 7335.5
2024-12-09 | C | 10.25 | 7308.03
2024-12-09 | C | 10.25 | 7320.74
2024-12-09 | C | 8.2 | 6361.75
2024-12-09 | C | 8.2 | 6340.33
2024-12-09 | C | 8.2 | 6299.88
2024-12-09 | C | 8.2 | 6299.5
2024-12-09 | C | 8.2 | 6269.71
2024-12-09 | C | 8.2 | 6242.37
2024-12-09 | C | 8.2 | 6272.85
2024-12-09 | C | 8.2 | 6328.02
2024-12-09 | C | 8.2 | 6292.47
2024-12-09 | C | 8.2 | 6201.04
2024-12-09 | C | 8.2 | 6247.56
2024-12-09 | C | 8.2 | 6286.52
2024-12-09 | C | 8.2 | 6285
2024-12-09 | C | 8.2 | 6302.76
2024-12-09 | C | 8.2 | 6221.92
2024-12-09 | C | 8.2 | 6257.99
2024-12-09 | C | 8.2 | 6245.83
2024-12-09 | C | 8.2 | 6250.98
2024-12-09 | C | 8.2 | 6212.83
2024-12-09 | C | 8.2 | 6182.83
2024-12-09 | C | 6.15 | 4827.94
2024-12-09 | C | 6.15 | 4959.3
2024-12-09 | C | 6.15 | 4886.78
2024-12-09 | C | 6.15 | 4894.09
2024-12-09 | C | 6.15 | 4888.04
2024-12-09 | C | 6.15 | 4935.97
2024-12-09 | C | 6.15 | 4884.61
2024-12-09 | C | 6.15 | 4972.69
2024-12-09 | C | 6.15 | 4945.65
2024-12-09 | C | 6.15 | 4895.96
2024-12-09 | C | 6.15 | 4873.26
2024-12-09 | C | 6.15 | 4981.63
2024-12-09 | C | 6.15 | 4861.09
2024-12-09 | C | 6.15 | 4870.86
2024-12-09 | C | 6.15 | 4954.98
2024-12-09 | C | 6.15 | 4939
2024-12-09 | C | 6.15 | 4875.36
2024-12-09 | C | 6.15 | 4951.81
2024-12-09 | C | 6.15 | 4968.83
2024-12-09 | C | 6.15 | 4959.1
2024-12-09 | C | 4.1 | 3337.51
2024-12-09 | C | 4.1 | 3326.63
2024-12-09 | C | 4.1 | 3267.49
2024-12-09 | C | 4.1 | 3364.08
2024-12-09 | C | 4.1 | 3311.64
2024-12-09 | C | 4.1 | 3323.44
2024-12-09 | C | 4.1 | 3360.02
2024-12-09 | C | 4.1 | 3343.36
2024-12-09 | C | 4.1 | 3251.13
2024-12-09 | C | 4.1 | 3270.45
2024-12-09 | C | 4.1 | 3364.55
2024-12-09 | C | 4.1 | 3269.82
2024-12-09 | C | 4.1 | 3321.48
2024-12-09 | C | 4.1 | 3298.95
2024-12-09 | C | 4.1 | 3284.47
2024-12-09 | C | 4.1 | 3398.23
2024-12-09 | C | 4.1 | 3335.71
2024-12-09 | C | 4.1 | 3308.56
2024-12-09 | C | 4.1 | 3299.6
2024-12-09 | C | 4.1 | 3301.9
2024-12-09 | C | 2.05 | 1707.44
2024-12-09 | C | 2.05 | 1718.67
2024-12-09 | C | 2.05 | 1752.64
2024-12-09 | C | 2.05 | 1706.57
2024-12-09 | C | 2.05 | 1743.92
2024-12-09 | C | 2.05 | 1782.26
2024-12-09 | C | 2.05 | 1733.69
2024-12-09 | C | 2.05 | 1743.56
2024-12-09 | C | 2.05 | 1733.85
2024-12-09 | C | 2.05 | 1742.64
2024-12-09 | C | 2.05 | 1719.95
2024-12-09 | C | 2.05 | 1735
2024-12-09 | C | 2.05 | 1719.51
2024-12-09 | C | 2.05 | 1713.76
2024-12-09 | C | 2.05 | 1724.13
2024-12-09 | C | 2.05 | 1694.19
2024-12-09 | C | 2.05 | 1739.21
2024-12-09 | C | 2.05 | 1750.86
2024-12-09 | C | 2.05 | 1742.2
2024-12-09 | C | 2.05 | 1749
2024-12-09 | C | 0 | 112.53
2024-12-09 | C | 0 | 121.13
2024-12-09 | C | 0 | 120.56
2024-12-09 | C | 0 | 113.89
2024-12-09 | C | 0 | 123.75
2024-12-09 | C | 0 | 114.32
2024-12-09 | C | 0 | 113.6
2024-12-09 | C | 0 | 119.76
2024-12-09 | C | 0 | 121.64
2024-12-09 | C | 0 | 116.17
2024-12-09 | C | 0 | 119.49
2024-12-09 | C | 0 | 111.65
2024-12-09 | C | 0 | 122.48
2024-12-09 | C | 0 | 120.57
2024-12-09 | C | 0 | 114.78
2024-12-09 | C | 0 | 120.48
2024-12-09 | C | 0 | 118.67
2024-12-09 | C | 0 | 125.12
2024-12-09 | C | 0 | 118.97
2024-12-09 | C | 0 | 112.66
2024-12-09 | C | NA | 4803.59
2024-12-09 | C | NA | 4924.56
2024-12-09 | C | NA | 4837.02
2024-12-09 | C | NA | 4829.65
2024-12-09 | C | NA | 4932.54
2024-12-09 | C | NA | 4940.35
2024-12-09 | C | NA | 4835.95
2024-12-09 | C | NA | 4882.83
2024-12-09 | C | NA | 4907.76
2024-12-09 | C | NA | 4966.73
2024-12-09 | C | NA | 4909.24
2024-12-09 | C | NA | 4871.58
2024-12-09 | C | NA | 4953.39
2024-12-09 | C | NA | 4927.27
2024-12-09 | C | NA | 4855.69
2024-12-09 | C | NA | 4879.05
2024-12-09 | C | NA | 4842.27
2024-12-09 | C | NA | 4883.39
2024-12-09 | C | NA | 4875.88
2024-12-09 | C | NA | 4860.93
2024-12-10 | C | 10.25 | 7333.64
2024-12-10 | C | 10.25 | 7358.49
2024-12-10 | C | 10.25 | 7313.83
2024-12-10 | C | 10.25 | 7279.78
2024-12-10 | C | 10.25 | 7297.16
2024-12-10 | C | 10.25 | 7353.46
2024-12-10 | C | 10.25 | 7387.81
2024-12-10 | C | 10.25 | 7253.4
2024-12-10 | C | 10.25 | 7293.89
2024-12-10 | C | 10.25 | 7314.95
2024-12-10 | C | 10.25 | 7284.33
2024-12-10 | C | 10.25 | 7306.27
2024-12-10 | C | 10.25 | 7287.42
2024-12-10 | C | 10.25 | 7355.37
2024-12-10 | C | 10.25 | 7318.71
2024-12-10 | C | 10.25 | 7243.11
2024-12-10 | C | 10.25 | 7268.74
2024-12-10 | C | 10.25 | 7403.86
2024-12-10 | C | 10.25 | 7397.9
2024-12-10 | C | 10.25 | 7404.08
2024-12-10 | C | 8.2 | 6294.5
2024-12-10 | C | 8.2 | 6302.67
2024-12-10 | C | 8.2 | 6355.78
2024-12-10 | C | 8.2 | 6312.3
2024-12-10 | C | 8.2 | 6254.04
2024-12-10 | C | 8.2 | 6353.58
2024-12-10 | C | 8.2 | 6227
2024-12-10 | C | 8.2 | 6230.81
2024-12-10 | C | 8.2 | 6208.71
2024-12-10 | C | 8.2 | 6311.67
2024-12-10 | C | 8.2 | 6299.96
2024-12-10 | C | 8.2 | 6226.88
2024-12-10 | C | 8.2 | 6242.89
2024-12-10 | C | 8.2 | 6286.56
2024-12-10 | C | 8.2 | 6240.26
2024-12-10 | C | 8.2 | 6237.53
2024-12-10 | C | 8.2 | 6184.27
2024-12-10 | C | 8.2 | 6323.4
2024-12-10 | C | 8.2 | 6249.62
2024-12-10 | C | 8.2 | 6269.16
2024-12-10 | C | 6.15 | 4875.65
2024-12-10 | C | 6.15 | 4896.49
2024-12-10 | C | 6.15 | 4879.36
2024-12-10 | C | 6.15 | 5006.91
2024-12-10 | C | 6.15 | 4952.82
2024-12-10 | C | 6.15 | 4848.49
2024-12-10 | C | 6.15 | 4911.32
2024-12-10 | C | 6.15 | 4938.16
2024-12-10 | C | 6.15 | 4925.11
2024-12-10 | C | 6.15 | 4904.18
2024-12-10 | C | 6.15 | 4872.35
2024-12-10 | C | 6.15 | 4878.53
2024-12-10 | C | 6.15 | 4971.49
2024-12-10 | C | 6.15 | 4868.94
2024-12-10 | C | 6.15 | 4910.71
2024-12-10 | C | 6.15 | 4904.29
2024-12-10 | C | 6.15 | 5003.71
2024-12-10 | C | 6.15 | 4903.94
2024-12-10 | C | 6.15 | 4939.92
2024-12-10 | C | 6.15 | 4887.56
2024-12-10 | C | 4.1 | 3234.25
2024-12-10 | C | 4.1 | 3342.42
2024-12-10 | C | 4.1 | 3318.19
2024-12-10 | C | 4.1 | 3304.49
2024-12-10 | C | 4.1 | 3344.57
2024-12-10 | C | 4.1 | 3336.34
2024-12-10 | C | 4.1 | 3265.71
2024-12-10 | C | 4.1 | 3329.98
2024-12-10 | C | 4.1 | 3369.7
2024-12-10 | C | 4.1 | 3338.84
2024-12-10 | C | 4.1 | 3273.51
2024-12-10 | C | 4.1 | 3293.18
2024-12-10 | C | 4.1 | 3257.09
2024-12-10 | C | 4.1 | 3307.43
2024-12-10 | C | 4.1 | 3288.11
2024-12-10 | C | 4.1 | 3316.99
2024-12-10 | C | 4.1 | 3296.02
2024-12-10 | C | 4.1 | 3274.79
2024-12-10 | C | 4.1 | 3400.63
2024-12-10 | C | 4.1 | 3344.11
2024-12-10 | C | 2.05 | 1796.01
2024-12-10 | C | 2.05 | 1797.39
2024-12-10 | C | 2.05 | 1698.9
2024-12-10 | C | 2.05 | 1725.63
2024-12-10 | C | 2.05 | 1770.74
2024-12-10 | C | 2.05 | 1734.43
2024-12-10 | C | 2.05 | 1725.11
2024-12-10 | C | 2.05 | 1782.27
2024-12-10 | C | 2.05 | 1706.29
2024-12-10 | C | 2.05 | 1733.39
2024-12-10 | C | 2.05 | 1761.34
2024-12-10 | C | 2.05 | 1797.45
2024-12-10 | C | 2.05 | 1713.54
2024-12-10 | C | 2.05 | 1720.62
2024-12-10 | C | 2.05 | 1731.79
2024-12-10 | C | 2.05 | 1760.56
2024-12-10 | C | 2.05 | 1741.93
2024-12-10 | C | 2.05 | 1761.94
2024-12-10 | C | 2.05 | 1704.11
2024-12-10 | C | 2.05 | 1761.37
2024-12-10 | C | 0 | 132.41
2024-12-10 | C | 0 | 114.34
2024-12-10 | C | 0 | 110.83
2024-12-10 | C | 0 | 120.81
2024-12-10 | C | 0 | 126.85
2024-12-10 | C | 0 | 116.69
2024-12-10 | C | 0 | 121.12
2024-12-10 | C | 0 | 110.5
2024-12-10 | C | 0 | 120.13
2024-12-10 | C | 0 | 124.09
2024-12-10 | C | 0 | 125.21
2024-12-10 | C | 0 | 123.13
2024-12-10 | C | 0 | 116.36
2024-12-10 | C | 0 | 123.01
2024-12-10 | C | 0 | 115.03
2024-12-10 | C | 0 | 125.46
2024-12-10 | C | 0 | 116.2
2024-12-10 | C | 0 | 119.99
2024-12-10 | C | 0 | 117.94
2024-12-10 | C | 0 | 116.56
2024-12-10 | C | NA | 4867.44
2024-12-10 | C | NA | 4962.7
2024-12-10 | C | NA | 4858.42
2024-12-10 | C | NA | 4876.75
2024-12-10 | C | NA | 4984.52
2024-12-10 | C | NA | 4913.91
2024-12-10 | C | NA | 4927.15
2024-12-10 | C | NA | 4818.34
2024-12-10 | C | NA | 4857.69
2024-12-10 | C | NA | 4821.69
2024-12-10 | C | NA | 4846.83
2024-12-10 | C | NA | 4989.46
2024-12-10 | C | NA | 4863.23
2024-12-10 | C | NA | 4843.14
2024-12-10 | C | NA | 4900.63
2024-12-10 | C | NA | 5017.29
2024-12-10 | C | NA | 4858.01
2024-12-10 | C | NA | 4866.67
2024-12-10 | C | NA | 4916.8
2024-12-10 | C | NA | 4928.3
2024-12-11 | C | 10.25 | 7251.59
2024-12-11 | C | 10.25 | 7348.24
2024-12-11 | C | 10.25 | 7317.27
2024-12-11 | C | 10.25 | 7306.66
2024-12-11 | C | 10.25 | 7345.04
2024-12-11 | C | 10.25 | 7307.33
2024-12-11 | C | 10.25 | 7370.53
2024-12-11 | C | 10.25 | 7352.03
2024-12-11 | C | 10.25 | 7430.12
2024-12-11 | C | 10.25 | 7285.24
2024-12-11 | C | 10.25 | 7363.99
2024-12-11 | C | 10.25 | 7309.42
2024-12-11 | C | 10.25 | 7448.62
2024-12-11 | C | 10.25 | 7296.34
2024-12-11 | C | 10.25 | 7313.15
2024-12-11 | C | 10.25 | 7442.81
2024-12-11 | C | 10.25 | 7325.09
2024-12-11 | C | 10.25 | 7430.65
2024-12-11 | C | 10.25 | 7332.88
2024-12-11 | C | 10.25 | 7326.48
2024-12-11 | C | 8.2 | 6328.48
2024-12-11 | C | 8.2 | 6260.58
2024-12-11 | C | 8.2 | 6232.03
2024-12-11 | C | 8.2 | 6339.18
2024-12-11 | C | 8.2 | 6209.63
2024-12-11 | C | 8.2 | 6233.88
2024-12-11 | C | 8.2 | 6249.74
2024-12-11 | C | 8.2 | 6375.91
2024-12-11 | C | 8.2 | 6261.24
2024-12-11 | C | 8.2 | 6189.71
2024-12-11 | C | 8.2 | 6279.42
2024-12-11 | C | 8.2 | 6302.41
2024-12-11 | C | 8.2 | 6281.11
2024-12-11 | C | 8.2 | 6312.95
2024-12-11 | C | 8.2 | 6256.87
2024-12-11 | C | 8.2 | 6310.23
2024-12-11 | C | 8.2 | 6275.04
2024-12-11 | C | 8.2 | 6271.26
2024-12-11 | C | 8.2 | 6384.88
2024-12-11 | C | 8.2 | 6337.38
2024-12-11 | C | 6.15 | 4967.39
2024-12-11 | C | 6.15 | 4900.28
2024-12-11 | C | 6.15 | 4923.77
2024-12-11 | C | 6.15 | 4914.52
2024-12-11 | C | 6.15 | 4860.66
2024-12-11 | C | 6.15 | 4862.43
2024-12-11 | C | 6.15 | 4917.22
2024-12-11 | C | 6.15 | 4907.39
2024-12-11 | C | 6.15 | 4870.05
2024-12-11 | C | 6.15 | 4882.05
2024-12-11 | C | 6.15 | 4923.09
2024-12-11 | C | 6.15 | 4874.03
2024-12-11 | C | 6.15 | 4936.82
2024-12-11 | C | 6.15 | 4942.84
2024-12-11 | C | 6.15 | 4915.98
2024-12-11 | C | 6.15 | 4958.99
2024-12-11 | C | 6.15 | 4926.46
2024-12-11 | C | 6.15 | 4866.76
2024-12-11 | C | 6.15 | 4848
2024-12-11 | C | 6.15 | 4915.05
2024-12-11 | C | 4.1 | 3338.44
2024-12-11 | C | 4.1 | 3315.64
2024-12-11 | C | 4.1 | 3376.54
2024-12-11 | C | 4.1 | 3359
2024-12-11 | C | 4.1 | 3327.03
2024-12-11 | C | 4.1 | 3326.54
2024-12-11 | C | 4.1 | 3327.42
2024-12-11 | C | 4.1 | 3309
2024-12-11 | C | 4.1 | 3417.91
2024-12-11 | C | 4.1 | 3366.27
2024-12-11 | C | 4.1 | 3329.28
2024-12-11 | C | 4.1 | 3374.16
2024-12-11 | C | 4.1 | 3286.14
2024-12-11 | C | 4.1 | 3401.24
2024-12-11 | C | 4.1 | 3295.52
2024-12-11 | C | 4.1 | 3371.55
2024-12-11 | C | 4.1 | 3323.66
2024-12-11 | C | 4.1 | 3335.55
2024-12-11 | C | 4.1 | 3321.4
2024-12-11 | C | 4.1 | 3355.9
2024-12-11 | C | 2.05 | 1700.89
2024-12-11 | C | 2.05 | 1719.44
2024-12-11 | C | 2.05 | 1784.71
2024-12-11 | C | 2.05 | 1811.56
2024-12-11 | C | 2.05 | 1768.15
2024-12-11 | C | 2.05 | 1809.77
2024-12-11 | C | 2.05 | 1737.52
2024-12-11 | C | 2.05 | 1731.97
2024-12-11 | C | 2.05 | 1783.45
2024-12-11 | C | 2.05 | 1739.89
2024-12-11 | C | 2.05 | 1777.5
2024-12-11 | C | 2.05 | 1745.11
2024-12-11 | C | 2.05 | 1738.87
2024-12-11 | C | 2.05 | 1737.47
2024-12-11 | C | 2.05 | 1749.46
2024-12-11 | C | 2.05 | 1742.52
2024-12-11 | C | 2.05 | 1745.62
2024-12-11 | C | 2.05 | 1740.17
2024-12-11 | C | 2.05 | 1734.83
2024-12-11 | C | 2.05 | 1786.15
2024-12-11 | C | 0 | 109.43
2024-12-11 | C | 0 | 119.03
2024-12-11 | C | 0 | 118.11
2024-12-11 | C | 0 | 125.39
2024-12-11 | C | 0 | 120.89
2024-12-11 | C | 0 | 110.38
2024-12-11 | C | 0 | 114.93
2024-12-11 | C | 0 | 112.3
2024-12-11 | C | 0 | 123.27
2024-12-11 | C | 0 | 113.5
2024-12-11 | C | 0 | 113.73
2024-12-11 | C | 0 | 120.03
2024-12-11 | C | 0 | 112.05
2024-12-11 | C | 0 | 115.71
2024-12-11 | C | 0 | 116.3
2024-12-11 | C | 0 | 124.2
2024-12-11 | C | 0 | 121.95
2024-12-11 | C | 0 | 122.03
2024-12-11 | C | 0 | 122.76
2024-12-11 | C | 0 | 113.44
2024-12-11 | C | NA | 4853.92
2024-12-11 | C | NA | 4930.21
2024-12-11 | C | NA | 4940.02
2024-12-11 | C | NA | 4850.15
2024-12-11 | C | NA | 4805.32
2024-12-11 | C | NA | 4901.57
2024-12-11 | C | NA | 4904.49
2024-12-11 | C | NA | 4881.95
2024-12-11 | C | NA | 4894.6
2024-12-11 | C | NA | 4868.78
2024-12-11 | C | NA | 4828.5
2024-12-11 | C | NA | 4842.9
2024-12-11 | C | NA | 4936.19
2024-12-11 | C | NA | 4936.86
2024-12-11 | C | NA | 4883.22
2024-12-11 | C | NA | 4894.72
2024-12-11 | C | NA | 4920.68
2024-12-11 | C | NA | 4920.74
2024-12-11 | C | NA | 4851.39
2024-12-11 | C | NA | 4882.71
2024-12-12 | C | 10.25 | 7305.16
2024-12-12 | C | 10.25 | 7313.55
2024-12-12 | C | 10.25 | 7345.61
2024-12-12 | C | 10.25 | 7275.99
2024-12-12 | C | 10.25 | 7350.68
2024-12-12 | C | 10.25 | 7243.99
2024-12-12 | C | 10.25 | 7395.38
2024-12-12 | C | 10.25 | 7373.22
2024-12-12 | C | 10.25 | 7353.65
2024-12-12 | C | 10.25 | 7403.15
2024-12-12 | C | 10.25 | 7345.95
2024-12-12 | C | 10.25 | 7287.49
2024-12-12 | C | 10.25 | 7375.39
2024-12-12 | C | 10.25 | 7414.68
2024-12-12 | C | 10.25 | 7302.96
2024-12-12 | C | 10.25 | 7303.49
2024-12-12 | C | 10.25 | 7390.25
2024-12-12 | C | 10.25 | 7336.14
2024-12-12 | C | 10.25 | 7324.6
2024-12-12 | C | 10.25 | 7429.36
2024-12-12 | C | 8.2 | 6282.18
2024-12-12 | C | 8.2 | 6243.27
2024-12-12 | C | 8.2 | 6265.93
2024-12-12 | C | 8.2 | 6306.72
2024-12-12 | C | 8.2 | 6255.23
2024-12-12 | C | 8.2 | 6239.06
2024-12-12 | C | 8.2 | 6228.27
2024-12-12 | C | 8.2 | 6284.45
2024-12-12 | C | 8.2 | 6249.54
2024-12-12 | C | 8.2 | 6275.97
2024-12-12 | C | 8.2 | 6257.44
2024-12-12 | C | 8.2 | 6266.42
2024-12-12 | C | 8.2 | 6286.44
2024-12-12 | C | 8.2 | 6251.84
2024-12-12 | C | 8.2 | 6278.13
2024-12-12 | C | 8.2 | 6253.01
2024-12-12 | C | 8.2 | 6246.42
2024-12-12 | C | 8.2 | 6298.45
2024-12-12 | C | 8.2 | 6343.44
2024-12-12 | C | 8.2 | 6284.24
2024-12-12 | C | 6.15 | 5003.27
2024-12-12 | C | 6.15 | 4894.47
2024-12-12 | C | 6.15 | 4896.92
2024-12-12 | C | 6.15 | 4913.03
2024-12-12 | C | 6.15 | 4867.36
2024-12-12 | C | 6.15 | 4846.06
2024-12-12 | C | 6.15 | 4933.84
2024-12-12 | C | 6.15 | 4921.82
2024-12-12 | C | 6.15 | 4905.86
2024-12-12 | C | 6.15 | 4865.27
2024-12-12 | C | 6.15 | 4914.53
2024-12-12 | C | 6.15 | 4918.52
2024-12-12 | C | 6.15 | 4942.18
2024-12-12 | C | 6.15 | 4847.92
2024-12-12 | C | 6.15 | 4881.61
2024-12-12 | C | 6.15 | 4965.87
2024-12-12 | C | 6.15 | 4915.59
2024-12-12 | C | 6.15 | 4908.57
2024-12-12 | C | 6.15 | 4944.11
2024-12-12 | C | 6.15 | 4928.37
2024-12-12 | C | 4.1 | 3339.07
2024-12-12 | C | 4.1 | 3335.69
2024-12-12 | C | 4.1 | 3304.87
2024-12-12 | C | 4.1 | 3334.64
2024-12-12 | C | 4.1 | 3290.69
2024-12-12 | C | 4.1 | 3360.46
2024-12-12 | C | 4.1 | 3353.7
2024-12-12 | C | 4.1 | 3355.63
2024-12-12 | C | 4.1 | 3307.68
2024-12-12 | C | 4.1 | 3291.63
2024-12-12 | C | 4.1 | 3332.17
2024-12-12 | C | 4.1 | 3306.79
2024-12-12 | C | 4.1 | 3365.82
2024-12-12 | C | 4.1 | 3253.28
2024-12-12 | C | 4.1 | 3369.77
2024-12-12 | C | 4.1 | 3253.02
2024-12-12 | C | 4.1 | 3342.8
2024-12-12 | C | 4.1 | 3303.33
2024-12-12 | C | 4.1 | 3381.03
2024-12-12 | C | 4.1 | 3316.97
2024-12-12 | C | 2.05 | 1716.32
2024-12-12 | C | 2.05 | 1745.44
2024-12-12 | C | 2.05 | 1674.52
2024-12-12 | C | 2.05 | 1721.2
2024-12-12 | C | 2.05 | 1702.65
2024-12-12 | C | 2.05 | 1769.81
2024-12-12 | C | 2.05 | 1729.76
2024-12-12 | C | 2.05 | 1721.67
2024-12-12 | C | 2.05 | 1723.35
2024-12-12 | C | 2.05 | 1731.85
2024-12-12 | C | 2.05 | 1726.02
2024-12-12 | C | 2.05 | 1791.55
2024-12-12 | C | 2.05 | 1754.27
2024-12-12 | C | 2.05 | 1667.99
2024-12-12 | C | 2.05 | 1720.63
2024-12-12 | C | 2.05 | 1729.58
2024-12-12 | C | 2.05 | 1695.61
2024-12-12 | C | 2.05 | 1763.03
2024-12-12 | C | 2.05 | 1771.09
2024-12-12 | C | 2.05 | 1781.54
2024-12-12 | C | 0 | 116.4
2024-12-12 | C | 0 | 122.26
2024-12-12 | C | 0 | 119.1
2024-12-12 | C | 0 | 122.9
2024-12-12 | C | 0 | 118.4
2024-12-12 | C | 0 | 114.78
2024-12-12 | C | 0 | 114.08
2024-12-12 | C | 0 | 119.28
2024-12-12 | C | 0 | 121.97
2024-12-12 | C | 0 | 104.16
2024-12-12 | C | 0 | 113.67
2024-12-12 | C | 0 | 119.5
2024-12-12 | C | 0 | 116.86
2024-12-12 | C | 0 | 111.26
2024-12-12 | C | 0 | 115.77
2024-12-12 | C | 0 | 111.58
2024-12-12 | C | 0 | 116.27
2024-12-12 | C | 0 | 123.85
2024-12-12 | C | 0 | 118.49
2024-12-12 | C | 0 | 117.53
2024-12-12 | C | NA | 4848.3
2024-12-12 | C | NA | 4895.67
2024-12-12 | C | NA | 4919.87
2024-12-12 | C | NA | 4810.93
2024-12-12 | C | NA | 4899.85
2024-12-12 | C | NA | 4846.15
2024-12-12 | C | NA | 4818.57
2024-12-12 | C | NA | 4847.48
2024-12-12 | C | NA | 4812.4
2024-12-12 | C | NA | 4907.89
2024-12-12 | C | NA | 4838.69
2024-12-12 | C | NA | 4904.45
2024-12-12 | C | NA | 4854.56
2024-12-12 | C | NA | 4802.97
2024-12-12 | C | NA | 4858.35
2024-12-12 | C | NA | 4915.81
2024-12-12 | C | NA | 4770.7
2024-12-12 | C | NA | 4798.78
2024-12-12 | C | NA | 4761.47
2024-12-12 | C | NA | 4899.29
2024-12-13 | C | 10.25 | 7265.14
2024-12-13 | C | 10.25 | 7360.89
2024-12-13 | C | 10.25 | 7388.48
2024-12-13 | C | 10.25 | 7286.25
2024-12-13 | C | 10.25 | 7475.62
2024-12-13 | C | 10.25 | 7256.37
2024-12-13 | C | 10.25 | 7334.46
2024-12-13 | C | 10.25 | 7258.7
2024-12-13 | C | 10.25 | 7247.13
2024-12-13 | C | 10.25 | 7341.21
2024-12-13 | C | 10.25 | 7412.54
2024-12-13 | C | 10.25 | 7294.86
2024-12-13 | C | 10.25 | 7251.25
2024-12-13 | C | 10.25 | 7347.15
2024-12-13 | C | 10.25 | 7335.49
2024-12-13 | C | 10.25 | 7296.73
2024-12-13 | C | 10.25 | 7265.77
2024-12-13 | C | 10.25 | 7263.91
2024-12-13 | C | 10.25 | 7336.33
2024-12-13 | C | 10.25 | 7377.93
2024-12-13 | C | 8.2 | 6308.79
2024-12-13 | C | 8.2 | 6218.89
2024-12-13 | C | 8.2 | 6215.85
2024-12-13 | C | 8.2 | 6355.94
2024-12-13 | C | 8.2 | 6303.11
2024-12-13 | C | 8.2 | 6313.32
2024-12-13 | C | 8.2 | 6241.22
2024-12-13 | C | 8.2 | 6246.5
2024-12-13 | C | 8.2 | 6345.62
2024-12-13 | C | 8.2 | 6290.56
2024-12-13 | C | 8.2 | 6258.08
2024-12-13 | C | 8.2 | 6317.16
2024-12-13 | C | 8.2 | 6294.23
2024-12-13 | C | 8.2 | 6400.78
2024-12-13 | C | 8.2 | 6213.15
2024-12-13 | C | 8.2 | 6197.58
2024-12-13 | C | 8.2 | 6195.32
2024-12-13 | C | 8.2 | 6269
2024-12-13 | C | 8.2 | 6269.89
2024-12-13 | C | 8.2 | 6300.35
2024-12-13 | C | 6.15 | 4963.32
2024-12-13 | C | 6.15 | 4970.92
2024-12-13 | C | 6.15 | 4902.81
2024-12-13 | C | 6.15 | 4928.94
2024-12-13 | C | 6.15 | 4962.26
2024-12-13 | C | 6.15 | 4881.17
2024-12-13 | C | 6.15 | 4843.03
2024-12-13 | C | 6.15 | 4915.48
2024-12-13 | C | 6.15 | 4860.04
2024-12-13 | C | 6.15 | 4933.72
2024-12-13 | C | 6.15 | 4970.98
2024-12-13 | C | 6.15 | 4962.97
2024-12-13 | C | 6.15 | 4962.03
2024-12-13 | C | 6.15 | 4888.81
2024-12-13 | C | 6.15 | 4905.77
2024-12-13 | C | 6.15 | 4901.19
2024-12-13 | C | 6.15 | 4853.29
2024-12-13 | C | 6.15 | 4884.25
2024-12-13 | C | 6.15 | 4923.43
2024-12-13 | C | 6.15 | 4948.38
2024-12-13 | C | 4.1 | 3337.49
2024-12-13 | C | 4.1 | 3328.45
2024-12-13 | C | 4.1 | 3289.39
2024-12-13 | C | 4.1 | 3262.98
2024-12-13 | C | 4.1 | 3257.95
2024-12-13 | C | 4.1 | 3328.11
2024-12-13 | C | 4.1 | 3295.96
2024-12-13 | C | 4.1 | 3318.82
2024-12-13 | C | 4.1 | 3311.1
2024-12-13 | C | 4.1 | 3330.45
2024-12-13 | C | 4.1 | 3346.51
2024-12-13 | C | 4.1 | 3316.46
2024-12-13 | C | 4.1 | 3340.02
2024-12-13 | C | 4.1 | 3305.22
2024-12-13 | C | 4.1 | 3341.59
2024-12-13 | C | 4.1 | 3306.9
2024-12-13 | C | 4.1 | 3350.96
2024-12-13 | C | 4.1 | 3319.89
2024-12-13 | C | 4.1 | 3364.44
2024-12-13 | C | 4.1 | 3323.63
2024-12-13 | C | 2.05 | 1705.42
2024-12-13 | C | 2.05 | 1719.92
2024-12-13 | C | 2.05 | 1711.87
2024-12-13 | C | 2.05 | 1750.44
2024-12-13 | C | 2.05 | 1710.93
2024-12-13 | C | 2.05 | 1752.23
2024-12-13 | C | 2.05 | 1792.76
2024-12-13 | C | 2.05 | 1712.74
2024-12-13 | C | 2.05 | 1789.78
2024-12-13 | C | 2.05 | 1741.48
2024-12-13 | C | 2.05 | 1767.27
2024-12-13 | C | 2.05 | 1697.35
2024-12-13 | C | 2.05 | 1742.32
2024-12-13 | C | 2.05 | 1731.52
2024-12-13 | C | 2.05 | 1742.82
2024-12-13 | C | 2.05 | 1733.3
2024-12-13 | C | 2.05 | 1726.48
2024-12-13 | C | 2.05 | 1749.58
2024-12-13 | C | 2.05 | 1724.42
2024-12-13 | C | 2.05 | 1756.65
2024-12-13 | C | 0 | 119.51
2024-12-13 | C | 0 | 110.99
2024-12-13 | C | 0 | 124.83
2024-12-13 | C | 0 | 119.26
2024-12-13 | C | 0 | 128.57
2024-12-13 | C | 0 | 127.92
2024-12-13 | C | 0 | 114.3
2024-12-13 | C | 0 | 116.82
2024-12-13 | C | 0 | 115.98
2024-12-13 | C | 0 | 116.4
2024-12-13 | C | 0 | 116.35
2024-12-13 | C | 0 | 122.23
2024-12-13 | C | 0 | 117.35
2024-12-13 | C | 0 | 116.09
2024-12-13 | C | 0 | 120.67
2024-12-13 | C | 0 | 119.7
2024-12-13 | C | 0 | 125.99
2024-12-13 | C | 0 | 127.35
2024-12-13 | C | 0 | 126.89
2024-12-13 | C | 0 | 126.22
2024-12-13 | C | NA | 4921.55
2024-12-13 | C | NA | 4930.23
2024-12-13 | C | NA | 4927.9
2024-12-13 | C | NA | 4755.95
2024-12-13 | C | NA | 4923.41
2024-12-13 | C | NA | 4850.47
2024-12-13 | C | NA | 4853.39
2024-12-13 | C | NA | 4921.97
2024-12-13 | C | NA | 4989.95
2024-12-13 | C | NA | 4886.85
2024-12-13 | C | NA | 4937.7
2024-12-13 | C | NA | 4957.38
2024-12-13 | C | NA | 4859.42
2024-12-13 | C | NA | 4818.36
2024-12-13 | C | NA | 4762.06
2024-12-13 | C | NA | 4742.42
2024-12-13 | C | NA | 4800.98
2024-12-13 | C | NA | 4877.5
2024-12-13 | C | NA | 4841.38
2024-12-13 | C | NA | 4864.85
2024-12-14 | C | 10.25 | 7290.35
2024-12-14 | C | 10.25 | 7269.79
2024-12-14 | C | 10.25 | 7342.57
2024-12-14 | C | 10.25 | 7240.96
2024-12-14 | C | 10.25 | 7339.71
2024-12-14 | C | 10.25 | 7297.77
2024-12-14 | C | 10.25 | 7305.41
2024-12-14 | C | 10.25 | 7261.01
2024-12-14 | C | 10.25 | 7375.85
2024-12-14 | C | 10.25 | 7292.83
2024-12-14 | C | 10.25 | 7369.72
2024-12-14 | C | 10.25 | 7363.9
2024-12-14 | C | 10.25 | 7393.17
2024-12-14 | C | 10.25 | 7278.87
2024-12-14 | C | 10.25 | 7412.99
2024-12-14 | C | 10.25 | 7395.95
2024-12-14 | C | 10.25 | 7299.7
2024-12-14 | C | 10.25 | 7316.89
2024-12-14 | C | 10.25 | 7238.07
2024-12-14 | C | 10.25 | 7320.65
2024-12-14 | C | 8.2 | 6303.01
2024-12-14 | C | 8.2 | 6305.97
2024-12-14 | C | 8.2 | 6255.62
2024-12-14 | C | 8.2 | 6202.59
2024-12-14 | C | 8.2 | 6326.03
2024-12-14 | C | 8.2 | 6309.65
2024-12-14 | C | 8.2 | 6249.11
2024-12-14 | C | 8.2 | 6224.61
2024-12-14 | C | 8.2 | 6254.72
2024-12-14 | C | 8.2 | 6269.59
2024-12-14 | C | 8.2 | 6285.93
2024-12-14 | C | 8.2 | 6279.21
2024-12-14 | C | 8.2 | 6258.46
2024-12-14 | C | 8.2 | 6308.73
2024-12-14 | C | 8.2 | 6334.66
2024-12-14 | C | 8.2 | 6256.08
2024-12-14 | C | 8.2 | 6326.67
2024-12-14 | C | 8.2 | 6264.89
2024-12-14 | C | 8.2 | 6278.35
2024-12-14 | C | 8.2 | 6308.7
2024-12-14 | C | 6.15 | 4869.68
2024-12-14 | C | 6.15 | 4958.3
2024-12-14 | C | 6.15 | 4960.16
2024-12-14 | C | 6.15 | 4942.78
2024-12-14 | C | 6.15 | 4880.99
2024-12-14 | C | 6.15 | 4888.79
2024-12-14 | C | 6.15 | 4962.65
2024-12-14 | C | 6.15 | 4913.59
2024-12-14 | C | 6.15 | 4906.87
2024-12-14 | C | 6.15 | 4951.97
2024-12-14 | C | 6.15 | 4902.5
2024-12-14 | C | 6.15 | 4883.3
2024-12-14 | C | 6.15 | 4929.16
2024-12-14 | C | 6.15 | 4891.35
2024-12-14 | C | 6.15 | 4890.97
2024-12-14 | C | 6.15 | 4998.14
2024-12-14 | C | 6.15 | 4946.35
2024-12-14 | C | 6.15 | 5016
2024-12-14 | C | 6.15 | 4908.39
2024-12-14 | C | 6.15 | 4905.89
2024-12-14 | C | 4.1 | 3294.22
2024-12-14 | C | 4.1 | 3343.63
2024-12-14 | C | 4.1 | 3244.99
2024-12-14 | C | 4.1 | 3289.06
2024-12-14 | C | 4.1 | 3285.9
2024-12-14 | C | 4.1 | 3307.37
2024-12-14 | C | 4.1 | 3265.27
2024-12-14 | C | 4.1 | 3340.98
2024-12-14 | C | 4.1 | 3321.86
2024-12-14 | C | 4.1 | 3320.56
2024-12-14 | C | 4.1 | 3331.76
2024-12-14 | C | 4.1 | 3252.75
2024-12-14 | C | 4.1 | 3313.1
2024-12-14 | C | 4.1 | 3330.57
2024-12-14 | C | 4.1 | 3291.53
2024-12-14 | C | 4.1 | 3295.2
2024-12-14 | C | 4.1 | 3291.76
2024-12-14 | C | 4.1 | 3285.19
2024-12-14 | C | 4.1 | 3333.73
2024-12-14 | C | 4.1 | 3317.97
2024-12-14 | C | 2.05 | 1754.36
2024-12-14 | C | 2.05 | 1735.39
2024-12-14 | C | 2.05 | 1722.86
2024-12-14 | C | 2.05 | 1766.98
2024-12-14 | C | 2.05 | 1726.01
2024-12-14 | C | 2.05 | 1714.68
2024-12-14 | C | 2.05 | 1800.97
2024-12-14 | C | 2.05 | 1817.04
2024-12-14 | C | 2.05 | 1732.58
2024-12-14 | C | 2.05 | 1780.36
2024-12-14 | C | 2.05 | 1738.5
2024-12-14 | C | 2.05 | 1737.46
2024-12-14 | C | 2.05 | 1729.02
2024-12-14 | C | 2.05 | 1778.02
2024-12-14 | C | 2.05 | 1723.35
2024-12-14 | C | 2.05 | 1757.38
2024-12-14 | C | 2.05 | 1756.44
2024-12-14 | C | 2.05 | 1743.05
2024-12-14 | C | 2.05 | 1745.67
2024-12-14 | C | 2.05 | 1783.07
2024-12-14 | C | 0 | 114.02
2024-12-14 | C | 0 | 124.98
2024-12-14 | C | 0 | 123.59
2024-12-14 | C | 0 | 107.16
2024-12-14 | C | 0 | 121.53
2024-12-14 | C | 0 | 121.87
2024-12-14 | C | 0 | 123.04
2024-12-14 | C | 0 | 115.05
2024-12-14 | C | 0 | 120.11
2024-12-14 | C | 0 | 120.65
2024-12-14 | C | 0 | 119.78
2024-12-14 | C | 0 | 121.45
2024-12-14 | C | 0 | 121.39
2024-12-14 | C | 0 | 109.79
2024-12-14 | C | 0 | 114.1
2024-12-14 | C | 0 | 106.69
2024-12-14 | C | 0 | 114.85
2024-12-14 | C | 0 | 121.59
2024-12-14 | C | 0 | 116.52
2024-12-14 | C | 0 | 118.53
2024-12-14 | C | NA | 4896.05
2024-12-14 | C | NA | 4853.63
2024-12-14 | C | NA | 4931.15
2024-12-14 | C | NA | 4880.8
2024-12-14 | C | NA | 4899.29
2024-12-14 | C | NA | 4835.19
2024-12-14 | C | NA | 4930.01
2024-12-14 | C | NA | 4930.61
2024-12-14 | C | NA | 4862.92
2024-12-14 | C | NA | 4857.65
2024-12-14 | C | NA | 4835.31
2024-12-14 | C | NA | 4798.54
2024-12-14 | C | NA | 4931.6
2024-12-14 | C | NA | 4923.18
2024-12-14 | C | NA | 5013.96
2024-12-14 | C | NA | 4902.61
2024-12-14 | C | NA | 4897.9
2024-12-14 | C | NA | 4814.68
2024-12-14 | C | NA | 4801.06
2024-12-14 | C | NA | 4926.21

## Datas
Date | Operateur | Concentration | Intensite
2024-11-18 | A | 10.25 | 7367.83
2024-11-18 | A | 10.25 | 7263.95
2024-11-18 | A | 10.25 | 7249.46
2024-11-18 | A | 10.25 | 7308.53
2024-11-18 | A | 10.25 | 7400.25
2024-11-18 | A | 10.25 | 7330.93
2024-11-18 | A | 10.25 | 7385.26
2024-11-18 | A | 10.25 | 7333.84
2024-11-18 | A | 10.25 | 7278.31
2024-11-18 | A | 10.25 | 7318.61
2024-11-18 | A | 10.25 | 7300.5
2024-11-18 | A | 10.25 | 7284.41
2024-11-18 | A | 10.25 | 7338.53
2024-11-18 | A | 10.25 | 7381.66
2024-11-18 | A | 10.25 | 7432
2024-11-18 | A | 10.25 | 7251.22
2024-11-18 | A | 10.25 | 7374.35
2024-11-18 | A | 10.25 | 7327.14
2024-11-18 | A | 10.25 | 7333.99
2024-11-18 | A | 10.25 | 7267.45
2024-11-18 | A | 8.2 | 6389.8
2024-11-18 | A | 8.2 | 6358.86
2024-11-18 | A | 8.2 | 6273.9
2024-11-18 | A | 8.2 | 6225.39
2024-11-18 | A | 8.2 | 6268.5
2024-11-18 | A | 8.2 | 6273.57
2024-11-18 | A | 8.2 | 6245.52
2024-11-18 | A | 8.2 | 6189.69
2024-11-18 | A | 8.2 | 6290.13
2024-11-18 | A | 8.2 | 6301.78
2024-11-18 | A | 8.2 | 6300.06
2024-11-18 | A | 8.2 | 6365.03
2024-11-18 | A | 8.2 | 6334.79
2024-11-18 | A | 8.2 | 6340.71
2024-11-18 | A | 8.2 | 6327.63
2024-11-18 | A | 8.2 | 6262.26
2024-11-18 | A | 8.2 | 6273.18
2024-11-18 | A | 8.2 | 6298.54
2024-11-18 | A | 8.2 | 6240.91
2024-11-18 | A | 8.2 | 6195
2024-11-18 | A | 6.15 | 5016.79
2024-11-18 | A | 6.15 | 4915.47
2024-11-18 | A | 6.15 | 4929.34
2024-11-18 | A | 6.15 | 4868.7
2024-11-18 | A | 6.15 | 4924.22
2024-11-18 | A | 6.15 | 4872.78
2024-11-18 | A | 6.15 | 5035.98
2024-11-18 | A | 6.15 | 4891.19
2024-11-18 | A | 6.15 | 4875.23
2024-11-18 | A | 6.15 | 4892.68
2024-11-18 | A | 6.15 | 4846.29
2024-11-18 | A | 6.15 | 4889.43
2024-11-18 | A | 6.15 | 4891.3
2024-11-18 | A | 6.15 | 4932.95
2024-11-18 | A | 6.15 | 4857.66
2024-11-18 | A | 6.15 | 4977.3
2024-11-18 | A | 6.15 | 4871.78
2024-11-18 | A | 6.15 | 4955.57
2024-11-18 | A | 6.15 | 4989.38
2024-11-18 | A | 6.15 | 4927.89
2024-11-18 | A | 4.1 | 3341.95
2024-11-18 | A | 4.1 | 3369.45
2024-11-18 | A | 4.1 | 3325.12
2024-11-18 | A | 4.1 | 3344.93
2024-11-18 | A | 4.1 | 3302.1
2024-11-18 | A | 4.1 | 3295.22
2024-11-18 | A | 4.1 | 3318.66
2024-11-18 | A | 4.1 | 3375.76
2024-11-18 | A | 4.1 | 3294.91
2024-11-18 | A | 4.1 | 3343.78
2024-11-18 | A | 4.1 | 3303.86
2024-11-18 | A | 4.1 | 3322.21
2024-11-18 | A | 4.1 | 3338.89
2024-11-18 | A | 4.1 | 3347.81
2024-11-18 | A | 4.1 | 3346.01
2024-11-18 | A | 4.1 | 3306.53
2024-11-18 | A | 4.1 | 3316.64
2024-11-18 | A | 4.1 | 3233.33
2024-11-18 | A | 4.1 | 3314.18
2024-11-18 | A | 4.1 | 3297.4
2024-11-18 | A | 2.05 | 1731.23
2024-11-18 | A | 2.05 | 1716.01
2024-11-18 | A | 2.05 | 1742.52
2024-11-18 | A | 2.05 | 1748.07
2024-11-18 | A | 2.05 | 1731.08
2024-11-18 | A | 2.05 | 1750.37
2024-11-18 | A | 2.05 | 1768.19
2024-11-18 | A | 2.05 | 1763.23
2024-11-18 | A | 2.05 | 1742.57
2024-11-18 | A | 2.05 | 1753.18
2024-11-18 | A | 2.05 | 1725.24
2024-11-18 | A | 2.05 | 1724.66
2024-11-18 | A | 2.05 | 1758.39
2024-11-18 | A | 2.05 | 1755.95
2024-11-18 | A | 2.05 | 1764.43
2024-11-18 | A | 2.05 | 1752.76
2024-11-18 | A | 2.05 | 1762.23
2024-11-18 | A | 2.05 | 1746.97
2024-11-18 | A | 2.05 | 1700.05
2024-11-18 | A | 2.05 | 1755.62
2024-11-18 | A | 0 | 119.54
2024-11-18 | A | 0 | 122.07
2024-11-18 | A | 0 | 130.4
2024-11-18 | A | 0 | 119.75
2024-11-18 | A | 0 | 120.38
2024-11-18 | A | 0 | 116.79
2024-11-18 | A | 0 | 125.75
2024-11-18 | A | 0 | 115.65
2024-11-18 | A | 0 | 124.07
2024-11-18 | A | 0 | 122.1
2024-11-18 | A | 0 | 120.37
2024-11-18 | A | 0 | 113.32
2024-11-18 | A | 0 | 123.96
2024-11-18 | A | 0 | 128.88
2024-11-18 | A | 0 | 112.67
2024-11-18 | A | 0 | 118.41
2024-11-18 | A | 0 | 128.81
2024-11-18 | A | 0 | 116.96
2024-11-18 | A | 0 | 114.2
2024-11-18 | A | 0 | 115.71
2024-11-18 | A | NA | 4914.46 | 5651.629
2024-11-18 | A | NA | 4887.74 | 5620.901
2024-11-18 | A | NA | 4908.91 | 5645.2465
2024-11-18 | A | NA | 4847.83 | 5575.0045
2024-11-18 | A | NA | 4878.65 | 5610.4475
2024-11-18 | A | NA | 4881.74 | 5614.001
2024-11-18 | A | NA | 4830.44 | 5555.006
2024-11-18 | A | NA | 4891.17 | 5624.8455
2024-11-18 | A | NA | 4845.7 | 5572.555
2024-11-18 | A | NA | 4945.94 | 5687.831
2024-11-18 | A | NA | 4831.15 | 5555.8225
2024-11-18 | A | NA | 4790 | 5508.5
2024-11-18 | A | NA | 4877.14 | 5608.711
2024-11-18 | A | NA | 4962.04 | 5706.346
2024-11-18 | A | NA | 4847.4 | 5574.51
2024-11-18 | A | NA | 4964.61 | 5709.3015
2024-11-18 | A | NA | 4855.7 | 5584.055
2024-11-18 | A | NA | 4855.88 | 5584.262
2024-11-18 | A | NA | 4982.89 | 5730.3235
2024-11-18 | A | NA | 4874.13 | 5605.2495
2024-11-19 | A | 10.25 | 7260.36
2024-11-19 | A | 10.25 | 7342.97
2024-11-19 | A | 10.25 | 7271.95
2024-11-19 | A | 10.25 | 7294.51
2024-11-19 | A | 10.25 | 7287.44
2024-11-19 | A | 10.25 | 7216.3
2024-11-19 | A | 10.25 | 7310.03
2024-11-19 | A | 10.25 | 7341.1
2024-11-19 | A | 10.25 | 7323.89
2024-11-19 | A | 10.25 | 7321.04
2024-11-19 | A | 10.25 | 7185.19
2024-11-19 | A | 10.25 | 7339.93
2024-11-19 | A | 10.25 | 7370.18
2024-11-19 | A | 10.25 | 7298.95
2024-11-19 | A | 10.25 | 7423.35
2024-11-19 | A | 10.25 | 7360.47
2024-11-19 | A | 10.25 | 7388.15
2024-11-19 | A | 10.25 | 7330.75
2024-11-19 | A | 10.25 | 7436.16
2024-11-19 | A | 10.25 | 7313.06
2024-11-19 | A | 8.2 | 6287.88
2024-11-19 | A | 8.2 | 6314.25
2024-11-19 | A | 8.2 | 6292.95
2024-11-19 | A | 8.2 | 6313.39
2024-11-19 | A | 8.2 | 6330.57
2024-11-19 | A | 8.2 | 6298.52
2024-11-19 | A | 8.2 | 6342.48
2024-11-19 | A | 8.2 | 6263.19
2024-11-19 | A | 8.2 | 6345.01
2024-11-19 | A | 8.2 | 6248.07
2024-11-19 | A | 8.2 | 6302.61
2024-11-19 | A | 8.2 | 6195.8
2024-11-19 | A | 8.2 | 6318.09
2024-11-19 | A | 8.2 | 6321
2024-11-19 | A | 8.2 | 6260.85
2024-11-19 | A | 8.2 | 6346.42
2024-11-19 | A | 8.2 | 6279.29
2024-11-19 | A | 8.2 | 6254.38
2024-11-19 | A | 8.2 | 6266.57
2024-11-19 | A | 8.2 | 6329.93
2024-11-19 | A | 6.15 | 4893.57
2024-11-19 | A | 6.15 | 4850.37
2024-11-19 | A | 6.15 | 4846.6
2024-11-19 | A | 6.15 | 4875.39
2024-11-19 | A | 6.15 | 4940.5
2024-11-19 | A | 6.15 | 4949.39
2024-11-19 | A | 6.15 | 4919.4
2024-11-19 | A | 6.15 | 4938.63
2024-11-19 | A | 6.15 | 4993.53
2024-11-19 | A | 6.15 | 4919.56
2024-11-19 | A | 6.15 | 4918.41
2024-11-19 | A | 6.15 | 4932.38
2024-11-19 | A | 6.15 | 4869.3
2024-11-19 | A | 6.15 | 4846.47
2024-11-19 | A | 6.15 | 4922.04
2024-11-19 | A | 6.15 | 4893.81
2024-11-19 | A | 6.15 | 4938.3
2024-11-19 | A | 6.15 | 4898.63
2024-11-19 | A | 6.15 | 4884.72
2024-11-19 | A | 6.15 | 4904.94
2024-11-19 | A | 4.1 | 3292.84
2024-11-19 | A | 4.1 | 3312.31
2024-11-19 | A | 4.1 | 3344.23
2024-11-19 | A | 4.1 | 3268.2
2024-11-19 | A | 4.1 | 3331.8
2024-11-19 | A | 4.1 | 3267.42
2024-11-19 | A | 4.1 | 3311.62
2024-11-19 | A | 4.1 | 3325.36
2024-11-19 | A | 4.1 | 3335.03
2024-11-19 | A | 4.1 | 3298.19
2024-11-19 | A | 4.1 | 3340.48
2024-11-19 | A | 4.1 | 3300.2
2024-11-19 | A | 4.1 | 3330.29
2024-11-19 | A | 4.1 | 3324.69
2024-11-19 | A | 4.1 | 3335.86
2024-11-19 | A | 4.1 | 3320.66
2024-11-19 | A | 4.1 | 3274.43
2024-11-19 | A | 4.1 | 3295.78
2024-11-19 | A | 4.1 | 3349.02
2024-11-19 | A | 4.1 | 3294.33
2024-11-19 | A | 2.05 | 1710.06
2024-11-19 | A | 2.05 | 1730.14
2024-11-19 | A | 2.05 | 1702.6
2024-11-19 | A | 2.05 | 1757.64
2024-11-19 | A | 2.05 | 1760.56
2024-11-19 | A | 2.05 | 1739.23
2024-11-19 | A | 2.05 | 1727.96
2024-11-19 | A | 2.05 | 1729.78
2024-11-19 | A | 2.05 | 1775.75
2024-11-19 | A | 2.05 | 1752.52
2024-11-19 | A | 2.05 | 1765.65
2024-11-19 | A | 2.05 | 1712.01
2024-11-19 | A | 2.05 | 1753.61
2024-11-19 | A | 2.05 | 1723.56
2024-11-19 | A | 2.05 | 1760.8
2024-11-19 | A | 2.05 | 1737.28
2024-11-19 | A | 2.05 | 1787.86
2024-11-19 | A | 2.05 | 1739.03
2024-11-19 | A | 2.05 | 1723.48
2024-11-19 | A | 2.05 | 1735.7
2024-11-19 | A | 0 | 117.29
2024-11-19 | A | 0 | 115.94
2024-11-19 | A | 0 | 114.04
2024-11-19 | A | 0 | 121.04
2024-11-19 | A | 0 | 128.71
2024-11-19 | A | 0 | 118.85
2024-11-19 | A | 0 | 115.78
2024-11-19 | A | 0 | 112.66
2024-11-19 | A | 0 | 117.55
2024-11-19 | A | 0 | 110.67
2024-11-19 | A | 0 | 116.94
2024-11-19 | A | 0 | 112.19
2024-11-19 | A | 0 | 119.77
2024-11-19 | A | 0 | 112.86
2024-11-19 | A | 0 | 124.53
2024-11-19 | A | 0 | 119.21
2024-11-19 | A | 0 | 119.71
2024-11-19 | A | 0 | 112.39
2024-11-19 | A | 0 | 120.03
2024-11-19 | A | 0 | 110.72
2024-11-19 | A | NA | 4906.99 | 5643.0385
2024-11-19 | A | NA | 4894.58 | 5628.767
2024-11-19 | A | NA | 4972.45 | 5718.3175
2024-11-19 | A | NA | 4885.33 | 5618.1295
2024-11-19 | A | NA | 4859.85 | 5588.8275
2024-11-19 | A | NA | 4922.87 | 5661.3005
2024-11-19 | A | NA | 4999.62 | 5749.563
2024-11-19 | A | NA | 4919.46 | 5657.379
2024-11-19 | A | NA | 4938.87 | 5679.7005
2024-11-19 | A | NA | 4923.04 | 5661.496
2024-11-19 | A | NA | 4843.22 | 5569.703
2024-11-19 | A | NA | 4782.07 | 5499.3805
2024-11-19 | A | NA | 4835.61 | 5560.9515
2024-11-19 | A | NA | 4885.03 | 5617.7845
2024-11-19 | A | NA | 4905.84 | 5641.716
2024-11-19 | A | NA | 5003.88 | 5754.462
2024-11-19 | A | NA | 4863.56 | 5593.094
2024-11-19 | A | NA | 4863.95 | 5593.5425
2024-11-19 | A | NA | 4924.71 | 5663.4165
2024-11-19 | A | NA | 4864.45 | 5594.1175
2024-11-20 | A | 10.25 | 7296.68
2024-11-20 | A | 10.25 | 7386.84
2024-11-20 | A | 10.25 | 7467.84
2024-11-20 | A | 10.25 | 7276.16
2024-11-20 | A | 10.25 | 7282.41
2024-11-20 | A | 10.25 | 7254.4
2024-11-20 | A | 10.25 | 7377.56
2024-11-20 | A | 10.25 | 7269.14
2024-11-20 | A | 10.25 | 7467.23
2024-11-20 | A | 10.25 | 7313.41
2024-11-20 | A | 10.25 | 7385.84
2024-11-20 | A | 10.25 | 7290.59
2024-11-20 | A | 10.25 | 7327.62
2024-11-20 | A | 10.25 | 7307.8
2024-11-20 | A | 10.25 | 7365.14
2024-11-20 | A | 10.25 | 7349.67
2024-11-20 | A | 10.25 | 7346.62
2024-11-20 | A | 10.25 | 7349.22
2024-11-20 | A | 10.25 | 7389.99
2024-11-20 | A | 10.25 | 7239.14
2024-11-20 | A | 8.2 | 6294.24
2024-11-20 | A | 8.2 | 6228.27
2024-11-20 | A | 8.2 | 6249.02
2024-11-20 | A | 8.2 | 6296.21
2024-11-20 | A | 8.2 | 6240.62
2024-11-20 | A | 8.2 | 6250.12
2024-11-20 | A | 8.2 | 6282.56
2024-11-20 | A | 8.2 | 6312.84
2024-11-20 | A | 8.2 | 6214.89
2024-11-20 | A | 8.2 | 6222.88
2024-11-20 | A | 8.2 | 6292.15
2024-11-20 | A | 8.2 | 6231.05
2024-11-20 | A | 8.2 | 6314.05
2024-11-20 | A | 8.2 | 6369.35
2024-11-20 | A | 8.2 | 6294.95
2024-11-20 | A | 8.2 | 6244.1
2024-11-20 | A | 8.2 | 6267.9
2024-11-20 | A | 8.2 | 6298.35
2024-11-20 | A | 8.2 | 6315.89
2024-11-20 | A | 8.2 | 6268.51
2024-11-20 | A | 6.15 | 4833.39
2024-11-20 | A | 6.15 | 4944.48
2024-11-20 | A | 6.15 | 4855.81
2024-11-20 | A | 6.15 | 4887.72
2024-11-20 | A | 6.15 | 4909.73
2024-11-20 | A | 6.15 | 4957.51
2024-11-20 | A | 6.15 | 4867.71
2024-11-20 | A | 6.15 | 4890.4
2024-11-20 | A | 6.15 | 4968.9
2024-11-20 | A | 6.15 | 4857.58
2024-11-20 | A | 6.15 | 4956.14
2024-11-20 | A | 6.15 | 4998.3
2024-11-20 | A | 6.15 | 4941.65
2024-11-20 | A | 6.15 | 4927.03
2024-11-20 | A | 6.15 | 4881.15
2024-11-20 | A | 6.15 | 4874.16
2024-11-20 | A | 6.15 | 4841.47
2024-11-20 | A | 6.15 | 4883.16
2024-11-20 | A | 6.15 | 4991.35
2024-11-20 | A | 6.15 | 4951.05
2024-11-20 | A | 4.1 | 3335.59
2024-11-20 | A | 4.1 | 3317.48
2024-11-20 | A | 4.1 | 3328.9
2024-11-20 | A | 4.1 | 3410.92
2024-11-20 | A | 4.1 | 3362.37
2024-11-20 | A | 4.1 | 3333.9
2024-11-20 | A | 4.1 | 3290.27
2024-11-20 | A | 4.1 | 3300.26
2024-11-20 | A | 4.1 | 3323.96
2024-11-20 | A | 4.1 | 3301.43
2024-11-20 | A | 4.1 | 3324.74
2024-11-20 | A | 4.1 | 3342.14
2024-11-20 | A | 4.1 | 3340.47
2024-11-20 | A | 4.1 | 3290.15
2024-11-20 | A | 4.1 | 3302.06
2024-11-20 | A | 4.1 | 3338.4
2024-11-20 | A | 4.1 | 3315.67
2024-11-20 | A | 4.1 | 3303.28
2024-11-20 | A | 4.1 | 3257.61
2024-11-20 | A | 4.1 | 3289.92
2024-11-20 | A | 2.05 | 1752.99
2024-11-20 | A | 2.05 | 1745.42
2024-11-20 | A | 2.05 | 1758.9
2024-11-20 | A | 2.05 | 1722.66
2024-11-20 | A | 2.05 | 1763.42
2024-11-20 | A | 2.05 | 1707.73
2024-11-20 | A | 2.05 | 1732.57
2024-11-20 | A | 2.05 | 1716.02
2024-11-20 | A | 2.05 | 1728.45
2024-11-20 | A | 2.05 | 1766.69
2024-11-20 | A | 2.05 | 1713.35
2024-11-20 | A | 2.05 | 1729.1
2024-11-20 | A | 2.05 | 1769.6
2024-11-20 | A | 2.05 | 1762.35
2024-11-20 | A | 2.05 | 1738.1
2024-11-20 | A | 2.05 | 1731.04
2024-11-20 | A | 2.05 | 1790.72
2024-11-20 | A | 2.05 | 1774.48
2024-11-20 | A | 2.05 | 1729.6
2024-11-20 | A | 2.05 | 1716.7
2024-11-20 | A | 0 | 121.04
2024-11-20 | A | 0 | 121.12
2024-11-20 | A | 0 | 119.79
2024-11-20 | A | 0 | 118.5
2024-11-20 | A | 0 | 113.02
2024-11-20 | A | 0 | 113.78
2024-11-20 | A | 0 | 114.33
2024-11-20 | A | 0 | 116.78
2024-11-20 | A | 0 | 116.72
2024-11-20 | A | 0 | 121.77
2024-11-20 | A | 0 | 121.15
2024-11-20 | A | 0 | 115.66
2024-11-20 | A | 0 | 121.5
2024-11-20 | A | 0 | 121.43
2024-11-20 | A | 0 | 118.67
2024-11-20 | A | 0 | 126.22
2024-11-20 | A | 0 | 113.42
2024-11-20 | A | 0 | 120.84
2024-11-20 | A | 0 | 113.7
2024-11-20 | A | 0 | 119.54
2024-11-20 | A | NA | 4877.82 | 5609.493
2024-11-20 | A | NA | 4905.94 | 5641.831
2024-11-20 | A | NA | 4890.25 | 5623.7875
2024-11-20 | A | NA | 4925.43 | 5664.2445
2024-11-20 | A | NA | 4963.67 | 5708.2205
2024-11-20 | A | NA | 4934.57 | 5674.7555
2024-11-20 | A | NA | 4898.47 | 5633.2405
2024-11-20 | A | NA | 4834.41 | 5559.5715
2024-11-20 | A | NA | 4830.98 | 5555.627
2024-11-20 | A | NA | 4996.66 | 5746.159
2024-11-20 | A | NA | 4889.77 | 5623.2355
2024-11-20 | A | NA | 4926.26 | 5665.199
2024-11-20 | A | NA | 4956.13 | 5699.5495
2024-11-20 | A | NA | 4781.46 | 5498.679
2024-11-20 | A | NA | 4928.79 | 5668.1085
2024-11-20 | A | NA | 4902.13 | 5637.4495
2024-11-20 | A | NA | 4842.66 | 5569.059
2024-11-20 | A | NA | 4955.13 | 5698.3995
2024-11-20 | A | NA | 4886.71 | 5619.7165
2024-11-20 | A | NA | 4963.6 | 5708.14
2024-11-21 | A | 10.25 | 7283.39
2024-11-21 | A | 10.25 | 7319.79
2024-11-21 | A | 10.25 | 7268.78
2024-11-21 | A | 10.25 | 7359.06
2024-11-21 | A | 10.25 | 7333.1
2024-11-21 | A | 10.25 | 7349.73
2024-11-21 | A | 10.25 | 7258.02
2024-11-21 | A | 10.25 | 7338.11
2024-11-21 | A | 10.25 | 7344.12
2024-11-21 | A | 10.25 | 7337.17
2024-11-21 | A | 10.25 | 7260.97
2024-11-21 | A | 10.25 | 7421.4
2024-11-21 | A | 10.25 | 7307.77
2024-11-21 | A | 10.25 | 7399.19
2024-11-21 | A | 10.25 | 7253.94
2024-11-21 | A | 10.25 | 7235.07
2024-11-21 | A | 10.25 | 7389.93
2024-11-21 | A | 10.25 | 7422.77
2024-11-21 | A | 10.25 | 7362.71
2024-11-21 | A | 10.25 | 7263.59
2024-11-21 | A | 8.2 | 6224.12
2024-11-21 | A | 8.2 | 6243.23
2024-11-21 | A | 8.2 | 6298.12
2024-11-21 | A | 8.2 | 6309.09
2024-11-21 | A | 8.2 | 6237.03
2024-11-21 | A | 8.2 | 6227.93
2024-11-21 | A | 8.2 | 6315.43
2024-11-21 | A | 8.2 | 6288.01
2024-11-21 | A | 8.2 | 6216.61
2024-11-21 | A | 8.2 | 6282.93
2024-11-21 | A | 8.2 | 6317.71
2024-11-21 | A | 8.2 | 6321.91
2024-11-21 | A | 8.2 | 6361.26
2024-11-21 | A | 8.2 | 6280.1
2024-11-21 | A | 8.2 | 6299.99
2024-11-21 | A | 8.2 | 6147.24
2024-11-21 | A | 8.2 | 6334.19
2024-11-21 | A | 8.2 | 6328.35
2024-11-21 | A | 8.2 | 6257.51
2024-11-21 | A | 8.2 | 6261.85
2024-11-21 | A | 6.15 | 4923.07
2024-11-21 | A | 6.15 | 4906.29
2024-11-21 | A | 6.15 | 4933.93
2024-11-21 | A | 6.15 | 4907.76
2024-11-21 | A | 6.15 | 4914.8
2024-11-21 | A | 6.15 | 4869.38
2024-11-21 | A | 6.15 | 4902.83
2024-11-21 | A | 6.15 | 4903.01
2024-11-21 | A | 6.15 | 4905.06
2024-11-21 | A | 6.15 | 4908.53
2024-11-21 | A | 6.15 | 4925.63
2024-11-21 | A | 6.15 | 4961.71
2024-11-21 | A | 6.15 | 4925.48
2024-11-21 | A | 6.15 | 4896.59
2024-11-21 | A | 6.15 | 4982.95
2024-11-21 | A | 6.15 | 4948.48
2024-11-21 | A | 6.15 | 4930.23
2024-11-21 | A | 6.15 | 4907.92
2024-11-21 | A | 6.15 | 4873.34
2024-11-21 | A | 6.15 | 4929.79
2024-11-21 | A | 4.1 | 3360.73
2024-11-21 | A | 4.1 | 3351.99
2024-11-21 | A | 4.1 | 3356.36
2024-11-21 | A | 4.1 | 3305.95
2024-11-21 | A | 4.1 | 3289.39
2024-11-21 | A | 4.1 | 3374.49
2024-11-21 | A | 4.1 | 3275.4
2024-11-21 | A | 4.1 | 3317.94
2024-11-21 | A | 4.1 | 3323.4
2024-11-21 | A | 4.1 | 3275.01
2024-11-21 | A | 4.1 | 3377.43
2024-11-21 | A | 4.1 | 3264.84
2024-11-21 | A | 4.1 | 3323.27
2024-11-21 | A | 4.1 | 3317.29
2024-11-21 | A | 4.1 | 3278.54
2024-11-21 | A | 4.1 | 3388.98
2024-11-21 | A | 4.1 | 3331.69
2024-11-21 | A | 4.1 | 3314.77
2024-11-21 | A | 4.1 | 3317.67
2024-11-21 | A | 4.1 | 3357.49
2024-11-21 | A | 2.05 | 1711.53
2024-11-21 | A | 2.05 | 1699.42
2024-11-21 | A | 2.05 | 1717.94
2024-11-21 | A | 2.05 | 1706.42
2024-11-21 | A | 2.05 | 1723.83
2024-11-21 | A | 2.05 | 1740.29
2024-11-21 | A | 2.05 | 1721.61
2024-11-21 | A | 2.05 | 1764.22
2024-11-21 | A | 2.05 | 1773.49
2024-11-21 | A | 2.05 | 1782.08
2024-11-21 | A | 2.05 | 1736.54
2024-11-21 | A | 2.05 | 1743.03
2024-11-21 | A | 2.05 | 1789.61
2024-11-21 | A | 2.05 | 1738.82
2024-11-21 | A | 2.05 | 1774.88
2024-11-21 | A | 2.05 | 1762.35
2024-11-21 | A | 2.05 | 1693.59
2024-11-21 | A | 2.05 | 1755.74
2024-11-21 | A | 2.05 | 1732.3
2024-11-21 | A | 2.05 | 1774.02
2024-11-21 | A | 0 | 122.01
2024-11-21 | A | 0 | 116.21
2024-11-21 | A | 0 | 116.36
2024-11-21 | A | 0 | 116.73
2024-11-21 | A | 0 | 119.6
2024-11-21 | A | 0 | 115.57
2024-11-21 | A | 0 | 120.07
2024-11-21 | A | 0 | 117.5
2024-11-21 | A | 0 | 121.25
2024-11-21 | A | 0 | 126.51
2024-11-21 | A | 0 | 118.03
2024-11-21 | A | 0 | 115.85
2024-11-21 | A | 0 | 113.21
2024-11-21 | A | 0 | 118.43
2024-11-21 | A | 0 | 116.69
2024-11-21 | A | 0 | 113.3
2024-11-21 | A | 0 | 120.57
2024-11-21 | A | 0 | 110.72
2024-11-21 | A | 0 | 115.26
2024-11-21 | A | 0 | 119.51
2024-11-21 | A | NA | 4879.96 | 5611.954
2024-11-21 | A | NA | 4829.65 | 5554.0975
2024-11-21 | A | NA | 4822.11 | 5545.4265
2024-11-21 | A | NA | 4891.05 | 5624.7075
2024-11-21 | A | NA | 4824.61 | 5548.3015
2024-11-21 | A | NA | 4935.84 | 5676.216
2024-11-21 | A | NA | 4816.2 | 5538.63
2024-11-21 | A | NA | 4879.91 | 5611.8965
2024-11-21 | A | NA | 4901.17 | 5636.3455
2024-11-21 | A | NA | 4902.98 | 5638.427
2024-11-21 | A | NA | 4892.41 | 5626.2715
2024-11-21 | A | NA | 4816.27 | 5538.7105
2024-11-21 | A | NA | 4916.55 | 5654.0325
2024-11-21 | A | NA | 4915.49 | 5652.8135
2024-11-21 | A | NA | 4899.94 | 5634.931
2024-11-21 | A | NA | 4850.66 | 5578.259
2024-11-21 | A | NA | 4916.45 | 5653.9175
2024-11-21 | A | NA | 4964.02 | 5708.623
2024-11-21 | A | NA | 4921.17 | 5659.3455
2024-11-21 | A | NA | 4915.91 | 5653.2965
2024-11-22 | A | 10.25 | 7287.97
2024-11-22 | A | 10.25 | 7441.28
2024-11-22 | A | 10.25 | 7304.32
2024-11-22 | A | 10.25 | 7318.78
2024-11-22 | A | 10.25 | 7368.99
2024-11-22 | A | 10.25 | 7384.35
2024-11-22 | A | 10.25 | 7332.89
2024-11-22 | A | 10.25 | 7331.8
2024-11-22 | A | 10.25 | 7200.54
2024-11-22 | A | 10.25 | 7396.95
2024-11-22 | A | 10.25 | 7215.41
2024-11-22 | A | 10.25 | 7341.38
2024-11-22 | A | 10.25 | 7339.51
2024-11-22 | A | 10.25 | 7342.66
2024-11-22 | A | 10.25 | 7281.71
2024-11-22 | A | 10.25 | 7233.59
2024-11-22 | A | 10.25 | 7339.94
2024-11-22 | A | 10.25 | 7320.66
2024-11-22 | A | 10.25 | 7367.88
2024-11-22 | A | 10.25 | 7351.81
2024-11-22 | A | 8.2 | 6313.98
2024-11-22 | A | 8.2 | 6292.86
2024-11-22 | A | 8.2 | 6274.07
2024-11-22 | A | 8.2 | 6325.05
2024-11-22 | A | 8.2 | 6257.35
2024-11-22 | A | 8.2 | 6225.19
2024-11-22 | A | 8.2 | 6330.52
2024-11-22 | A | 8.2 | 6273.08
2024-11-22 | A | 8.2 | 6272.54
2024-11-22 | A | 8.2 | 6268.81
2024-11-22 | A | 8.2 | 6361.44
2024-11-22 | A | 8.2 | 6344.47
2024-11-22 | A | 8.2 | 6312.69
2024-11-22 | A | 8.2 | 6301.75
2024-11-22 | A | 8.2 | 6227.56
2024-11-22 | A | 8.2 | 6291.41
2024-11-22 | A | 8.2 | 6252.41
2024-11-22 | A | 8.2 | 6339.11
2024-11-22 | A | 8.2 | 6301.11
2024-11-22 | A | 8.2 | 6340.36
2024-11-22 | A | 6.15 | 4931.64
2024-11-22 | A | 6.15 | 4921.58
2024-11-22 | A | 6.15 | 4938.21
2024-11-22 | A | 6.15 | 4933.76
2024-11-22 | A | 6.15 | 4940.13
2024-11-22 | A | 6.15 | 4883.44
2024-11-22 | A | 6.15 | 4930.84
2024-11-22 | A | 6.15 | 4954.49
2024-11-22 | A | 6.15 | 4939.44
2024-11-22 | A | 6.15 | 4852.74
2024-11-22 | A | 6.15 | 4997.28
2024-11-22 | A | 6.15 | 4933.19
2024-11-22 | A | 6.15 | 4901.48
2024-11-22 | A | 6.15 | 4868.89
2024-11-22 | A | 6.15 | 4948.25
2024-11-22 | A | 6.15 | 4873.84
2024-11-22 | A | 6.15 | 4887.68
2024-11-22 | A | 6.15 | 4941.94
2024-11-22 | A | 6.15 | 4868.23
2024-11-22 | A | 6.15 | 4881.32
2024-11-22 | A | 4.1 | 3329.52
2024-11-22 | A | 4.1 | 3356.68
2024-11-22 | A | 4.1 | 3328.53
2024-11-22 | A | 4.1 | 3384
2024-11-22 | A | 4.1 | 3272.4
2024-11-22 | A | 4.1 | 3376.89
2024-11-22 | A | 4.1 | 3362.71
2024-11-22 | A | 4.1 | 3296.87
2024-11-22 | A | 4.1 | 3370.14
2024-11-22 | A | 4.1 | 3379.96
2024-11-22 | A | 4.1 | 3348.5
2024-11-22 | A | 4.1 | 3325.46
2024-11-22 | A | 4.1 | 3283.19
2024-11-22 | A | 4.1 | 3314.78
2024-11-22 | A | 4.1 | 3403.94
2024-11-22 | A | 4.1 | 3409.27
2024-11-22 | A | 4.1 | 3270.09
2024-11-22 | A | 4.1 | 3303.85
2024-11-22 | A | 4.1 | 3408.57
2024-11-22 | A | 4.1 | 3296.97
2024-11-22 | A | 2.05 | 1732.41
2024-11-22 | A | 2.05 | 1747.46
2024-11-22 | A | 2.05 | 1716.26
2024-11-22 | A | 2.05 | 1744.52
2024-11-22 | A | 2.05 | 1744.93
2024-11-22 | A | 2.05 | 1713.5
2024-11-22 | A | 2.05 | 1705.23
2024-11-22 | A | 2.05 | 1758.2
2024-11-22 | A | 2.05 | 1763.01
2024-11-22 | A | 2.05 | 1745.15
2024-11-22 | A | 2.05 | 1740.67
2024-11-22 | A | 2.05 | 1766.91
2024-11-22 | A | 2.05 | 1760.22
2024-11-22 | A | 2.05 | 1685.29
2024-11-22 | A | 2.05 | 1777.13
2024-11-22 | A | 2.05 | 1748.23
2024-11-22 | A | 2.05 | 1794.91
2024-11-22 | A | 2.05 | 1759
2024-11-22 | A | 2.05 | 1726.85
2024-11-22 | A | 2.05 | 1724.61
2024-11-22 | A | 0 | 116.34
2024-11-22 | A | 0 | 119.7
2024-11-22 | A | 0 | 117.51
2024-11-22 | A | 0 | 113.7
2024-11-22 | A | 0 | 115.28
2024-11-22 | A | 0 | 117.87
2024-11-22 | A | 0 | 113
2024-11-22 | A | 0 | 120.51
2024-11-22 | A | 0 | 116.63
2024-11-22 | A | 0 | 130.37
2024-11-22 | A | 0 | 121.31
2024-11-22 | A | 0 | 120.38
2024-11-22 | A | 0 | 114.95
2024-11-22 | A | 0 | 111.28
2024-11-22 | A | 0 | 114.82
2024-11-22 | A | 0 | 118.36
2024-11-22 | A | 0 | 119.48
2024-11-22 | A | 0 | 121.07
2024-11-22 | A | 0 | 115.9
2024-11-22 | A | 0 | 122.34
2024-11-22 | A | NA | 4863.19 | 5592.6685
2024-11-22 | A | NA | 4908.39 | 5644.6485
2024-11-22 | A | NA | 4899.3 | 5634.195
2024-11-22 | A | NA | 4894.71 | 5628.9165
2024-11-22 | A | NA | 4889.27 | 5622.6605
2024-11-22 | A | NA | 4943.76 | 5685.324
2024-11-22 | A | NA | 4958.62 | 5702.413
2024-11-22 | A | NA | 4843.81 | 5570.3815
2024-11-22 | A | NA | 4964.68 | 5709.382
2024-11-22 | A | NA | 4914.4 | 5651.56
2024-11-22 | A | NA | 4874.59 | 5605.7785
2024-11-22 | A | NA | 4806.88 | 5527.912
2024-11-22 | A | NA | 4991.87 | 5740.6505
2024-11-22 | A | NA | 4902.4 | 5637.76
2024-11-22 | A | NA | 4940.03 | 5681.0345
2024-11-22 | A | NA | 4817.74 | 5540.401
2024-11-22 | A | NA | 4876.15 | 5607.5725
2024-11-22 | A | NA | 5003.83 | 5754.4045
2024-11-22 | A | NA | 4850.8 | 5578.42
2024-11-22 | A | NA | 4805.26 | 5526.049
2024-11-24 | A | 10.25 | 7296.15
2024-11-24 | A | 10.25 | 7327.25
2024-11-24 | A | 10.25 | 7305.73
2024-11-24 | A | 10.25 | 7414.78
2024-11-24 | A | 10.25 | 7305.81
2024-11-24 | A | 10.25 | 7341.38
2024-11-24 | A | 10.25 | 7368.15
2024-11-24 | A | 10.25 | 7383.34
2024-11-24 | A | 10.25 | 7231.62
2024-11-24 | A | 10.25 | 7250.25
2024-11-24 | A | 10.25 | 7365.22
2024-11-24 | A | 10.25 | 7247.66
2024-11-24 | A | 10.25 | 7340.15
2024-11-24 | A | 10.25 | 7301.25
2024-11-24 | A | 10.25 | 7350.78
2024-11-24 | A | 10.25 | 7331.91
2024-11-24 | A | 10.25 | 7314.41
2024-11-24 | A | 10.25 | 7345.29
2024-11-24 | A | 10.25 | 7372.13
2024-11-24 | A | 10.25 | 7306.83
2024-11-24 | A | 8.2 | 6316.88
2024-11-24 | A | 8.2 | 6334.99
2024-11-24 | A | 8.2 | 6297.27
2024-11-24 | A | 8.2 | 6344.57
2024-11-24 | A | 8.2 | 6395.48
2024-11-24 | A | 8.2 | 6353.79
2024-11-24 | A | 8.2 | 6302.38
2024-11-24 | A | 8.2 | 6317.54
2024-11-24 | A | 8.2 | 6198.23
2024-11-24 | A | 8.2 | 6316.87
2024-11-24 | A | 8.2 | 6237.89
2024-11-24 | A | 8.2 | 6300.84
2024-11-24 | A | 8.2 | 6271.01
2024-11-24 | A | 8.2 | 6274.06
2024-11-24 | A | 8.2 | 6360.27
2024-11-24 | A | 8.2 | 6198.37
2024-11-24 | A | 8.2 | 6280.27
2024-11-24 | A | 8.2 | 6190.65
2024-11-24 | A | 8.2 | 6275.93
2024-11-24 | A | 8.2 | 6275.75
2024-11-24 | A | 6.15 | 4941.11
2024-11-24 | A | 6.15 | 4922.94
2024-11-24 | A | 6.15 | 4929.93
2024-11-24 | A | 6.15 | 4932.96
2024-11-24 | A | 6.15 | 4913.12
2024-11-24 | A | 6.15 | 4892.08
2024-11-24 | A | 6.15 | 4912.77
2024-11-24 | A | 6.15 | 4899.51
2024-11-24 | A | 6.15 | 4905.37
2024-11-24 | A | 6.15 | 4940.36
2024-11-24 | A | 6.15 | 4897.3
2024-11-24 | A | 6.15 | 4968.01
2024-11-24 | A | 6.15 | 4850.64
2024-11-24 | A | 6.15 | 4945.91
2024-11-24 | A | 6.15 | 4936.41
2024-11-24 | A | 6.15 | 4962.33
2024-11-24 | A | 6.15 | 4946.79
2024-11-24 | A | 6.15 | 4934.86
2024-11-24 | A | 6.15 | 4876.31
2024-11-24 | A | 6.15 | 4917.85
2024-11-24 | A | 4.1 | 3341.25
2024-11-24 | A | 4.1 | 3301.13
2024-11-24 | A | 4.1 | 3293.27
2024-11-24 | A | 4.1 | 3362.83
2024-11-24 | A | 4.1 | 3344.41
2024-11-24 | A | 4.1 | 3322.88
2024-11-24 | A | 4.1 | 3369.7
2024-11-24 | A | 4.1 | 3289.9
2024-11-24 | A | 4.1 | 3328.75
2024-11-24 | A | 4.1 | 3344.65
2024-11-24 | A | 4.1 | 3344.12
2024-11-24 | A | 4.1 | 3347.96
2024-11-24 | A | 4.1 | 3272.72
2024-11-24 | A | 4.1 | 3388.09
2024-11-24 | A | 4.1 | 3332.51
2024-11-24 | A | 4.1 | 3317.41
2024-11-24 | A | 4.1 | 3336.59
2024-11-24 | A | 4.1 | 3334.41
2024-11-24 | A | 4.1 | 3360.85
2024-11-24 | A | 4.1 | 3315.68
2024-11-24 | A | 2.05 | 1708.83
2024-11-24 | A | 2.05 | 1778.74
2024-11-24 | A | 2.05 | 1725.18
2024-11-24 | A | 2.05 | 1747.52
2024-11-24 | A | 2.05 | 1710.56
2024-11-24 | A | 2.05 | 1753.04
2024-11-24 | A | 2.05 | 1722.64
2024-11-24 | A | 2.05 | 1784.93
2024-11-24 | A | 2.05 | 1729.67
2024-11-24 | A | 2.05 | 1740.84
2024-11-24 | A | 2.05 | 1767.95
2024-11-24 | A | 2.05 | 1723.94
2024-11-24 | A | 2.05 | 1764.11
2024-11-24 | A | 2.05 | 1798.37
2024-11-24 | A | 2.05 | 1742.56
2024-11-24 | A | 2.05 | 1782.69
2024-11-24 | A | 2.05 | 1739.71
2024-11-24 | A | 2.05 | 1772.63
2024-11-24 | A | 2.05 | 1773.19
2024-11-24 | A | 2.05 | 1773.07
2024-11-24 | A | 0 | 121.2
2024-11-24 | A | 0 | 121.97
2024-11-24 | A | 0 | 121.56
2024-11-24 | A | 0 | 122.1
2024-11-24 | A | 0 | 110.5
2024-11-24 | A | 0 | 119.85
2024-11-24 | A | 0 | 113.36
2024-11-24 | A | 0 | 121.55
2024-11-24 | A | 0 | 123.06
2024-11-24 | A | 0 | 113
2024-11-24 | A | 0 | 114.09
2024-11-24 | A | 0 | 126.29
2024-11-24 | A | 0 | 117.46
2024-11-24 | A | 0 | 114.92
2024-11-24 | A | 0 | 120.96
2024-11-24 | A | 0 | 112.78
2024-11-24 | A | 0 | 116.78
2024-11-24 | A | 0 | 114.33
2024-11-24 | A | 0 | 120.84
2024-11-24 | A | 0 | 123.21
2024-11-24 | A | NA | 4877.25 | 5608.8375
2024-11-24 | A | NA | 4860.54 | 5589.621
2024-11-24 | A | NA | 4917.51 | 5655.1365
2024-11-24 | A | NA | 4888.94 | 5622.281
2024-11-24 | A | NA | 4802.76 | 5523.174
2024-11-24 | A | NA | 4888.8 | 5622.12
2024-11-24 | A | NA | 4911.71 | 5648.4665
2024-11-24 | A | NA | 4889.48 | 5622.902
2024-11-24 | A | NA | 4931.99 | 5671.7885
2024-11-24 | A | NA | 4853.99 | 5582.0885
2024-11-24 | A | NA | 4833.37 | 5558.3755
2024-11-24 | A | NA | 4889.49 | 5622.9135
2024-11-24 | A | NA | 4867.83 | 5598.0045
2024-11-24 | A | NA | 4831.95 | 5556.7425
2024-11-24 | A | NA | 4899.53 | 5634.4595
2024-11-24 | A | NA | 4919 | 5656.85
2024-11-24 | A | NA | 4875.03 | 5606.2845
2024-11-24 | A | NA | 4881.38 | 5613.587
2024-11-24 | A | NA | 4782.75 | 5500.1625
2024-11-24 | A | NA | 4885.11 | 5617.8765
2024-11-25 | A | 10.25 | 7328.89
2024-11-25 | A | 10.25 | 7339.99
2024-11-25 | A | 10.25 | 7327.73
2024-11-25 | A | 10.25 | 7422.05
2024-11-25 | A | 10.25 | 7306.46
2024-11-25 | A | 10.25 | 7313.81
2024-11-25 | A | 10.25 | 7343.7
2024-11-25 | A | 10.25 | 7383.08
2024-11-25 | A | 10.25 | 7326.65
2024-11-25 | A | 10.25 | 7453.93
2024-11-25 | A | 10.25 | 7237.41
2024-11-25 | A | 10.25 | 7316.33
2024-11-25 | A | 10.25 | 7331.97
2024-11-25 | A | 10.25 | 7367
2024-11-25 | A | 10.25 | 7355.75
2024-11-25 | A | 10.25 | 7217.07
2024-11-25 | A | 10.25 | 7332.73
2024-11-25 | A | 10.25 | 7284.65
2024-11-25 | A | 10.25 | 7312.64
2024-11-25 | A | 10.25 | 7309.87
2024-11-25 | A | 8.2 | 6268.65
2024-11-25 | A | 8.2 | 6313.33
2024-11-25 | A | 8.2 | 6246.29
2024-11-25 | A | 8.2 | 6318.47
2024-11-25 | A | 8.2 | 6290.45
2024-11-25 | A | 8.2 | 6244.12
2024-11-25 | A | 8.2 | 6277.6
2024-11-25 | A | 8.2 | 6313.86
2024-11-25 | A | 8.2 | 6353.54
2024-11-25 | A | 8.2 | 6344.41
2024-11-25 | A | 8.2 | 6259.08
2024-11-25 | A | 8.2 | 6338.75
2024-11-25 | A | 8.2 | 6287.54
2024-11-25 | A | 8.2 | 6304.07
2024-11-25 | A | 8.2 | 6242.12
2024-11-25 | A | 8.2 | 6290
2024-11-25 | A | 8.2 | 6352.21
2024-11-25 | A | 8.2 | 6280.36
2024-11-25 | A | 8.2 | 6210.35
2024-11-25 | A | 8.2 | 6250.94
2024-11-25 | A | 6.15 | 4846.25
2024-11-25 | A | 6.15 | 4951.33
2024-11-25 | A | 6.15 | 4959.16
2024-11-25 | A | 6.15 | 4932.11
2024-11-25 | A | 6.15 | 4921.33
2024-11-25 | A | 6.15 | 4860.08
2024-11-25 | A | 6.15 | 4931.86
2024-11-25 | A | 6.15 | 4930
2024-11-25 | A | 6.15 | 4911.74
2024-11-25 | A | 6.15 | 4871.43
2024-11-25 | A | 6.15 | 4874.94
2024-11-25 | A | 6.15 | 4858.15
2024-11-25 | A | 6.15 | 4903.64
2024-11-25 | A | 6.15 | 4883.36
2024-11-25 | A | 6.15 | 4946.41
2024-11-25 | A | 6.15 | 4863.48
2024-11-25 | A | 6.15 | 4947.99
2024-11-25 | A | 6.15 | 4931.37
2024-11-25 | A | 6.15 | 4839.52
2024-11-25 | A | 6.15 | 4875.25
2024-11-25 | A | 4.1 | 3333.82
2024-11-25 | A | 4.1 | 3329.63
2024-11-25 | A | 4.1 | 3323.06
2024-11-25 | A | 4.1 | 3370.15
2024-11-25 | A | 4.1 | 3337.12
2024-11-25 | A | 4.1 | 3294.6
2024-11-25 | A | 4.1 | 3342.15
2024-11-25 | A | 4.1 | 3287.49
2024-11-25 | A | 4.1 | 3269.27
2024-11-25 | A | 4.1 | 3291.73
2024-11-25 | A | 4.1 | 3294.98
2024-11-25 | A | 4.1 | 3320.57
2024-11-25 | A | 4.1 | 3404.86
2024-11-25 | A | 4.1 | 3296.89
2024-11-25 | A | 4.1 | 3274.68
2024-11-25 | A | 4.1 | 3268.76
2024-11-25 | A | 4.1 | 3313.96
2024-11-25 | A | 4.1 | 3350.96
2024-11-25 | A | 4.1 | 3323.51
2024-11-25 | A | 4.1 | 3305.39
2024-11-25 | A | 2.05 | 1724.79
2024-11-25 | A | 2.05 | 1737.33
2024-11-25 | A | 2.05 | 1734.09
2024-11-25 | A | 2.05 | 1769.44
2024-11-25 | A | 2.05 | 1674.29
2024-11-25 | A | 2.05 | 1726.27
2024-11-25 | A | 2.05 | 1753.09
2024-11-25 | A | 2.05 | 1746.79
2024-11-25 | A | 2.05 | 1767.28
2024-11-25 | A | 2.05 | 1722.68
2024-11-25 | A | 2.05 | 1799.65
2024-11-25 | A | 2.05 | 1709.16
2024-11-25 | A | 2.05 | 1770.66
2024-11-25 | A | 2.05 | 1753.73
2024-11-25 | A | 2.05 | 1741.55
2024-11-25 | A | 2.05 | 1693.83
2024-11-25 | A | 2.05 | 1750.66
2024-11-25 | A | 2.05 | 1767
2024-11-25 | A | 2.05 | 1745.06
2024-11-25 | A | 2.05 | 1756.99
2024-11-25 | A | 0 | 120.2
2024-11-25 | A | 0 | 119.55
2024-11-25 | A | 0 | 118.85
2024-11-25 | A | 0 | 109.7
2024-11-25 | A | 0 | 118.52
2024-11-25 | A | 0 | 109.32
2024-11-25 | A | 0 | 115.98
2024-11-25 | A | 0 | 114.77
2024-11-25 | A | 0 | 122.38
2024-11-25 | A | 0 | 120.57
2024-11-25 | A | 0 | 112.94
2024-11-25 | A | 0 | 116.8
2024-11-25 | A | 0 | 118.07
2024-11-25 | A | 0 | 115.73
2024-11-25 | A | 0 | 115.19
2024-11-25 | A | 0 | 114.91
2024-11-25 | A | 0 | 121.1
2024-11-25 | A | 0 | 120.66
2024-11-25 | A | 0 | 119.32
2024-11-25 | A | 0 | 107.58
2024-11-25 | A | NA | 4912.34 | 5649.191
2024-11-25 | A | NA | 4823.03 | 5546.4845
2024-11-25 | A | NA | 4878.78 | 5610.597
2024-11-25 | A | NA | 4876.97 | 5608.5155
2024-11-25 | A | NA | 4846.94 | 5573.981
2024-11-25 | A | NA | 4935.9 | 5676.285
2024-11-25 | A | NA | 4896.99 | 5631.5385
2024-11-25 | A | NA | 4892.09 | 5625.9035
2024-11-25 | A | NA | 4983.12 | 5730.588
2024-11-25 | A | NA | 4899.82 | 5634.793
2024-11-25 | A | NA | 4894.58 | 5628.767
2024-11-25 | A | NA | 4877.4 | 5609.01
2024-11-25 | A | NA | 4899.02 | 5633.873
2024-11-25 | A | NA | 4900.92 | 5636.058
2024-11-25 | A | NA | 4844.8 | 5571.52
2024-11-25 | A | NA | 4877.91 | 5609.5965
2024-11-25 | A | NA | 4855.99 | 5584.3885
2024-11-25 | A | NA | 4855.32 | 5583.618
2024-11-25 | A | NA | 4897.59 | 5632.2285
2024-11-25 | A | NA | 4927.99 | 5667.1885
2024-11-26 | B | 10.25 | 7307.83
2024-11-26 | B | 10.25 | 7295.06
2024-11-26 | B | 10.25 | 7433.81
2024-11-26 | B | 10.25 | 7301.87
2024-11-26 | B | 10.25 | 7355.29
2024-11-26 | B | 10.25 | 7305.76
2024-11-26 | B | 10.25 | 7335.95
2024-11-26 | B | 10.25 | 7352.24
2024-11-26 | B | 10.25 | 7395.85
2024-11-26 | B | 10.25 | 7377.37
2024-11-26 | B | 10.25 | 7408.29
2024-11-26 | B | 10.25 | 7316.95
2024-11-26 | B | 10.25 | 7325.36
2024-11-26 | B | 10.25 | 7279.03
2024-11-26 | B | 10.25 | 7300.26
2024-11-26 | B | 10.25 | 7292.44
2024-11-26 | B | 10.25 | 7282.32
2024-11-26 | B | 10.25 | 7329.51
2024-11-26 | B | 10.25 | 7284.62
2024-11-26 | B | 10.25 | 7294.98
2024-11-26 | B | 8.2 | 6326.55
2024-11-26 | B | 8.2 | 6228.3
2024-11-26 | B | 8.2 | 6334.19
2024-11-26 | B | 8.2 | 6354.12
2024-11-26 | B | 8.2 | 6308.92
2024-11-26 | B | 8.2 | 6288.78
2024-11-26 | B | 8.2 | 6272.73
2024-11-26 | B | 8.2 | 6303.26
2024-11-26 | B | 8.2 | 6316.98
2024-11-26 | B | 8.2 | 6325.86
2024-11-26 | B | 8.2 | 6189.24
2024-11-26 | B | 8.2 | 6270.76
2024-11-26 | B | 8.2 | 6319.48
2024-11-26 | B | 8.2 | 6204.29
2024-11-26 | B | 8.2 | 6324.61
2024-11-26 | B | 8.2 | 6261.87
2024-11-26 | B | 8.2 | 6302.53
2024-11-26 | B | 8.2 | 6352.53
2024-11-26 | B | 8.2 | 6337.49
2024-11-26 | B | 8.2 | 6251.33
2024-11-26 | B | 6.15 | 4975.41
2024-11-26 | B | 6.15 | 4881.48
2024-11-26 | B | 6.15 | 4899.3
2024-11-26 | B | 6.15 | 4898.9
2024-11-26 | B | 6.15 | 4867
2024-11-26 | B | 6.15 | 4909.86
2024-11-26 | B | 6.15 | 4922.89
2024-11-26 | B | 6.15 | 4953.44
2024-11-26 | B | 6.15 | 4974.6
2024-11-26 | B | 6.15 | 4953.49
2024-11-26 | B | 6.15 | 5005.55
2024-11-26 | B | 6.15 | 4912.56
2024-11-26 | B | 6.15 | 4897.92
2024-11-26 | B | 6.15 | 4865.91
2024-11-26 | B | 6.15 | 4958.9
2024-11-26 | B | 6.15 | 4992.27
2024-11-26 | B | 6.15 | 4902.62
2024-11-26 | B | 6.15 | 4879.57
2024-11-26 | B | 6.15 | 4920.18
2024-11-26 | B | 6.15 | 4948.69
2024-11-26 | B | 4.1 | 3352.68
2024-11-26 | B | 4.1 | 3355.56
2024-11-26 | B | 4.1 | 3378.31
2024-11-26 | B | 4.1 | 3302.09
2024-11-26 | B | 4.1 | 3311
2024-11-26 | B | 4.1 | 3368.79
2024-11-26 | B | 4.1 | 3293.01
2024-11-26 | B | 4.1 | 3350.51
2024-11-26 | B | 4.1 | 3354.82
2024-11-26 | B | 4.1 | 3397.99
2024-11-26 | B | 4.1 | 3346.17
2024-11-26 | B | 4.1 | 3372.76
2024-11-26 | B | 4.1 | 3365.2
2024-11-26 | B | 4.1 | 3287.19
2024-11-26 | B | 4.1 | 3355.24
2024-11-26 | B | 4.1 | 3340.35
2024-11-26 | B | 4.1 | 3340
2024-11-26 | B | 4.1 | 3344.44
2024-11-26 | B | 4.1 | 3359.71
2024-11-26 | B | 4.1 | 3409.79
2024-11-26 | B | 2.05 | 1682.88
2024-11-26 | B | 2.05 | 1767.28
2024-11-26 | B | 2.05 | 1720.5
2024-11-26 | B | 2.05 | 1768.22
2024-11-26 | B | 2.05 | 1728.95
2024-11-26 | B | 2.05 | 1702.88
2024-11-26 | B | 2.05 | 1721.03
2024-11-26 | B | 2.05 | 1726.57
2024-11-26 | B | 2.05 | 1704.51
2024-11-26 | B | 2.05 | 1742.58
2024-11-26 | B | 2.05 | 1740.21
2024-11-26 | B | 2.05 | 1780.13
2024-11-26 | B | 2.05 | 1702.43
2024-11-26 | B | 2.05 | 1691.82
2024-11-26 | B | 2.05 | 1737.53
2024-11-26 | B | 2.05 | 1731.63
2024-11-26 | B | 2.05 | 1720.32
2024-11-26 | B | 2.05 | 1714.45
2024-11-26 | B | 2.05 | 1689.72
2024-11-26 | B | 2.05 | 1718.13
2024-11-26 | B | 0 | 113.63
2024-11-26 | B | 0 | 117.34
2024-11-26 | B | 0 | 120.68
2024-11-26 | B | 0 | 106
2024-11-26 | B | 0 | 114.84
2024-11-26 | B | 0 | 115.98
2024-11-26 | B | 0 | 115.44
2024-11-26 | B | 0 | 117.15
2024-11-26 | B | 0 | 114.79
2024-11-26 | B | 0 | 116.42
2024-11-26 | B | 0 | 115.96
2024-11-26 | B | 0 | 122.12
2024-11-26 | B | 0 | 121.93
2024-11-26 | B | 0 | 113.19
2024-11-26 | B | 0 | 117.83
2024-11-26 | B | 0 | 116.87
2024-11-26 | B | 0 | 115.21
2024-11-26 | B | 0 | 117.68
2024-11-26 | B | 0 | 112.32
2024-11-26 | B | 0 | 118.06
2024-11-26 | B | NA | 4982.26 | 5729.599
2024-11-26 | B | NA | 4885.49 | 5618.3135
2024-11-26 | B | NA | 4882.78 | 5615.197
2024-11-26 | B | NA | 4885.19 | 5617.9685
2024-11-26 | B | NA | 4883.08 | 5615.542
2024-11-26 | B | NA | 4942.24 | 5683.576
2024-11-26 | B | NA | 4970 | 5715.5
2024-11-26 | B | NA | 4904.58 | 5640.267
2024-11-26 | B | NA | 4798.85 | 5518.6775
2024-11-26 | B | NA | 4867.83 | 5598.0045
2024-11-26 | B | NA | 4904.39 | 5640.0485
2024-11-26 | B | NA | 4843.88 | 5570.462
2024-11-26 | B | NA | 4816.06 | 5538.469
2024-11-26 | B | NA | 4926.43 | 5665.3945
2024-11-26 | B | NA | 4830.9 | 5555.535
2024-11-26 | B | NA | 4769.39 | 5484.7985
2024-11-26 | B | NA | 4883.26 | 5615.749
2024-11-26 | B | NA | 4933.37 | 5673.3755
2024-11-26 | B | NA | 4900.02 | 5635.023
2024-11-26 | B | NA | 4833.83 | 5558.9045
2024-11-27 | B | 10.25 | 7400.57
2024-11-27 | B | 10.25 | 7375.33
2024-11-27 | B | 10.25 | 7246.6
2024-11-27 | B | 10.25 | 7380.15
2024-11-27 | B | 10.25 | 7372.66
2024-11-27 | B | 10.25 | 7318.6
2024-11-27 | B | 10.25 | 7295.58
2024-11-27 | B | 10.25 | 7395.38
2024-11-27 | B | 10.25 | 7405.11
2024-11-27 | B | 10.25 | 7286.62
2024-11-27 | B | 10.25 | 7199.82
2024-11-27 | B | 10.25 | 7231.97
2024-11-27 | B | 10.25 | 7383.44
2024-11-27 | B | 10.25 | 7322.11
2024-11-27 | B | 10.25 | 7344.19
2024-11-27 | B | 10.25 | 7381.38
2024-11-27 | B | 10.25 | 7311.93
2024-11-27 | B | 10.25 | 7297.96
2024-11-27 | B | 10.25 | 7288.86
2024-11-27 | B | 10.25 | 7299.12
2024-11-27 | B | 8.2 | 6342.33
2024-11-27 | B | 8.2 | 6295.97
2024-11-27 | B | 8.2 | 6324.81
2024-11-27 | B | 8.2 | 6292.42
2024-11-27 | B | 8.2 | 6274.12
2024-11-27 | B | 8.2 | 6225.82
2024-11-27 | B | 8.2 | 6308.58
2024-11-27 | B | 8.2 | 6223.96
2024-11-27 | B | 8.2 | 6272.78
2024-11-27 | B | 8.2 | 6319.62
2024-11-27 | B | 8.2 | 6326.83
2024-11-27 | B | 8.2 | 6261.3
2024-11-27 | B | 8.2 | 6290.47
2024-11-27 | B | 8.2 | 6304.86
2024-11-27 | B | 8.2 | 6295.93
2024-11-27 | B | 8.2 | 6276.77
2024-11-27 | B | 8.2 | 6283.25
2024-11-27 | B | 8.2 | 6295.08
2024-11-27 | B | 8.2 | 6258.62
2024-11-27 | B | 8.2 | 6355.29
2024-11-27 | B | 6.15 | 4913.77
2024-11-27 | B | 6.15 | 4915.69
2024-11-27 | B | 6.15 | 4902.26
2024-11-27 | B | 6.15 | 4847.44
2024-11-27 | B | 6.15 | 4966.78
2024-11-27 | B | 6.15 | 4972.51
2024-11-27 | B | 6.15 | 4893.53
2024-11-27 | B | 6.15 | 4970.76
2024-11-27 | B | 6.15 | 4862.5
2024-11-27 | B | 6.15 | 4868.09
2024-11-27 | B | 6.15 | 4930.86
2024-11-27 | B | 6.15 | 4955.55
2024-11-27 | B | 6.15 | 4938.53
2024-11-27 | B | 6.15 | 4965.25
2024-11-27 | B | 6.15 | 4898.6
2024-11-27 | B | 6.15 | 4914.2
2024-11-27 | B | 6.15 | 4900.81
2024-11-27 | B | 6.15 | 4919.34
2024-11-27 | B | 6.15 | 4913.9
2024-11-27 | B | 6.15 | 4911.43
2024-11-27 | B | 4.1 | 3359.65
2024-11-27 | B | 4.1 | 3294.41
2024-11-27 | B | 4.1 | 3364.94
2024-11-27 | B | 4.1 | 3372.8
2024-11-27 | B | 4.1 | 3312.68
2024-11-27 | B | 4.1 | 3341.2
2024-11-27 | B | 4.1 | 3311.65
2024-11-27 | B | 4.1 | 3336.14
2024-11-27 | B | 4.1 | 3274.44
2024-11-27 | B | 4.1 | 3337.04
2024-11-27 | B | 4.1 | 3295.47
2024-11-27 | B | 4.1 | 3335.46
2024-11-27 | B | 4.1 | 3321.24
2024-11-27 | B | 4.1 | 3307.81
2024-11-27 | B | 4.1 | 3269.12
2024-11-27 | B | 4.1 | 3258.79
2024-11-27 | B | 4.1 | 3256.36
2024-11-27 | B | 4.1 | 3362.38
2024-11-27 | B | 4.1 | 3281.19
2024-11-27 | B | 4.1 | 3357.12
2024-11-27 | B | 2.05 | 1778.81
2024-11-27 | B | 2.05 | 1694.72
2024-11-27 | B | 2.05 | 1757.48
2024-11-27 | B | 2.05 | 1738.73
2024-11-27 | B | 2.05 | 1749.04
2024-11-27 | B | 2.05 | 1777.06
2024-11-27 | B | 2.05 | 1770.13
2024-11-27 | B | 2.05 | 1773.69
2024-11-27 | B | 2.05 | 1753.09
2024-11-27 | B | 2.05 | 1702.05
2024-11-27 | B | 2.05 | 1755.71
2024-11-27 | B | 2.05 | 1767.25
2024-11-27 | B | 2.05 | 1742.65
2024-11-27 | B | 2.05 | 1791.93
2024-11-27 | B | 2.05 | 1747.53
2024-11-27 | B | 2.05 | 1767.48
2024-11-27 | B | 2.05 | 1814.49
2024-11-27 | B | 2.05 | 1778.28
2024-11-27 | B | 2.05 | 1755.71
2024-11-27 | B | 2.05 | 1700.17
2024-11-27 | B | 0 | 123.19
2024-11-27 | B | 0 | 113.78
2024-11-27 | B | 0 | 119.83
2024-11-27 | B | 0 | 110.87
2024-11-27 | B | 0 | 108.56
2024-11-27 | B | 0 | 110.28
2024-11-27 | B | 0 | 128.06
2024-11-27 | B | 0 | 113.25
2024-11-27 | B | 0 | 115.09
2024-11-27 | B | 0 | 117.45
2024-11-27 | B | 0 | 110.49
2024-11-27 | B | 0 | 113.58
2024-11-27 | B | 0 | 116.77
2024-11-27 | B | 0 | 116.09
2024-11-27 | B | 0 | 118.82
2024-11-27 | B | 0 | 109.64
2024-11-27 | B | 0 | 120.49
2024-11-27 | B | 0 | 113.4
2024-11-27 | B | 0 | 115.7
2024-11-27 | B | 0 | 119.03
2024-11-27 | B | NA | 4819.89 | 5542.8735
2024-11-27 | B | NA | 4880.23 | 5612.2645
2024-11-27 | B | NA | 4901.25 | 5636.4375
2024-11-27 | B | NA | 4903.5 | 5639.025
2024-11-27 | B | NA | 4874.38 | 5605.537
2024-11-27 | B | NA | 4918.2 | 5655.93
2024-11-27 | B | NA | 4943.95 | 5685.5425
2024-11-27 | B | NA | 4899.49 | 5634.4135
2024-11-27 | B | NA | 4841.81 | 5568.0815
2024-11-27 | B | NA | 4902.37 | 5637.7255
2024-11-27 | B | NA | 4917.28 | 5654.872
2024-11-27 | B | NA | 4891.82 | 5625.593
2024-11-27 | B | NA | 4880.93 | 5613.0695
2024-11-27 | B | NA | 4820.81 | 5543.9315
2024-11-27 | B | NA | 4868.44 | 5598.706
2024-11-27 | B | NA | 4945.42 | 5687.233
2024-11-27 | B | NA | 4907.5 | 5643.625
2024-11-27 | B | NA | 4878.01 | 5609.7115
2024-11-27 | B | NA | 4840.96 | 5567.104
2024-11-27 | B | NA | 4834.16 | 5559.284
2024-11-28 | B | 10.25 | 7251.45
2024-11-28 | B | 10.25 | 7305.85
2024-11-28 | B | 10.25 | 7349.32
2024-11-28 | B | 10.25 | 7268.95
2024-11-28 | B | 10.25 | 7321.36
2024-11-28 | B | 10.25 | 7372.4
2024-11-28 | B | 10.25 | 7362.34
2024-11-28 | B | 10.25 | 7423.95
2024-11-28 | B | 10.25 | 7366.38
2024-11-28 | B | 10.25 | 7259.19
2024-11-28 | B | 10.25 | 7283.01
2024-11-28 | B | 10.25 | 7272.9
2024-11-28 | B | 10.25 | 7341.55
2024-11-28 | B | 10.25 | 7288.72
2024-11-28 | B | 10.25 | 7377.27
2024-11-28 | B | 10.25 | 7444.69
2024-11-28 | B | 10.25 | 7352.53
2024-11-28 | B | 10.25 | 7209.92
2024-11-28 | B | 10.25 | 7296.78
2024-11-28 | B | 10.25 | 7337.89
2024-11-28 | B | 8.2 | 6256.17
2024-11-28 | B | 8.2 | 6361.79
2024-11-28 | B | 8.2 | 6258.54
2024-11-28 | B | 8.2 | 6367.68
2024-11-28 | B | 8.2 | 6333.02
2024-11-28 | B | 8.2 | 6329.34
2024-11-28 | B | 8.2 | 6300.28
2024-11-28 | B | 8.2 | 6258.74
2024-11-28 | B | 8.2 | 6171.54
2024-11-28 | B | 8.2 | 6261.87
2024-11-28 | B | 8.2 | 6226.04
2024-11-28 | B | 8.2 | 6357.73
2024-11-28 | B | 8.2 | 6296.2
2024-11-28 | B | 8.2 | 6275.09
2024-11-28 | B | 8.2 | 6244.88
2024-11-28 | B | 8.2 | 6263.56
2024-11-28 | B | 8.2 | 6324.75
2024-11-28 | B | 8.2 | 6339.37
2024-11-28 | B | 8.2 | 6243.05
2024-11-28 | B | 8.2 | 6239.41
2024-11-28 | B | 6.15 | 4880.91
2024-11-28 | B | 6.15 | 4837.27
2024-11-28 | B | 6.15 | 4996.75
2024-11-28 | B | 6.15 | 4911.31
2024-11-28 | B | 6.15 | 4917.22
2024-11-28 | B | 6.15 | 4900.11
2024-11-28 | B | 6.15 | 4880.85
2024-11-28 | B | 6.15 | 4859.02
2024-11-28 | B | 6.15 | 4852.58
2024-11-28 | B | 6.15 | 4932.3
2024-11-28 | B | 6.15 | 4977.21
2024-11-28 | B | 6.15 | 4859.73
2024-11-28 | B | 6.15 | 4920.26
2024-11-28 | B | 6.15 | 4942.52
2024-11-28 | B | 6.15 | 4959.13
2024-11-28 | B | 6.15 | 4907.4
2024-11-28 | B | 6.15 | 4975.81
2024-11-28 | B | 6.15 | 4961.3
2024-11-28 | B | 6.15 | 4909.35
2024-11-28 | B | 6.15 | 4898.11
2024-11-28 | B | 4.1 | 3301.2
2024-11-28 | B | 4.1 | 3327.55
2024-11-28 | B | 4.1 | 3309.32
2024-11-28 | B | 4.1 | 3307.22
2024-11-28 | B | 4.1 | 3353.36
2024-11-28 | B | 4.1 | 3394.83
2024-11-28 | B | 4.1 | 3303.53
2024-11-28 | B | 4.1 | 3285.25
2024-11-28 | B | 4.1 | 3340.05
2024-11-28 | B | 4.1 | 3357.14
2024-11-28 | B | 4.1 | 3333.98
2024-11-28 | B | 4.1 | 3250.09
2024-11-28 | B | 4.1 | 3352.54
2024-11-28 | B | 4.1 | 3321.95
2024-11-28 | B | 4.1 | 3368.51
2024-11-28 | B | 4.1 | 3333.16
2024-11-28 | B | 4.1 | 3355.83
2024-11-28 | B | 4.1 | 3373.56
2024-11-28 | B | 4.1 | 3340.18
2024-11-28 | B | 4.1 | 3331.21
2024-11-28 | B | 2.05 | 1752.18
2024-11-28 | B | 2.05 | 1775.33
2024-11-28 | B | 2.05 | 1719.81
2024-11-28 | B | 2.05 | 1734.75
2024-11-28 | B | 2.05 | 1746.33
2024-11-28 | B | 2.05 | 1740.84
2024-11-28 | B | 2.05 | 1766.16
2024-11-28 | B | 2.05 | 1703.6
2024-11-28 | B | 2.05 | 1741.61
2024-11-28 | B | 2.05 | 1775.21
2024-11-28 | B | 2.05 | 1724.06
2024-11-28 | B | 2.05 | 1728.55
2024-11-28 | B | 2.05 | 1685.36
2024-11-28 | B | 2.05 | 1703.01
2024-11-28 | B | 2.05 | 1764.05
2024-11-28 | B | 2.05 | 1734.04
2024-11-28 | B | 2.05 | 1783.21
2024-11-28 | B | 2.05 | 1750.13
2024-11-28 | B | 2.05 | 1713.93
2024-11-28 | B | 2.05 | 1752.42
2024-11-28 | B | 0 | 122.3
2024-11-28 | B | 0 | 114.44
2024-11-28 | B | 0 | 124.87
2024-11-28 | B | 0 | 112.24
2024-11-28 | B | 0 | 122.55
2024-11-28 | B | 0 | 115.07
2024-11-28 | B | 0 | 119.06
2024-11-28 | B | 0 | 110.19
2024-11-28 | B | 0 | 120.02
2024-11-28 | B | 0 | 123.02
2024-11-28 | B | 0 | 117.35
2024-11-28 | B | 0 | 121.77
2024-11-28 | B | 0 | 110.75
2024-11-28 | B | 0 | 121.12
2024-11-28 | B | 0 | 125.33
2024-11-28 | B | 0 | 104.78
2024-11-28 | B | 0 | 116.73
2024-11-28 | B | 0 | 122.91
2024-11-28 | B | 0 | 118.89
2024-11-28 | B | 0 | 114.28
2024-11-28 | B | NA | 4845.76 | 5572.624
2024-11-28 | B | NA | 4855.93 | 5584.3195
2024-11-28 | B | NA | 4820.03 | 5543.0345
2024-11-28 | B | NA | 4869.33 | 5599.7295
2024-11-28 | B | NA | 4743.84 | 5455.416
2024-11-28 | B | NA | 4905.01 | 5640.7615
2024-11-28 | B | NA | 4783.78 | 5501.347
2024-11-28 | B | NA | 4868.35 | 5598.6025
2024-11-28 | B | NA | 4880.35 | 5612.4025
2024-11-28 | B | NA | 4886.97 | 5620.0155
2024-11-28 | B | NA | 4853.44 | 5581.456
2024-11-28 | B | NA | 4896.39 | 5630.8485
2024-11-28 | B | NA | 4892.91 | 5626.8465
2024-11-28 | B | NA | 4896.82 | 5631.343
2024-11-28 | B | NA | 4904.14 | 5639.761
2024-11-28 | B | NA | 4931.55 | 5671.2825
2024-11-28 | B | NA | 4863.35 | 5592.8525
2024-11-28 | B | NA | 4857.58 | 5586.217
2024-11-28 | B | NA | 4913.57 | 5650.6055
2024-11-28 | B | NA | 4864.82 | 5594.543
2024-12-02 | B | 10.25 | 7281.12
2024-12-02 | B | 10.25 | 7257.35
2024-12-02 | B | 10.25 | 7292.07
2024-12-02 | B | 10.25 | 7393.85
2024-12-02 | B | 10.25 | 7264.4
2024-12-02 | B | 10.25 | 7412.75
2024-12-02 | B | 10.25 | 7392.2
2024-12-02 | B | 10.25 | 7278.77
2024-12-02 | B | 10.25 | 7381.21
2024-12-02 | B | 10.25 | 7327.93
2024-12-02 | B | 10.25 | 7374.87
2024-12-02 | B | 10.25 | 7263.38
2024-12-02 | B | 10.25 | 7280.05
2024-12-02 | B | 10.25 | 7203.75
2024-12-02 | B | 10.25 | 7336.07
2024-12-02 | B | 10.25 | 7284.18
2024-12-02 | B | 10.25 | 7375.36
2024-12-02 | B | 10.25 | 7400.42
2024-12-02 | B | 10.25 | 7250.11
2024-12-02 | B | 10.25 | 7262.97
2024-12-02 | B | 8.2 | 6320.07
2024-12-02 | B | 8.2 | 6295.47
2024-12-02 | B | 8.2 | 6268.55
2024-12-02 | B | 8.2 | 6303.12
2024-12-02 | B | 8.2 | 6306.41
2024-12-02 | B | 8.2 | 6292.94
2024-12-02 | B | 8.2 | 6305.68
2024-12-02 | B | 8.2 | 6197.19
2024-12-02 | B | 8.2 | 6332.31
2024-12-02 | B | 8.2 | 6275.18
2024-12-02 | B | 8.2 | 6320.62
2024-12-02 | B | 8.2 | 6256.44
2024-12-02 | B | 8.2 | 6343.24
2024-12-02 | B | 8.2 | 6211
2024-12-02 | B | 8.2 | 6264.76
2024-12-02 | B | 8.2 | 6279.66
2024-12-02 | B | 8.2 | 6288.31
2024-12-02 | B | 8.2 | 6349.48
2024-12-02 | B | 8.2 | 6324.76
2024-12-02 | B | 8.2 | 6276.31
2024-12-02 | B | 6.15 | 4955.74
2024-12-02 | B | 6.15 | 4944.38
2024-12-02 | B | 6.15 | 4939.38
2024-12-02 | B | 6.15 | 4984.09
2024-12-02 | B | 6.15 | 4951.98
2024-12-02 | B | 6.15 | 4957.56
2024-12-02 | B | 6.15 | 4994.74
2024-12-02 | B | 6.15 | 4895.01
2024-12-02 | B | 6.15 | 4947.23
2024-12-02 | B | 6.15 | 4863.34
2024-12-02 | B | 6.15 | 4914.07
2024-12-02 | B | 6.15 | 4858.78
2024-12-02 | B | 6.15 | 4892.21
2024-12-02 | B | 6.15 | 4992.12
2024-12-02 | B | 6.15 | 4930
2024-12-02 | B | 6.15 | 4941.81
2024-12-02 | B | 6.15 | 4928.73
2024-12-02 | B | 6.15 | 4874.41
2024-12-02 | B | 6.15 | 4961.89
2024-12-02 | B | 6.15 | 4865.82
2024-12-02 | B | 4.1 | 3364.36
2024-12-02 | B | 4.1 | 3399.78
2024-12-02 | B | 4.1 | 3312.7
2024-12-02 | B | 4.1 | 3404.71
2024-12-02 | B | 4.1 | 3294.87
2024-12-02 | B | 4.1 | 3299.73
2024-12-02 | B | 4.1 | 3391.65
2024-12-02 | B | 4.1 | 3328.65
2024-12-02 | B | 4.1 | 3314.65
2024-12-02 | B | 4.1 | 3325.18
2024-12-02 | B | 4.1 | 3295
2024-12-02 | B | 4.1 | 3337.61
2024-12-02 | B | 4.1 | 3327.39
2024-12-02 | B | 4.1 | 3295.58
2024-12-02 | B | 4.1 | 3336.3
2024-12-02 | B | 4.1 | 3322.09
2024-12-02 | B | 4.1 | 3275.14
2024-12-02 | B | 4.1 | 3307.14
2024-12-02 | B | 4.1 | 3312.57
2024-12-02 | B | 4.1 | 3285
2024-12-02 | B | 2.05 | 1784.32
2024-12-02 | B | 2.05 | 1767.86
2024-12-02 | B | 2.05 | 1754.07
2024-12-02 | B | 2.05 | 1719.73
2024-12-02 | B | 2.05 | 1772.79
2024-12-02 | B | 2.05 | 1691.26
2024-12-02 | B | 2.05 | 1699.74
2024-12-02 | B | 2.05 | 1750.32
2024-12-02 | B | 2.05 | 1760.31
2024-12-02 | B | 2.05 | 1762.44
2024-12-02 | B | 2.05 | 1732.48
2024-12-02 | B | 2.05 | 1726.22
2024-12-02 | B | 2.05 | 1770.53
2024-12-02 | B | 2.05 | 1788.35
2024-12-02 | B | 2.05 | 1756.44
2024-12-02 | B | 2.05 | 1767.29
2024-12-02 | B | 2.05 | 1722.15
2024-12-02 | B | 2.05 | 1759.52
2024-12-02 | B | 2.05 | 1706.52
2024-12-02 | B | 2.05 | 1772.7
2024-12-02 | B | 0 | 115.83
2024-12-02 | B | 0 | 123.48
2024-12-02 | B | 0 | 125.64
2024-12-02 | B | 0 | 119.19
2024-12-02 | B | 0 | 111.34
2024-12-02 | B | 0 | 118.16
2024-12-02 | B | 0 | 115.36
2024-12-02 | B | 0 | 116.63
2024-12-02 | B | 0 | 121.59
2024-12-02 | B | 0 | 125.45
2024-12-02 | B | 0 | 111.72
2024-12-02 | B | 0 | 123.71
2024-12-02 | B | 0 | 123.43
2024-12-02 | B | 0 | 119.73
2024-12-02 | B | 0 | 111.4
2024-12-02 | B | 0 | 111.89
2024-12-02 | B | 0 | 122.76
2024-12-02 | B | 0 | 114.02
2024-12-02 | B | 0 | 111.88
2024-12-02 | B | 0 | 113.78
2024-12-02 | B | NA | 4826.95 | 5550.9925
2024-12-02 | B | NA | 4897.61 | 5632.2515
2024-12-02 | B | NA | 4823.51 | 5547.0365
2024-12-02 | B | NA | 4924.64 | 5663.336
2024-12-02 | B | NA | 4817.52 | 5540.148
2024-12-02 | B | NA | 4939.05 | 5679.9075
2024-12-02 | B | NA | 4857.75 | 5586.4125
2024-12-02 | B | NA | 4925.13 | 5663.8995
2024-12-02 | B | NA | 4858.06 | 5586.769
2024-12-02 | B | NA | 4822.24 | 5545.576
2024-12-02 | B | NA | 4908.07 | 5644.2805
2024-12-02 | B | NA | 4958.56 | 5702.344
2024-12-02 | B | NA | 4925.01 | 5663.7615
2024-12-02 | B | NA | 4859.53 | 5588.4595
2024-12-02 | B | NA | 4857.8 | 5586.47
2024-12-02 | B | NA | 4828.44 | 5552.706
2024-12-02 | B | NA | 4938.69 | 5679.4935
2024-12-02 | B | NA | 4916.13 | 5653.5495
2024-12-02 | B | NA | 4888.09 | 5621.3035
2024-12-02 | B | NA | 4885.1 | 5617.865
2024-12-03 | B | 10.25 | 7425.02
2024-12-03 | B | 10.25 | 7218.92
2024-12-03 | B | 10.25 | 7305.41
2024-12-03 | B | 10.25 | 7428.88
2024-12-03 | B | 10.25 | 7345.72
2024-12-03 | B | 10.25 | 7351.73
2024-12-03 | B | 10.25 | 7347.55
2024-12-03 | B | 10.25 | 7323.58
2024-12-03 | B | 10.25 | 7307.95
2024-12-03 | B | 10.25 | 7352.98
2024-12-03 | B | 10.25 | 7311.79
2024-12-03 | B | 10.25 | 7343.39
2024-12-03 | B | 10.25 | 7387.57
2024-12-03 | B | 10.25 | 7253.88
2024-12-03 | B | 10.25 | 7345.21
2024-12-03 | B | 10.25 | 7355.39
2024-12-03 | B | 10.25 | 7209.66
2024-12-03 | B | 10.25 | 7369.01
2024-12-03 | B | 10.25 | 7421.1
2024-12-03 | B | 10.25 | 7248.22
2024-12-03 | B | 8.2 | 6288.53
2024-12-03 | B | 8.2 | 6268.85
2024-12-03 | B | 8.2 | 6245.96
2024-12-03 | B | 8.2 | 6235.36
2024-12-03 | B | 8.2 | 6299.63
2024-12-03 | B | 8.2 | 6234.82
2024-12-03 | B | 8.2 | 6245.27
2024-12-03 | B | 8.2 | 6285.05
2024-12-03 | B | 8.2 | 6336.67
2024-12-03 | B | 8.2 | 6271.51
2024-12-03 | B | 8.2 | 6275.15
2024-12-03 | B | 8.2 | 6271.05
2024-12-03 | B | 8.2 | 6285.73
2024-12-03 | B | 8.2 | 6338.81
2024-12-03 | B | 8.2 | 6276.96
2024-12-03 | B | 8.2 | 6283.24
2024-12-03 | B | 8.2 | 6312.55
2024-12-03 | B | 8.2 | 6150.32
2024-12-03 | B | 8.2 | 6288.6
2024-12-03 | B | 8.2 | 6267.78
2024-12-03 | B | 6.15 | 4966.87
2024-12-03 | B | 6.15 | 4933.21
2024-12-03 | B | 6.15 | 4869.46
2024-12-03 | B | 6.15 | 4914.72
2024-12-03 | B | 6.15 | 4907.49
2024-12-03 | B | 6.15 | 4885.29
2024-12-03 | B | 6.15 | 4860.52
2024-12-03 | B | 6.15 | 4893.45
2024-12-03 | B | 6.15 | 4982.03
2024-12-03 | B | 6.15 | 4908.78
2024-12-03 | B | 6.15 | 4867.67
2024-12-03 | B | 6.15 | 4858.32
2024-12-03 | B | 6.15 | 4926.03
2024-12-03 | B | 6.15 | 4908.67
2024-12-03 | B | 6.15 | 4910.67
2024-12-03 | B | 6.15 | 4920.91
2024-12-03 | B | 6.15 | 4883.18
2024-12-03 | B | 6.15 | 4927.94
2024-12-03 | B | 6.15 | 4951.81
2024-12-03 | B | 6.15 | 4961.63
2024-12-03 | B | 4.1 | 3334.98
2024-12-03 | B | 4.1 | 3251.43
2024-12-03 | B | 4.1 | 3323.96
2024-12-03 | B | 4.1 | 3312.37
2024-12-03 | B | 4.1 | 3348.33
2024-12-03 | B | 4.1 | 3289.14
2024-12-03 | B | 4.1 | 3328.02
2024-12-03 | B | 4.1 | 3354.74
2024-12-03 | B | 4.1 | 3316.37
2024-12-03 | B | 4.1 | 3308.71
2024-12-03 | B | 4.1 | 3377.29
2024-12-03 | B | 4.1 | 3335.66
2024-12-03 | B | 4.1 | 3345.75
2024-12-03 | B | 4.1 | 3349.94
2024-12-03 | B | 4.1 | 3307.01
2024-12-03 | B | 4.1 | 3336.93
2024-12-03 | B | 4.1 | 3306.7
2024-12-03 | B | 4.1 | 3276.65
2024-12-03 | B | 4.1 | 3288.03
2024-12-03 | B | 4.1 | 3336.74
2024-12-03 | B | 2.05 | 1734.92
2024-12-03 | B | 2.05 | 1759.57
2024-12-03 | B | 2.05 | 1759.79
2024-12-03 | B | 2.05 | 1745.17
2024-12-03 | B | 2.05 | 1743.7
2024-12-03 | B | 2.05 | 1725.44
2024-12-03 | B | 2.05 | 1723.67
2024-12-03 | B | 2.05 | 1739.68
2024-12-03 | B | 2.05 | 1722.2
2024-12-03 | B | 2.05 | 1729.86
2024-12-03 | B | 2.05 | 1724.69
2024-12-03 | B | 2.05 | 1766.88
2024-12-03 | B | 2.05 | 1793.53
2024-12-03 | B | 2.05 | 1701.21
2024-12-03 | B | 2.05 | 1709.18
2024-12-03 | B | 2.05 | 1708.07
2024-12-03 | B | 2.05 | 1754.8
2024-12-03 | B | 2.05 | 1718.33
2024-12-03 | B | 2.05 | 1751.94
2024-12-03 | B | 2.05 | 1768.33
2024-12-03 | B | 0 | 120.44
2024-12-03 | B | 0 | 116.68
2024-12-03 | B | 0 | 121.2
2024-12-03 | B | 0 | 123.14
2024-12-03 | B | 0 | 125.94
2024-12-03 | B | 0 | 106.06
2024-12-03 | B | 0 | 126.84
2024-12-03 | B | 0 | 124.71
2024-12-03 | B | 0 | 120.96
2024-12-03 | B | 0 | 119.98
2024-12-03 | B | 0 | 114.91
2024-12-03 | B | 0 | 110.82
2024-12-03 | B | 0 | 121.44
2024-12-03 | B | 0 | 124.67
2024-12-03 | B | 0 | 115.39
2024-12-03 | B | 0 | 113.66
2024-12-03 | B | 0 | 111.11
2024-12-03 | B | 0 | 122.88
2024-12-03 | B | 0 | 120.07
2024-12-03 | B | 0 | 119.5
2024-12-03 | B | NA | 4910.02 | 5646.523
2024-12-03 | B | NA | 4911.46 | 5648.179
2024-12-03 | B | NA | 4856.05 | 5584.4575
2024-12-03 | B | NA | 4886.48 | 5619.452
2024-12-03 | B | NA | 4896.23 | 5630.6645
2024-12-03 | B | NA | 4897.79 | 5632.4585
2024-12-03 | B | NA | 4883.39 | 5615.8985
2024-12-03 | B | NA | 4886.72 | 5619.728
2024-12-03 | B | NA | 4928.14 | 5667.361
2024-12-03 | B | NA | 4823.87 | 5547.4505
2024-12-03 | B | NA | 4848.31 | 5575.5565
2024-12-03 | B | NA | 4855.79 | 5584.1585
2024-12-03 | B | NA | 4851.41 | 5579.1215
2024-12-03 | B | NA | 4818.52 | 5541.298
2024-12-03 | B | NA | 4870.72 | 5601.328
2024-12-03 | B | NA | 4859.03 | 5587.8845
2024-12-03 | B | NA | 4859.7 | 5588.655
2024-12-03 | B | NA | 4912.18 | 5649.007
2024-12-03 | B | NA | 4911.75 | 5648.5125
2024-12-03 | B | NA | 4919.92 | 5657.908
2024-12-04 | B | 10.25 | 7267.46
2024-12-04 | B | 10.25 | 7382.62
2024-12-04 | B | 10.25 | 7308.54
2024-12-04 | B | 10.25 | 7309.87
2024-12-04 | B | 10.25 | 7390.97
2024-12-04 | B | 10.25 | 7272.98
2024-12-04 | B | 10.25 | 7280.12
2024-12-04 | B | 10.25 | 7337.29
2024-12-04 | B | 10.25 | 7312.84
2024-12-04 | B | 10.25 | 7263.19
2024-12-04 | B | 10.25 | 7320.85
2024-12-04 | B | 10.25 | 7293.52
2024-12-04 | B | 10.25 | 7319.2
2024-12-04 | B | 10.25 | 7349.03
2024-12-04 | B | 10.25 | 7285
2024-12-04 | B | 10.25 | 7195.01
2024-12-04 | B | 10.25 | 7347.9
2024-12-04 | B | 10.25 | 7380.93
2024-12-04 | B | 10.25 | 7396.17
2024-12-04 | B | 10.25 | 7342.61
2024-12-04 | B | 8.2 | 6274.21
2024-12-04 | B | 8.2 | 6307.83
2024-12-04 | B | 8.2 | 6296.8
2024-12-04 | B | 8.2 | 6307.53
2024-12-04 | B | 8.2 | 6330.53
2024-12-04 | B | 8.2 | 6342.62
2024-12-04 | B | 8.2 | 6339.11
2024-12-04 | B | 8.2 | 6290.29
2024-12-04 | B | 8.2 | 6182.15
2024-12-04 | B | 8.2 | 6339.79
2024-12-04 | B | 8.2 | 6291.7
2024-12-04 | B | 8.2 | 6264.88
2024-12-04 | B | 8.2 | 6299.53
2024-12-04 | B | 8.2 | 6303.54
2024-12-04 | B | 8.2 | 6281.94
2024-12-04 | B | 8.2 | 6307.65
2024-12-04 | B | 8.2 | 6413.85
2024-12-04 | B | 8.2 | 6321.79
2024-12-04 | B | 8.2 | 6272.75
2024-12-04 | B | 8.2 | 6311.85
2024-12-04 | B | 6.15 | 4892.91
2024-12-04 | B | 6.15 | 4866.52
2024-12-04 | B | 6.15 | 4888.77
2024-12-04 | B | 6.15 | 4924.59
2024-12-04 | B | 6.15 | 4871.21
2024-12-04 | B | 6.15 | 4943.31
2024-12-04 | B | 6.15 | 4973.14
2024-12-04 | B | 6.15 | 4954.31
2024-12-04 | B | 6.15 | 4866.66
2024-12-04 | B | 6.15 | 4880.24
2024-12-04 | B | 6.15 | 4918.02
2024-12-04 | B | 6.15 | 4872.52
2024-12-04 | B | 6.15 | 4913.23
2024-12-04 | B | 6.15 | 4877.6
2024-12-04 | B | 6.15 | 4931.97
2024-12-04 | B | 6.15 | 4941.77
2024-12-04 | B | 6.15 | 4960.37
2024-12-04 | B | 6.15 | 4911.91
2024-12-04 | B | 6.15 | 4878.89
2024-12-04 | B | 6.15 | 4949.43
2024-12-04 | B | 4.1 | 3315.31
2024-12-04 | B | 4.1 | 3286.2
2024-12-04 | B | 4.1 | 3354.83
2024-12-04 | B | 4.1 | 3267.19
2024-12-04 | B | 4.1 | 3330.28
2024-12-04 | B | 4.1 | 3383.78
2024-12-04 | B | 4.1 | 3329.17
2024-12-04 | B | 4.1 | 3388.01
2024-12-04 | B | 4.1 | 3264.24
2024-12-04 | B | 4.1 | 3275.68
2024-12-04 | B | 4.1 | 3284.56
2024-12-04 | B | 4.1 | 3327.45
2024-12-04 | B | 4.1 | 3313.89
2024-12-04 | B | 4.1 | 3361.17
2024-12-04 | B | 4.1 | 3308.3
2024-12-04 | B | 4.1 | 3321.71
2024-12-04 | B | 4.1 | 3326.65
2024-12-04 | B | 4.1 | 3304.02
2024-12-04 | B | 4.1 | 3320
2024-12-04 | B | 4.1 | 3295.18
2024-12-04 | B | 2.05 | 1781.92
2024-12-04 | B | 2.05 | 1780.27
2024-12-04 | B | 2.05 | 1744.78
2024-12-04 | B | 2.05 | 1712.13
2024-12-04 | B | 2.05 | 1786.64
2024-12-04 | B | 2.05 | 1699.83
2024-12-04 | B | 2.05 | 1756.83
2024-12-04 | B | 2.05 | 1751.67
2024-12-04 | B | 2.05 | 1768.66
2024-12-04 | B | 2.05 | 1784.72
2024-12-04 | B | 2.05 | 1748.7
2024-12-04 | B | 2.05 | 1738.17
2024-12-04 | B | 2.05 | 1804.17
2024-12-04 | B | 2.05 | 1768.03
2024-12-04 | B | 2.05 | 1697.23
2024-12-04 | B | 2.05 | 1725.58
2024-12-04 | B | 2.05 | 1728.03
2024-12-04 | B | 2.05 | 1748.47
2024-12-04 | B | 2.05 | 1731.38
2024-12-04 | B | 2.05 | 1698.87
2024-12-04 | B | 0 | 120.36
2024-12-04 | B | 0 | 106.96
2024-12-04 | B | 0 | 128.36
2024-12-04 | B | 0 | 122.3
2024-12-04 | B | 0 | 117.72
2024-12-04 | B | 0 | 123.69
2024-12-04 | B | 0 | 118.65
2024-12-04 | B | 0 | 124.56
2024-12-04 | B | 0 | 128.67
2024-12-04 | B | 0 | 116.21
2024-12-04 | B | 0 | 117.28
2024-12-04 | B | 0 | 114.13
2024-12-04 | B | 0 | 120.81
2024-12-04 | B | 0 | 121.92
2024-12-04 | B | 0 | 125.55
2024-12-04 | B | 0 | 115.14
2024-12-04 | B | 0 | 122.3
2024-12-04 | B | 0 | 125.46
2024-12-04 | B | 0 | 115.85
2024-12-04 | B | 0 | 112.67
2024-12-04 | B | NA | 4884.95 | 5617.6925
2024-12-04 | B | NA | 4887.36 | 5620.464
2024-12-04 | B | NA | 4991.12 | 5739.788
2024-12-04 | B | NA | 4883.41 | 5615.9215
2024-12-04 | B | NA | 4940.43 | 5681.4945
2024-12-04 | B | NA | 4846.49 | 5573.4635
2024-12-04 | B | NA | 4903.99 | 5639.5885
2024-12-04 | B | NA | 4924.48 | 5663.152
2024-12-04 | B | NA | 5009.77 | 5761.2355
2024-12-04 | B | NA | 4850.93 | 5578.5695
2024-12-04 | B | NA | 4844.01 | 5570.6115
2024-12-04 | B | NA | 4954.82 | 5698.043
2024-12-04 | B | NA | 4931.58 | 5671.317
2024-12-04 | B | NA | 4894.87 | 5629.1005
2024-12-04 | B | NA | 4845.22 | 5572.003
2024-12-04 | B | NA | 4919.68 | 5657.632
2024-12-04 | B | NA | 4972.49 | 5718.3635
2024-12-04 | B | NA | 4906.55 | 5642.5325
2024-12-04 | B | NA | 4901.14 | 5636.311
2024-12-04 | B | NA | 4841.37 | 5567.5755
2024-12-05 | B | 10.25 | 7357.19
2024-12-05 | B | 10.25 | 7322.38
2024-12-05 | B | 10.25 | 7315.72
2024-12-05 | B | 10.25 | 7322.59
2024-12-05 | B | 10.25 | 7273.92
2024-12-05 | B | 10.25 | 7227.35
2024-12-05 | B | 10.25 | 7354.69
2024-12-05 | B | 10.25 | 7311.99
2024-12-05 | B | 10.25 | 7319.87
2024-12-05 | B | 10.25 | 7288.13
2024-12-05 | B | 10.25 | 7318.34
2024-12-05 | B | 10.25 | 7431.29
2024-12-05 | B | 10.25 | 7392.6
2024-12-05 | B | 10.25 | 7378.54
2024-12-05 | B | 10.25 | 7345.92
2024-12-05 | B | 10.25 | 7309.47
2024-12-05 | B | 10.25 | 7292.5
2024-12-05 | B | 10.25 | 7324.26
2024-12-05 | B | 10.25 | 7190.89
2024-12-05 | B | 10.25 | 7278.62
2024-12-05 | B | 8.2 | 6339.11
2024-12-05 | B | 8.2 | 6255.58
2024-12-05 | B | 8.2 | 6346.69
2024-12-05 | B | 8.2 | 6301.49
2024-12-05 | B | 8.2 | 6295.73
2024-12-05 | B | 8.2 | 6185.89
2024-12-05 | B | 8.2 | 6247.49
2024-12-05 | B | 8.2 | 6281.41
2024-12-05 | B | 8.2 | 6310.05
2024-12-05 | B | 8.2 | 6284.47
2024-12-05 | B | 8.2 | 6222.76
2024-12-05 | B | 8.2 | 6169.56
2024-12-05 | B | 8.2 | 6267.8
2024-12-05 | B | 8.2 | 6270.41
2024-12-05 | B | 8.2 | 6283.31
2024-12-05 | B | 8.2 | 6280.23
2024-12-05 | B | 8.2 | 6321.72
2024-12-05 | B | 8.2 | 6295.1
2024-12-05 | B | 8.2 | 6276.87
2024-12-05 | B | 8.2 | 6264.2
2024-12-05 | B | 6.15 | 4869.95
2024-12-05 | B | 6.15 | 4944.95
2024-12-05 | B | 6.15 | 4884.62
2024-12-05 | B | 6.15 | 4863.77
2024-12-05 | B | 6.15 | 4891.96
2024-12-05 | B | 6.15 | 4888.15
2024-12-05 | B | 6.15 | 4913.71
2024-12-05 | B | 6.15 | 4935.41
2024-12-05 | B | 6.15 | 4909.11
2024-12-05 | B | 6.15 | 4893.35
2024-12-05 | B | 6.15 | 4884.29
2024-12-05 | B | 6.15 | 4940.47
2024-12-05 | B | 6.15 | 4950.43
2024-12-05 | B | 6.15 | 4925.47
2024-12-05 | B | 6.15 | 4923.19
2024-12-05 | B | 6.15 | 4941.91
2024-12-05 | B | 6.15 | 4939.64
2024-12-05 | B | 6.15 | 4916.74
2024-12-05 | B | 6.15 | 4893.94
2024-12-05 | B | 6.15 | 4876.17
2024-12-05 | B | 4.1 | 3282.78
2024-12-05 | B | 4.1 | 3301.47
2024-12-05 | B | 4.1 | 3283.01
2024-12-05 | B | 4.1 | 3303.02
2024-12-05 | B | 4.1 | 3300.56
2024-12-05 | B | 4.1 | 3340.9
2024-12-05 | B | 4.1 | 3330.07
2024-12-05 | B | 4.1 | 3350.88
2024-12-05 | B | 4.1 | 3287.36
2024-12-05 | B | 4.1 | 3275.47
2024-12-05 | B | 4.1 | 3304.67
2024-12-05 | B | 4.1 | 3366.58
2024-12-05 | B | 4.1 | 3309.16
2024-12-05 | B | 4.1 | 3346.97
2024-12-05 | B | 4.1 | 3294.22
2024-12-05 | B | 4.1 | 3317.15
2024-12-05 | B | 4.1 | 3325.8
2024-12-05 | B | 4.1 | 3324.14
2024-12-05 | B | 4.1 | 3393.56
2024-12-05 | B | 4.1 | 3296.23
2024-12-05 | B | 2.05 | 1763.3
2024-12-05 | B | 2.05 | 1750.53
2024-12-05 | B | 2.05 | 1716.85
2024-12-05 | B | 2.05 | 1759.43
2024-12-05 | B | 2.05 | 1769.59
2024-12-05 | B | 2.05 | 1753.64
2024-12-05 | B | 2.05 | 1783.86
2024-12-05 | B | 2.05 | 1842.73
2024-12-05 | B | 2.05 | 1752.74
2024-12-05 | B | 2.05 | 1775.37
2024-12-05 | B | 2.05 | 1741.07
2024-12-05 | B | 2.05 | 1752.86
2024-12-05 | B | 2.05 | 1773.68
2024-12-05 | B | 2.05 | 1760.28
2024-12-05 | B | 2.05 | 1759.78
2024-12-05 | B | 2.05 | 1738.44
2024-12-05 | B | 2.05 | 1778.11
2024-12-05 | B | 2.05 | 1760.1
2024-12-05 | B | 2.05 | 1725.4
2024-12-05 | B | 2.05 | 1741.37
2024-12-05 | B | 0 | 110.7
2024-12-05 | B | 0 | 115.1
2024-12-05 | B | 0 | 113.5
2024-12-05 | B | 0 | 122.33
2024-12-05 | B | 0 | 116.16
2024-12-05 | B | 0 | 118.79
2024-12-05 | B | 0 | 115.62
2024-12-05 | B | 0 | 123.11
2024-12-05 | B | 0 | 128.52
2024-12-05 | B | 0 | 123.34
2024-12-05 | B | 0 | 112.89
2024-12-05 | B | 0 | 115.76
2024-12-05 | B | 0 | 116.98
2024-12-05 | B | 0 | 111.3
2024-12-05 | B | 0 | 113.26
2024-12-05 | B | 0 | 118.8
2024-12-05 | B | 0 | 120.26
2024-12-05 | B | 0 | 116.14
2024-12-05 | B | 0 | 115.31
2024-12-05 | B | 0 | 116.55
2024-12-05 | B | NA | 4881.56 | 5613.794
2024-12-05 | B | NA | 4934.65 | 5674.8475
2024-12-05 | B | NA | 4843.49 | 5570.0135
2024-12-05 | B | NA | 4950.57 | 5693.1555
2024-12-05 | B | NA | 4802.44 | 5522.806
2024-12-05 | B | NA | 4841.98 | 5568.277
2024-12-05 | B | NA | 4909.6 | 5646.04
2024-12-05 | B | NA | 4774.94 | 5491.181
2024-12-05 | B | NA | 4851.83 | 5579.6045
2024-12-05 | B | NA | 4801.52 | 5521.748
2024-12-05 | B | NA | 4867.45 | 5597.5675
2024-12-05 | B | NA | 4892 | 5625.8
2024-12-05 | B | NA | 4905.83 | 5641.7045
2024-12-05 | B | NA | 4870.33 | 5600.8795
2024-12-05 | B | NA | 4836.82 | 5562.343
2024-12-05 | B | NA | 4892.7 | 5626.605
2024-12-05 | B | NA | 4829.37 | 5553.7755
2024-12-05 | B | NA | 4893.98 | 5628.077
2024-12-05 | B | NA | 4819.2 | 5542.08
2024-12-05 | B | NA | 4869.74 | 5600.201
2024-12-06 | C | 10.25 | 7354.22
2024-12-06 | C | 10.25 | 7373.47
2024-12-06 | C | 10.25 | 7375.29
2024-12-06 | C | 10.25 | 7328.68
2024-12-06 | C | 10.25 | 7345.66
2024-12-06 | C | 10.25 | 7301.03
2024-12-06 | C | 10.25 | 7298.17
2024-12-06 | C | 10.25 | 7342.94
2024-12-06 | C | 10.25 | 7341.59
2024-12-06 | C | 10.25 | 7307.72
2024-12-06 | C | 10.25 | 7364.31
2024-12-06 | C | 10.25 | 7349.15
2024-12-06 | C | 10.25 | 7313.68
2024-12-06 | C | 10.25 | 7347.44
2024-12-06 | C | 10.25 | 7376.58
2024-12-06 | C | 10.25 | 7341.7
2024-12-06 | C | 10.25 | 7268.64
2024-12-06 | C | 10.25 | 7355.51
2024-12-06 | C | 10.25 | 7371.11
2024-12-06 | C | 10.25 | 7436.03
2024-12-06 | C | 8.2 | 6194.79
2024-12-06 | C | 8.2 | 6341.36
2024-12-06 | C | 8.2 | 6263.27
2024-12-06 | C | 8.2 | 6185.02
2024-12-06 | C | 8.2 | 6245
2024-12-06 | C | 8.2 | 6277.91
2024-12-06 | C | 8.2 | 6292.96
2024-12-06 | C | 8.2 | 6277.6
2024-12-06 | C | 8.2 | 6266.56
2024-12-06 | C | 8.2 | 6238.48
2024-12-06 | C | 8.2 | 6337.35
2024-12-06 | C | 8.2 | 6343.12
2024-12-06 | C | 8.2 | 6299.37
2024-12-06 | C | 8.2 | 6273.87
2024-12-06 | C | 8.2 | 6326.63
2024-12-06 | C | 8.2 | 6291.14
2024-12-06 | C | 8.2 | 6222.35
2024-12-06 | C | 8.2 | 6323.2
2024-12-06 | C | 8.2 | 6358.47
2024-12-06 | C | 8.2 | 6303.22
2024-12-06 | C | 6.15 | 4960.88
2024-12-06 | C | 6.15 | 4897.95
2024-12-06 | C | 6.15 | 4884.55
2024-12-06 | C | 6.15 | 4961.09
2024-12-06 | C | 6.15 | 4830
2024-12-06 | C | 6.15 | 4932.56
2024-12-06 | C | 6.15 | 4884.45
2024-12-06 | C | 6.15 | 4829.7
2024-12-06 | C | 6.15 | 4865.66
2024-12-06 | C | 6.15 | 4926.7
2024-12-06 | C | 6.15 | 4821.69
2024-12-06 | C | 6.15 | 4887.23
2024-12-06 | C | 6.15 | 4911.19
2024-12-06 | C | 6.15 | 4931.84
2024-12-06 | C | 6.15 | 4922.52
2024-12-06 | C | 6.15 | 4938.61
2024-12-06 | C | 6.15 | 4885.33
2024-12-06 | C | 6.15 | 4842.88
2024-12-06 | C | 6.15 | 4871.63
2024-12-06 | C | 6.15 | 5016.5
2024-12-06 | C | 4.1 | 3318.53
2024-12-06 | C | 4.1 | 3380.46
2024-12-06 | C | 4.1 | 3366.29
2024-12-06 | C | 4.1 | 3312.78
2024-12-06 | C | 4.1 | 3269.85
2024-12-06 | C | 4.1 | 3350.33
2024-12-06 | C | 4.1 | 3311.88
2024-12-06 | C | 4.1 | 3320.56
2024-12-06 | C | 4.1 | 3388.9
2024-12-06 | C | 4.1 | 3328.5
2024-12-06 | C | 4.1 | 3292.05
2024-12-06 | C | 4.1 | 3297.3
2024-12-06 | C | 4.1 | 3350.39
2024-12-06 | C | 4.1 | 3341.94
2024-12-06 | C | 4.1 | 3276.05
2024-12-06 | C | 4.1 | 3357.68
2024-12-06 | C | 4.1 | 3356.45
2024-12-06 | C | 4.1 | 3347.31
2024-12-06 | C | 4.1 | 3340.72
2024-12-06 | C | 4.1 | 3285.71
2024-12-06 | C | 2.05 | 1756.82
2024-12-06 | C | 2.05 | 1724.12
2024-12-06 | C | 2.05 | 1797.2
2024-12-06 | C | 2.05 | 1744.68
2024-12-06 | C | 2.05 | 1705.05
2024-12-06 | C | 2.05 | 1767.04
2024-12-06 | C | 2.05 | 1765.27
2024-12-06 | C | 2.05 | 1743.47
2024-12-06 | C | 2.05 | 1727.01
2024-12-06 | C | 2.05 | 1758.11
2024-12-06 | C | 2.05 | 1737.18
2024-12-06 | C | 2.05 | 1785.32
2024-12-06 | C | 2.05 | 1722.47
2024-12-06 | C | 2.05 | 1775.14
2024-12-06 | C | 2.05 | 1772.13
2024-12-06 | C | 2.05 | 1721.19
2024-12-06 | C | 2.05 | 1743.89
2024-12-06 | C | 2.05 | 1732.84
2024-12-06 | C | 2.05 | 1729.98
2024-12-06 | C | 2.05 | 1704.15
2024-12-06 | C | 0 | 116.01
2024-12-06 | C | 0 | 121.82
2024-12-06 | C | 0 | 117.43
2024-12-06 | C | 0 | 117.92
2024-12-06 | C | 0 | 111.51
2024-12-06 | C | 0 | 111.91
2024-12-06 | C | 0 | 121.36
2024-12-06 | C | 0 | 120.11
2024-12-06 | C | 0 | 113.38
2024-12-06 | C | 0 | 121.62
2024-12-06 | C | 0 | 107.08
2024-12-06 | C | 0 | 116.62
2024-12-06 | C | 0 | 104.37
2024-12-06 | C | 0 | 122.79
2024-12-06 | C | 0 | 120.94
2024-12-06 | C | 0 | 118.61
2024-12-06 | C | 0 | 127.59
2024-12-06 | C | 0 | 123.17
2024-12-06 | C | 0 | 119.87
2024-12-06 | C | 0 | 124.81
2024-12-06 | C | NA | 4882.07 | 5614.3805
2024-12-06 | C | NA | 4910.58 | 5647.167
2024-12-06 | C | NA | 4897.19 | 5631.7685
2024-12-06 | C | NA | 4832.43 | 5557.2945
2024-12-06 | C | NA | 4869.88 | 5600.362
2024-12-06 | C | NA | 4853.04 | 5580.996
2024-12-06 | C | NA | 4890.46 | 5624.029
2024-12-06 | C | NA | 4977.2 | 5723.78
2024-12-06 | C | NA | 4827.52 | 5551.648
2024-12-06 | C | NA | 4948.35 | 5690.6025
2024-12-06 | C | NA | 4912.28 | 5649.122
2024-12-06 | C | NA | 4807.46 | 5528.579
2024-12-06 | C | NA | 4839.12 | 5564.988
2024-12-06 | C | NA | 4883.19 | 5615.6685
2024-12-06 | C | NA | 4875.39 | 5606.6985
2024-12-06 | C | NA | 4883.19 | 5615.6685
2024-12-06 | C | NA | 4857.99 | 5586.6885
2024-12-06 | C | NA | 4769.95 | 5485.4425
2024-12-06 | C | NA | 4908.52 | 5644.798
2024-12-06 | C | NA | 4854.07 | 5582.1805
2024-12-09 | C | 10.25 | 7398.91
2024-12-09 | C | 10.25 | 7301.53
2024-12-09 | C | 10.25 | 7378.62
2024-12-09 | C | 10.25 | 7327.16
2024-12-09 | C | 10.25 | 7420.69
2024-12-09 | C | 10.25 | 7252.76
2024-12-09 | C | 10.25 | 7380.45
2024-12-09 | C | 10.25 | 7323.68
2024-12-09 | C | 10.25 | 7348.58
2024-12-09 | C | 10.25 | 7345.9
2024-12-09 | C | 10.25 | 7300.95
2024-12-09 | C | 10.25 | 7217.82
2024-12-09 | C | 10.25 | 7399.02
2024-12-09 | C | 10.25 | 7330.57
2024-12-09 | C | 10.25 | 7411.6
2024-12-09 | C | 10.25 | 7256.16
2024-12-09 | C | 10.25 | 7301.27
2024-12-09 | C | 10.25 | 7335.5
2024-12-09 | C | 10.25 | 7308.03
2024-12-09 | C | 10.25 | 7320.74
2024-12-09 | C | 8.2 | 6361.75
2024-12-09 | C | 8.2 | 6340.33
2024-12-09 | C | 8.2 | 6299.88
2024-12-09 | C | 8.2 | 6299.5
2024-12-09 | C | 8.2 | 6269.71
2024-12-09 | C | 8.2 | 6242.37
2024-12-09 | C | 8.2 | 6272.85
2024-12-09 | C | 8.2 | 6328.02
2024-12-09 | C | 8.2 | 6292.47
2024-12-09 | C | 8.2 | 6201.04
2024-12-09 | C | 8.2 | 6247.56
2024-12-09 | C | 8.2 | 6286.52
2024-12-09 | C | 8.2 | 6285
2024-12-09 | C | 8.2 | 6302.76
2024-12-09 | C | 8.2 | 6221.92
2024-12-09 | C | 8.2 | 6257.99
2024-12-09 | C | 8.2 | 6245.83
2024-12-09 | C | 8.2 | 6250.98
2024-12-09 | C | 8.2 | 6212.83
2024-12-09 | C | 8.2 | 6182.83
2024-12-09 | C | 6.15 | 4827.94
2024-12-09 | C | 6.15 | 4959.3
2024-12-09 | C | 6.15 | 4886.78
2024-12-09 | C | 6.15 | 4894.09
2024-12-09 | C | 6.15 | 4888.04
2024-12-09 | C | 6.15 | 4935.97
2024-12-09 | C | 6.15 | 4884.61
2024-12-09 | C | 6.15 | 4972.69
2024-12-09 | C | 6.15 | 4945.65
2024-12-09 | C | 6.15 | 4895.96
2024-12-09 | C | 6.15 | 4873.26
2024-12-09 | C | 6.15 | 4981.63
2024-12-09 | C | 6.15 | 4861.09
2024-12-09 | C | 6.15 | 4870.86
2024-12-09 | C | 6.15 | 4954.98
2024-12-09 | C | 6.15 | 4939
2024-12-09 | C | 6.15 | 4875.36
2024-12-09 | C | 6.15 | 4951.81
2024-12-09 | C | 6.15 | 4968.83
2024-12-09 | C | 6.15 | 4959.1
2024-12-09 | C | 4.1 | 3337.51
2024-12-09 | C | 4.1 | 3326.63
2024-12-09 | C | 4.1 | 3267.49
2024-12-09 | C | 4.1 | 3364.08
2024-12-09 | C | 4.1 | 3311.64
2024-12-09 | C | 4.1 | 3323.44
2024-12-09 | C | 4.1 | 3360.02
2024-12-09 | C | 4.1 | 3343.36
2024-12-09 | C | 4.1 | 3251.13
2024-12-09 | C | 4.1 | 3270.45
2024-12-09 | C | 4.1 | 3364.55
2024-12-09 | C | 4.1 | 3269.82
2024-12-09 | C | 4.1 | 3321.48
2024-12-09 | C | 4.1 | 3298.95
2024-12-09 | C | 4.1 | 3284.47
2024-12-09 | C | 4.1 | 3398.23
2024-12-09 | C | 4.1 | 3335.71
2024-12-09 | C | 4.1 | 3308.56
2024-12-09 | C | 4.1 | 3299.6
2024-12-09 | C | 4.1 | 3301.9
2024-12-09 | C | 2.05 | 1707.44
2024-12-09 | C | 2.05 | 1718.67
2024-12-09 | C | 2.05 | 1752.64
2024-12-09 | C | 2.05 | 1706.57
2024-12-09 | C | 2.05 | 1743.92
2024-12-09 | C | 2.05 | 1782.26
2024-12-09 | C | 2.05 | 1733.69
2024-12-09 | C | 2.05 | 1743.56
2024-12-09 | C | 2.05 | 1733.85
2024-12-09 | C | 2.05 | 1742.64
2024-12-09 | C | 2.05 | 1719.95
2024-12-09 | C | 2.05 | 1735
2024-12-09 | C | 2.05 | 1719.51
2024-12-09 | C | 2.05 | 1713.76
2024-12-09 | C | 2.05 | 1724.13
2024-12-09 | C | 2.05 | 1694.19
2024-12-09 | C | 2.05 | 1739.21
2024-12-09 | C | 2.05 | 1750.86
2024-12-09 | C | 2.05 | 1742.2
2024-12-09 | C | 2.05 | 1749
2024-12-09 | C | 0 | 112.53
2024-12-09 | C | 0 | 121.13
2024-12-09 | C | 0 | 120.56
2024-12-09 | C | 0 | 113.89
2024-12-09 | C | 0 | 123.75
2024-12-09 | C | 0 | 114.32
2024-12-09 | C | 0 | 113.6
2024-12-09 | C | 0 | 119.76
2024-12-09 | C | 0 | 121.64
2024-12-09 | C | 0 | 116.17
2024-12-09 | C | 0 | 119.49
2024-12-09 | C | 0 | 111.65
2024-12-09 | C | 0 | 122.48
2024-12-09 | C | 0 | 120.57
2024-12-09 | C | 0 | 114.78
2024-12-09 | C | 0 | 120.48
2024-12-09 | C | 0 | 118.67
2024-12-09 | C | 0 | 125.12
2024-12-09 | C | 0 | 118.97
2024-12-09 | C | 0 | 112.66
2024-12-09 | C | NA | 4803.59 | 5524.1285
2024-12-09 | C | NA | 4924.56 | 5663.244
2024-12-09 | C | NA | 4837.02 | 5562.573
2024-12-09 | C | NA | 4829.65 | 5554.0975
2024-12-09 | C | NA | 4932.54 | 5672.421
2024-12-09 | C | NA | 4940.35 | 5681.4025
2024-12-09 | C | NA | 4835.95 | 5561.3425
2024-12-09 | C | NA | 4882.83 | 5615.2545
2024-12-09 | C | NA | 4907.76 | 5643.924
2024-12-09 | C | NA | 4966.73 | 5711.7395
2024-12-09 | C | NA | 4909.24 | 5645.626
2024-12-09 | C | NA | 4871.58 | 5602.317
2024-12-09 | C | NA | 4953.39 | 5696.3985
2024-12-09 | C | NA | 4927.27 | 5666.3605
2024-12-09 | C | NA | 4855.69 | 5584.0435
2024-12-09 | C | NA | 4879.05 | 5610.9075
2024-12-09 | C | NA | 4842.27 | 5568.6105
2024-12-09 | C | NA | 4883.39 | 5615.8985
2024-12-09 | C | NA | 4875.88 | 5607.262
2024-12-09 | C | NA | 4860.93 | 5590.0695
2024-12-10 | C | 10.25 | 7333.64
2024-12-10 | C | 10.25 | 7358.49
2024-12-10 | C | 10.25 | 7313.83
2024-12-10 | C | 10.25 | 7279.78
2024-12-10 | C | 10.25 | 7297.16
2024-12-10 | C | 10.25 | 7353.46
2024-12-10 | C | 10.25 | 7387.81
2024-12-10 | C | 10.25 | 7253.4
2024-12-10 | C | 10.25 | 7293.89
2024-12-10 | C | 10.25 | 7314.95
2024-12-10 | C | 10.25 | 7284.33
2024-12-10 | C | 10.25 | 7306.27
2024-12-10 | C | 10.25 | 7287.42
2024-12-10 | C | 10.25 | 7355.37
2024-12-10 | C | 10.25 | 7318.71
2024-12-10 | C | 10.25 | 7243.11
2024-12-10 | C | 10.25 | 7268.74
2024-12-10 | C | 10.25 | 7403.86
2024-12-10 | C | 10.25 | 7397.9
2024-12-10 | C | 10.25 | 7404.08
2024-12-10 | C | 8.2 | 6294.5
2024-12-10 | C | 8.2 | 6302.67
2024-12-10 | C | 8.2 | 6355.78
2024-12-10 | C | 8.2 | 6312.3
2024-12-10 | C | 8.2 | 6254.04
2024-12-10 | C | 8.2 | 6353.58
2024-12-10 | C | 8.2 | 6227
2024-12-10 | C | 8.2 | 6230.81
2024-12-10 | C | 8.2 | 6208.71
2024-12-10 | C | 8.2 | 6311.67
2024-12-10 | C | 8.2 | 6299.96
2024-12-10 | C | 8.2 | 6226.88
2024-12-10 | C | 8.2 | 6242.89
2024-12-10 | C | 8.2 | 6286.56
2024-12-10 | C | 8.2 | 6240.26
2024-12-10 | C | 8.2 | 6237.53
2024-12-10 | C | 8.2 | 6184.27
2024-12-10 | C | 8.2 | 6323.4
2024-12-10 | C | 8.2 | 6249.62
2024-12-10 | C | 8.2 | 6269.16
2024-12-10 | C | 6.15 | 4875.65
2024-12-10 | C | 6.15 | 4896.49
2024-12-10 | C | 6.15 | 4879.36
2024-12-10 | C | 6.15 | 5006.91
2024-12-10 | C | 6.15 | 4952.82
2024-12-10 | C | 6.15 | 4848.49
2024-12-10 | C | 6.15 | 4911.32
2024-12-10 | C | 6.15 | 4938.16
2024-12-10 | C | 6.15 | 4925.11
2024-12-10 | C | 6.15 | 4904.18
2024-12-10 | C | 6.15 | 4872.35
2024-12-10 | C | 6.15 | 4878.53
2024-12-10 | C | 6.15 | 4971.49
2024-12-10 | C | 6.15 | 4868.94
2024-12-10 | C | 6.15 | 4910.71
2024-12-10 | C | 6.15 | 4904.29
2024-12-10 | C | 6.15 | 5003.71
2024-12-10 | C | 6.15 | 4903.94
2024-12-10 | C | 6.15 | 4939.92
2024-12-10 | C | 6.15 | 4887.56
2024-12-10 | C | 4.1 | 3234.25
2024-12-10 | C | 4.1 | 3342.42
2024-12-10 | C | 4.1 | 3318.19
2024-12-10 | C | 4.1 | 3304.49
2024-12-10 | C | 4.1 | 3344.57
2024-12-10 | C | 4.1 | 3336.34
2024-12-10 | C | 4.1 | 3265.71
2024-12-10 | C | 4.1 | 3329.98
2024-12-10 | C | 4.1 | 3369.7
2024-12-10 | C | 4.1 | 3338.84
2024-12-10 | C | 4.1 | 3273.51
2024-12-10 | C | 4.1 | 3293.18
2024-12-10 | C | 4.1 | 3257.09
2024-12-10 | C | 4.1 | 3307.43
2024-12-10 | C | 4.1 | 3288.11
2024-12-10 | C | 4.1 | 3316.99
2024-12-10 | C | 4.1 | 3296.02
2024-12-10 | C | 4.1 | 3274.79
2024-12-10 | C | 4.1 | 3400.63
2024-12-10 | C | 4.1 | 3344.11
2024-12-10 | C | 2.05 | 1796.01
2024-12-10 | C | 2.05 | 1797.39
2024-12-10 | C | 2.05 | 1698.9
2024-12-10 | C | 2.05 | 1725.63
2024-12-10 | C | 2.05 | 1770.74
2024-12-10 | C | 2.05 | 1734.43
2024-12-10 | C | 2.05 | 1725.11
2024-12-10 | C | 2.05 | 1782.27
2024-12-10 | C | 2.05 | 1706.29
2024-12-10 | C | 2.05 | 1733.39
2024-12-10 | C | 2.05 | 1761.34
2024-12-10 | C | 2.05 | 1797.45
2024-12-10 | C | 2.05 | 1713.54
2024-12-10 | C | 2.05 | 1720.62
2024-12-10 | C | 2.05 | 1731.79
2024-12-10 | C | 2.05 | 1760.56
2024-12-10 | C | 2.05 | 1741.93
2024-12-10 | C | 2.05 | 1761.94
2024-12-10 | C | 2.05 | 1704.11
2024-12-10 | C | 2.05 | 1761.37
2024-12-10 | C | 0 | 132.41
2024-12-10 | C | 0 | 114.34
2024-12-10 | C | 0 | 110.83
2024-12-10 | C | 0 | 120.81
2024-12-10 | C | 0 | 126.85
2024-12-10 | C | 0 | 116.69
2024-12-10 | C | 0 | 121.12
2024-12-10 | C | 0 | 110.5
2024-12-10 | C | 0 | 120.13
2024-12-10 | C | 0 | 124.09
2024-12-10 | C | 0 | 125.21
2024-12-10 | C | 0 | 123.13
2024-12-10 | C | 0 | 116.36
2024-12-10 | C | 0 | 123.01
2024-12-10 | C | 0 | 115.03
2024-12-10 | C | 0 | 125.46
2024-12-10 | C | 0 | 116.2
2024-12-10 | C | 0 | 119.99
2024-12-10 | C | 0 | 117.94
2024-12-10 | C | 0 | 116.56
2024-12-10 | C | NA | 4867.44 | 5597.556
2024-12-10 | C | NA | 4962.7 | 5707.105
2024-12-10 | C | NA | 4858.42 | 5587.183
2024-12-10 | C | NA | 4876.75 | 5608.2625
2024-12-10 | C | NA | 4984.52 | 5732.198
2024-12-10 | C | NA | 4913.91 | 5650.9965
2024-12-10 | C | NA | 4927.15 | 5666.2225
2024-12-10 | C | NA | 4818.34 | 5541.091
2024-12-10 | C | NA | 4857.69 | 5586.3435
2024-12-10 | C | NA | 4821.69 | 5544.9435
2024-12-10 | C | NA | 4846.83 | 5573.8545
2024-12-10 | C | NA | 4989.46 | 5737.879
2024-12-10 | C | NA | 4863.23 | 5592.7145
2024-12-10 | C | NA | 4843.14 | 5569.611
2024-12-10 | C | NA | 4900.63 | 5635.7245
2024-12-10 | C | NA | 5017.29 | 5769.8835
2024-12-10 | C | NA | 4858.01 | 5586.7115
2024-12-10 | C | NA | 4866.67 | 5596.6705
2024-12-10 | C | NA | 4916.8 | 5654.32
2024-12-10 | C | NA | 4928.3 | 5667.545
2024-12-11 | C | 10.25 | 7251.59
2024-12-11 | C | 10.25 | 7348.24
2024-12-11 | C | 10.25 | 7317.27
2024-12-11 | C | 10.25 | 7306.66
2024-12-11 | C | 10.25 | 7345.04
2024-12-11 | C | 10.25 | 7307.33
2024-12-11 | C | 10.25 | 7370.53
2024-12-11 | C | 10.25 | 7352.03
2024-12-11 | C | 10.25 | 7430.12
2024-12-11 | C | 10.25 | 7285.24
2024-12-11 | C | 10.25 | 7363.99
2024-12-11 | C | 10.25 | 7309.42
2024-12-11 | C | 10.25 | 7448.62
2024-12-11 | C | 10.25 | 7296.34
2024-12-11 | C | 10.25 | 7313.15
2024-12-11 | C | 10.25 | 7442.81
2024-12-11 | C | 10.25 | 7325.09
2024-12-11 | C | 10.25 | 7430.65
2024-12-11 | C | 10.25 | 7332.88
2024-12-11 | C | 10.25 | 7326.48
2024-12-11 | C | 8.2 | 6328.48
2024-12-11 | C | 8.2 | 6260.58
2024-12-11 | C | 8.2 | 6232.03
2024-12-11 | C | 8.2 | 6339.18
2024-12-11 | C | 8.2 | 6209.63
2024-12-11 | C | 8.2 | 6233.88
2024-12-11 | C | 8.2 | 6249.74
2024-12-11 | C | 8.2 | 6375.91
2024-12-11 | C | 8.2 | 6261.24
2024-12-11 | C | 8.2 | 6189.71
2024-12-11 | C | 8.2 | 6279.42
2024-12-11 | C | 8.2 | 6302.41
2024-12-11 | C | 8.2 | 6281.11
2024-12-11 | C | 8.2 | 6312.95
2024-12-11 | C | 8.2 | 6256.87
2024-12-11 | C | 8.2 | 6310.23
2024-12-11 | C | 8.2 | 6275.04
2024-12-11 | C | 8.2 | 6271.26
2024-12-11 | C | 8.2 | 6384.88
2024-12-11 | C | 8.2 | 6337.38
2024-12-11 | C | 6.15 | 4967.39
2024-12-11 | C | 6.15 | 4900.28
2024-12-11 | C | 6.15 | 4923.77
2024-12-11 | C | 6.15 | 4914.52
2024-12-11 | C | 6.15 | 4860.66
2024-12-11 | C | 6.15 | 4862.43
2024-12-11 | C | 6.15 | 4917.22
2024-12-11 | C | 6.15 | 4907.39
2024-12-11 | C | 6.15 | 4870.05
2024-12-11 | C | 6.15 | 4882.05
2024-12-11 | C | 6.15 | 4923.09
2024-12-11 | C | 6.15 | 4874.03
2024-12-11 | C | 6.15 | 4936.82
2024-12-11 | C | 6.15 | 4942.84
2024-12-11 | C | 6.15 | 4915.98
2024-12-11 | C | 6.15 | 4958.99
2024-12-11 | C | 6.15 | 4926.46
2024-12-11 | C | 6.15 | 4866.76
2024-12-11 | C | 6.15 | 4848
2024-12-11 | C | 6.15 | 4915.05
2024-12-11 | C | 4.1 | 3338.44
2024-12-11 | C | 4.1 | 3315.64
2024-12-11 | C | 4.1 | 3376.54
2024-12-11 | C | 4.1 | 3359
2024-12-11 | C | 4.1 | 3327.03
2024-12-11 | C | 4.1 | 3326.54
2024-12-11 | C | 4.1 | 3327.42
2024-12-11 | C | 4.1 | 3309
2024-12-11 | C | 4.1 | 3417.91
2024-12-11 | C | 4.1 | 3366.27
2024-12-11 | C | 4.1 | 3329.28
2024-12-11 | C | 4.1 | 3374.16
2024-12-11 | C | 4.1 | 3286.14
2024-12-11 | C | 4.1 | 3401.24
2024-12-11 | C | 4.1 | 3295.52
2024-12-11 | C | 4.1 | 3371.55
2024-12-11 | C | 4.1 | 3323.66
2024-12-11 | C | 4.1 | 3335.55
2024-12-11 | C | 4.1 | 3321.4
2024-12-11 | C | 4.1 | 3355.9
2024-12-11 | C | 2.05 | 1700.89
2024-12-11 | C | 2.05 | 1719.44
2024-12-11 | C | 2.05 | 1784.71
2024-12-11 | C | 2.05 | 1811.56
2024-12-11 | C | 2.05 | 1768.15
2024-12-11 | C | 2.05 | 1809.77
2024-12-11 | C | 2.05 | 1737.52
2024-12-11 | C | 2.05 | 1731.97
2024-12-11 | C | 2.05 | 1783.45
2024-12-11 | C | 2.05 | 1739.89
2024-12-11 | C | 2.05 | 1777.5
2024-12-11 | C | 2.05 | 1745.11
2024-12-11 | C | 2.05 | 1738.87
2024-12-11 | C | 2.05 | 1737.47
2024-12-11 | C | 2.05 | 1749.46
2024-12-11 | C | 2.05 | 1742.52
2024-12-11 | C | 2.05 | 1745.62
2024-12-11 | C | 2.05 | 1740.17
2024-12-11 | C | 2.05 | 1734.83
2024-12-11 | C | 2.05 | 1786.15
2024-12-11 | C | 0 | 109.43
2024-12-11 | C | 0 | 119.03
2024-12-11 | C | 0 | 118.11
2024-12-11 | C | 0 | 125.39
2024-12-11 | C | 0 | 120.89
2024-12-11 | C | 0 | 110.38
2024-12-11 | C | 0 | 114.93
2024-12-11 | C | 0 | 112.3
2024-12-11 | C | 0 | 123.27
2024-12-11 | C | 0 | 113.5
2024-12-11 | C | 0 | 113.73
2024-12-11 | C | 0 | 120.03
2024-12-11 | C | 0 | 112.05
2024-12-11 | C | 0 | 115.71
2024-12-11 | C | 0 | 116.3
2024-12-11 | C | 0 | 124.2
2024-12-11 | C | 0 | 121.95
2024-12-11 | C | 0 | 122.03
2024-12-11 | C | 0 | 122.76
2024-12-11 | C | 0 | 113.44
2024-12-11 | C | NA | 4853.92 | 5582.008
2024-12-11 | C | NA | 4930.21 | 5669.7415
2024-12-11 | C | NA | 4940.02 | 5681.023
2024-12-11 | C | NA | 4850.15 | 5577.6725
2024-12-11 | C | NA | 4805.32 | 5526.118
2024-12-11 | C | NA | 4901.57 | 5636.8055
2024-12-11 | C | NA | 4904.49 | 5640.1635
2024-12-11 | C | NA | 4881.95 | 5614.2425
2024-12-11 | C | NA | 4894.6 | 5628.79
2024-12-11 | C | NA | 4868.78 | 5599.097
2024-12-11 | C | NA | 4828.5 | 5552.775
2024-12-11 | C | NA | 4842.9 | 5569.335
2024-12-11 | C | NA | 4936.19 | 5676.6185
2024-12-11 | C | NA | 4936.86 | 5677.389
2024-12-11 | C | NA | 4883.22 | 5615.703
2024-12-11 | C | NA | 4894.72 | 5628.928
2024-12-11 | C | NA | 4920.68 | 5658.782
2024-12-11 | C | NA | 4920.74 | 5658.851
2024-12-11 | C | NA | 4851.39 | 5579.0985
2024-12-11 | C | NA | 4882.71 | 5615.1165
2024-12-12 | C | 10.25 | 7305.16
2024-12-12 | C | 10.25 | 7313.55
2024-12-12 | C | 10.25 | 7345.61
2024-12-12 | C | 10.25 | 7275.99
2024-12-12 | C | 10.25 | 7350.68
2024-12-12 | C | 10.25 | 7243.99
2024-12-12 | C | 10.25 | 7395.38
2024-12-12 | C | 10.25 | 7373.22
2024-12-12 | C | 10.25 | 7353.65
2024-12-12 | C | 10.25 | 7403.15
2024-12-12 | C | 10.25 | 7345.95
2024-12-12 | C | 10.25 | 7287.49
2024-12-12 | C | 10.25 | 7375.39
2024-12-12 | C | 10.25 | 7414.68
2024-12-12 | C | 10.25 | 7302.96
2024-12-12 | C | 10.25 | 7303.49
2024-12-12 | C | 10.25 | 7390.25
2024-12-12 | C | 10.25 | 7336.14
2024-12-12 | C | 10.25 | 7324.6
2024-12-12 | C | 10.25 | 7429.36
2024-12-12 | C | 8.2 | 6282.18
2024-12-12 | C | 8.2 | 6243.27
2024-12-12 | C | 8.2 | 6265.93
2024-12-12 | C | 8.2 | 6306.72
2024-12-12 | C | 8.2 | 6255.23
2024-12-12 | C | 8.2 | 6239.06
2024-12-12 | C | 8.2 | 6228.27
2024-12-12 | C | 8.2 | 6284.45
2024-12-12 | C | 8.2 | 6249.54
2024-12-12 | C | 8.2 | 6275.97
2024-12-12 | C | 8.2 | 6257.44
2024-12-12 | C | 8.2 | 6266.42
2024-12-12 | C | 8.2 | 6286.44
2024-12-12 | C | 8.2 | 6251.84
2024-12-12 | C | 8.2 | 6278.13
2024-12-12 | C | 8.2 | 6253.01
2024-12-12 | C | 8.2 | 6246.42
2024-12-12 | C | 8.2 | 6298.45
2024-12-12 | C | 8.2 | 6343.44
2024-12-12 | C | 8.2 | 6284.24
2024-12-12 | C | 6.15 | 5003.27
2024-12-12 | C | 6.15 | 4894.47
2024-12-12 | C | 6.15 | 4896.92
2024-12-12 | C | 6.15 | 4913.03
2024-12-12 | C | 6.15 | 4867.36
2024-12-12 | C | 6.15 | 4846.06
2024-12-12 | C | 6.15 | 4933.84
2024-12-12 | C | 6.15 | 4921.82
2024-12-12 | C | 6.15 | 4905.86
2024-12-12 | C | 6.15 | 4865.27
2024-12-12 | C | 6.15 | 4914.53
2024-12-12 | C | 6.15 | 4918.52
2024-12-12 | C | 6.15 | 4942.18
2024-12-12 | C | 6.15 | 4847.92
2024-12-12 | C | 6.15 | 4881.61
2024-12-12 | C | 6.15 | 4965.87
2024-12-12 | C | 6.15 | 4915.59
2024-12-12 | C | 6.15 | 4908.57
2024-12-12 | C | 6.15 | 4944.11
2024-12-12 | C | 6.15 | 4928.37
2024-12-12 | C | 4.1 | 3339.07
2024-12-12 | C | 4.1 | 3335.69
2024-12-12 | C | 4.1 | 3304.87
2024-12-12 | C | 4.1 | 3334.64
2024-12-12 | C | 4.1 | 3290.69
2024-12-12 | C | 4.1 | 3360.46
2024-12-12 | C | 4.1 | 3353.7
2024-12-12 | C | 4.1 | 3355.63
2024-12-12 | C | 4.1 | 3307.68
2024-12-12 | C | 4.1 | 3291.63
2024-12-12 | C | 4.1 | 3332.17
2024-12-12 | C | 4.1 | 3306.79
2024-12-12 | C | 4.1 | 3365.82
2024-12-12 | C | 4.1 | 3253.28
2024-12-12 | C | 4.1 | 3369.77
2024-12-12 | C | 4.1 | 3253.02
2024-12-12 | C | 4.1 | 3342.8
2024-12-12 | C | 4.1 | 3303.33
2024-12-12 | C | 4.1 | 3381.03
2024-12-12 | C | 4.1 | 3316.97
2024-12-12 | C | 2.05 | 1716.32
2024-12-12 | C | 2.05 | 1745.44
2024-12-12 | C | 2.05 | 1674.52
2024-12-12 | C | 2.05 | 1721.2
2024-12-12 | C | 2.05 | 1702.65
2024-12-12 | C | 2.05 | 1769.81
2024-12-12 | C | 2.05 | 1729.76
2024-12-12 | C | 2.05 | 1721.67
2024-12-12 | C | 2.05 | 1723.35
2024-12-12 | C | 2.05 | 1731.85
2024-12-12 | C | 2.05 | 1726.02
2024-12-12 | C | 2.05 | 1791.55
2024-12-12 | C | 2.05 | 1754.27
2024-12-12 | C | 2.05 | 1667.99
2024-12-12 | C | 2.05 | 1720.63
2024-12-12 | C | 2.05 | 1729.58
2024-12-12 | C | 2.05 | 1695.61
2024-12-12 | C | 2.05 | 1763.03
2024-12-12 | C | 2.05 | 1771.09
2024-12-12 | C | 2.05 | 1781.54
2024-12-12 | C | 0 | 116.4
2024-12-12 | C | 0 | 122.26
2024-12-12 | C | 0 | 119.1
2024-12-12 | C | 0 | 122.9
2024-12-12 | C | 0 | 118.4
2024-12-12 | C | 0 | 114.78
2024-12-12 | C | 0 | 114.08
2024-12-12 | C | 0 | 119.28
2024-12-12 | C | 0 | 121.97
2024-12-12 | C | 0 | 104.16
2024-12-12 | C | 0 | 113.67
2024-12-12 | C | 0 | 119.5
2024-12-12 | C | 0 | 116.86
2024-12-12 | C | 0 | 111.26
2024-12-12 | C | 0 | 115.77
2024-12-12 | C | 0 | 111.58
2024-12-12 | C | 0 | 116.27
2024-12-12 | C | 0 | 123.85
2024-12-12 | C | 0 | 118.49
2024-12-12 | C | 0 | 117.53
2024-12-12 | C | NA | 4848.3 | 5575.545
2024-12-12 | C | NA | 4895.67 | 5630.0205
2024-12-12 | C | NA | 4919.87 | 5657.8505
2024-12-12 | C | NA | 4810.93 | 5532.5695
2024-12-12 | C | NA | 4899.85 | 5634.8275
2024-12-12 | C | NA | 4846.15 | 5573.0725
2024-12-12 | C | NA | 4818.57 | 5541.3555
2024-12-12 | C | NA | 4847.48 | 5574.602
2024-12-12 | C | NA | 4812.4 | 5534.26
2024-12-12 | C | NA | 4907.89 | 5644.0735
2024-12-12 | C | NA | 4838.69 | 5564.4935
2024-12-12 | C | NA | 4904.45 | 5640.1175
2024-12-12 | C | NA | 4854.56 | 5582.744
2024-12-12 | C | NA | 4802.97 | 5523.4155
2024-12-12 | C | NA | 4858.35 | 5587.1025
2024-12-12 | C | NA | 4915.81 | 5653.1815
2024-12-12 | C | NA | 4770.7 | 5486.305
2024-12-12 | C | NA | 4798.78 | 5518.597
2024-12-12 | C | NA | 4761.47 | 5475.6905
2024-12-12 | C | NA | 4899.29 | 5634.1835
2024-12-13 | C | 10.25 | 7265.14
2024-12-13 | C | 10.25 | 7360.89
2024-12-13 | C | 10.25 | 7388.48
2024-12-13 | C | 10.25 | 7286.25
2024-12-13 | C | 10.25 | 7475.62
2024-12-13 | C | 10.25 | 7256.37
2024-12-13 | C | 10.25 | 7334.46
2024-12-13 | C | 10.25 | 7258.7
2024-12-13 | C | 10.25 | 7247.13
2024-12-13 | C | 10.25 | 7341.21
2024-12-13 | C | 10.25 | 7412.54
2024-12-13 | C | 10.25 | 7294.86
2024-12-13 | C | 10.25 | 7251.25
2024-12-13 | C | 10.25 | 7347.15
2024-12-13 | C | 10.25 | 7335.49
2024-12-13 | C | 10.25 | 7296.73
2024-12-13 | C | 10.25 | 7265.77
2024-12-13 | C | 10.25 | 7263.91
2024-12-13 | C | 10.25 | 7336.33
2024-12-13 | C | 10.25 | 7377.93
2024-12-13 | C | 8.2 | 6308.79
2024-12-13 | C | 8.2 | 6218.89
2024-12-13 | C | 8.2 | 6215.85
2024-12-13 | C | 8.2 | 6355.94
2024-12-13 | C | 8.2 | 6303.11
2024-12-13 | C | 8.2 | 6313.32
2024-12-13 | C | 8.2 | 6241.22
2024-12-13 | C | 8.2 | 6246.5
2024-12-13 | C | 8.2 | 6345.62
2024-12-13 | C | 8.2 | 6290.56
2024-12-13 | C | 8.2 | 6258.08
2024-12-13 | C | 8.2 | 6317.16
2024-12-13 | C | 8.2 | 6294.23
2024-12-13 | C | 8.2 | 6400.78
2024-12-13 | C | 8.2 | 6213.15
2024-12-13 | C | 8.2 | 6197.58
2024-12-13 | C | 8.2 | 6195.32
2024-12-13 | C | 8.2 | 6269
2024-12-13 | C | 8.2 | 6269.89
2024-12-13 | C | 8.2 | 6300.35
2024-12-13 | C | 6.15 | 4963.32
2024-12-13 | C | 6.15 | 4970.92
2024-12-13 | C | 6.15 | 4902.81
2024-12-13 | C | 6.15 | 4928.94
2024-12-13 | C | 6.15 | 4962.26
2024-12-13 | C | 6.15 | 4881.17
2024-12-13 | C | 6.15 | 4843.03
2024-12-13 | C | 6.15 | 4915.48
2024-12-13 | C | 6.15 | 4860.04
2024-12-13 | C | 6.15 | 4933.72
2024-12-13 | C | 6.15 | 4970.98
2024-12-13 | C | 6.15 | 4962.97
2024-12-13 | C | 6.15 | 4962.03
2024-12-13 | C | 6.15 | 4888.81
2024-12-13 | C | 6.15 | 4905.77
2024-12-13 | C | 6.15 | 4901.19
2024-12-13 | C | 6.15 | 4853.29
2024-12-13 | C | 6.15 | 4884.25
2024-12-13 | C | 6.15 | 4923.43
2024-12-13 | C | 6.15 | 4948.38
2024-12-13 | C | 4.1 | 3337.49
2024-12-13 | C | 4.1 | 3328.45
2024-12-13 | C | 4.1 | 3289.39
2024-12-13 | C | 4.1 | 3262.98
2024-12-13 | C | 4.1 | 3257.95
2024-12-13 | C | 4.1 | 3328.11
2024-12-13 | C | 4.1 | 3295.96
2024-12-13 | C | 4.1 | 3318.82
2024-12-13 | C | 4.1 | 3311.1
2024-12-13 | C | 4.1 | 3330.45
2024-12-13 | C | 4.1 | 3346.51
2024-12-13 | C | 4.1 | 3316.46
2024-12-13 | C | 4.1 | 3340.02
2024-12-13 | C | 4.1 | 3305.22
2024-12-13 | C | 4.1 | 3341.59
2024-12-13 | C | 4.1 | 3306.9
2024-12-13 | C | 4.1 | 3350.96
2024-12-13 | C | 4.1 | 3319.89
2024-12-13 | C | 4.1 | 3364.44
2024-12-13 | C | 4.1 | 3323.63
2024-12-13 | C | 2.05 | 1705.42
2024-12-13 | C | 2.05 | 1719.92
2024-12-13 | C | 2.05 | 1711.87
2024-12-13 | C | 2.05 | 1750.44
2024-12-13 | C | 2.05 | 1710.93
2024-12-13 | C | 2.05 | 1752.23
2024-12-13 | C | 2.05 | 1792.76
2024-12-13 | C | 2.05 | 1712.74
2024-12-13 | C | 2.05 | 1789.78
2024-12-13 | C | 2.05 | 1741.48
2024-12-13 | C | 2.05 | 1767.27
2024-12-13 | C | 2.05 | 1697.35
2024-12-13 | C | 2.05 | 1742.32
2024-12-13 | C | 2.05 | 1731.52
2024-12-13 | C | 2.05 | 1742.82
2024-12-13 | C | 2.05 | 1733.3
2024-12-13 | C | 2.05 | 1726.48
2024-12-13 | C | 2.05 | 1749.58
2024-12-13 | C | 2.05 | 1724.42
2024-12-13 | C | 2.05 | 1756.65
2024-12-13 | C | 0 | 119.51
2024-12-13 | C | 0 | 110.99
2024-12-13 | C | 0 | 124.83
2024-12-13 | C | 0 | 119.26
2024-12-13 | C | 0 | 128.57
2024-12-13 | C | 0 | 127.92
2024-12-13 | C | 0 | 114.3
2024-12-13 | C | 0 | 116.82
2024-12-13 | C | 0 | 115.98
2024-12-13 | C | 0 | 116.4
2024-12-13 | C | 0 | 116.35
2024-12-13 | C | 0 | 122.23
2024-12-13 | C | 0 | 117.35
2024-12-13 | C | 0 | 116.09
2024-12-13 | C | 0 | 120.67
2024-12-13 | C | 0 | 119.7
2024-12-13 | C | 0 | 125.99
2024-12-13 | C | 0 | 127.35
2024-12-13 | C | 0 | 126.89
2024-12-13 | C | 0 | 126.22
2024-12-13 | C | NA | 4921.55 | 5659.7825
2024-12-13 | C | NA | 4930.23 | 5669.7645
2024-12-13 | C | NA | 4927.9 | 5667.085
2024-12-13 | C | NA | 4755.95 | 5469.3425
2024-12-13 | C | NA | 4923.41 | 5661.9215
2024-12-13 | C | NA | 4850.47 | 5578.0405
2024-12-13 | C | NA | 4853.39 | 5581.3985
2024-12-13 | C | NA | 4921.97 | 5660.2655
2024-12-13 | C | NA | 4989.95 | 5738.4425
2024-12-13 | C | NA | 4886.85 | 5619.8775
2024-12-13 | C | NA | 4937.7 | 5678.355
2024-12-13 | C | NA | 4957.38 | 5700.987
2024-12-13 | C | NA | 4859.42 | 5588.333
2024-12-13 | C | NA | 4818.36 | 5541.114
2024-12-13 | C | NA | 4762.06 | 5476.369
2024-12-13 | C | NA | 4742.42 | 5453.783
2024-12-13 | C | NA | 4800.98 | 5521.127
2024-12-13 | C | NA | 4877.5 | 5609.125
2024-12-13 | C | NA | 4841.38 | 5567.587
2024-12-13 | C | NA | 4864.85 | 5594.5775
2024-12-14 | C | 10.25 | 7290.35
2024-12-14 | C | 10.25 | 7269.79
2024-12-14 | C | 10.25 | 7342.57
2024-12-14 | C | 10.25 | 7240.96
2024-12-14 | C | 10.25 | 7339.71
2024-12-14 | C | 10.25 | 7297.77
2024-12-14 | C | 10.25 | 7305.41
2024-12-14 | C | 10.25 | 7261.01
2024-12-14 | C | 10.25 | 7375.85
2024-12-14 | C | 10.25 | 7292.83
2024-12-14 | C | 10.25 | 7369.72
2024-12-14 | C | 10.25 | 7363.9
2024-12-14 | C | 10.25 | 7393.17
2024-12-14 | C | 10.25 | 7278.87
2024-12-14 | C | 10.25 | 7412.99
2024-12-14 | C | 10.25 | 7395.95
2024-12-14 | C | 10.25 | 7299.7
2024-12-14 | C | 10.25 | 7316.89
2024-12-14 | C | 10.25 | 7238.07
2024-12-14 | C | 10.25 | 7320.65
2024-12-14 | C | 8.2 | 6303.01
2024-12-14 | C | 8.2 | 6305.97
2024-12-14 | C | 8.2 | 6255.62
2024-12-14 | C | 8.2 | 6202.59
2024-12-14 | C | 8.2 | 6326.03
2024-12-14 | C | 8.2 | 6309.65
2024-12-14 | C | 8.2 | 6249.11
2024-12-14 | C | 8.2 | 6224.61
2024-12-14 | C | 8.2 | 6254.72
2024-12-14 | C | 8.2 | 6269.59
2024-12-14 | C | 8.2 | 6285.93
2024-12-14 | C | 8.2 | 6279.21
2024-12-14 | C | 8.2 | 6258.46
2024-12-14 | C | 8.2 | 6308.73
2024-12-14 | C | 8.2 | 6334.66
2024-12-14 | C | 8.2 | 6256.08
2024-12-14 | C | 8.2 | 6326.67
2024-12-14 | C | 8.2 | 6264.89
2024-12-14 | C | 8.2 | 6278.35
2024-12-14 | C | 8.2 | 6308.7
2024-12-14 | C | 6.15 | 4869.68
2024-12-14 | C | 6.15 | 4958.3
2024-12-14 | C | 6.15 | 4960.16
2024-12-14 | C | 6.15 | 4942.78
2024-12-14 | C | 6.15 | 4880.99
2024-12-14 | C | 6.15 | 4888.79
2024-12-14 | C | 6.15 | 4962.65
2024-12-14 | C | 6.15 | 4913.59
2024-12-14 | C | 6.15 | 4906.87
2024-12-14 | C | 6.15 | 4951.97
2024-12-14 | C | 6.15 | 4902.5
2024-12-14 | C | 6.15 | 4883.3
2024-12-14 | C | 6.15 | 4929.16
2024-12-14 | C | 6.15 | 4891.35
2024-12-14 | C | 6.15 | 4890.97
2024-12-14 | C | 6.15 | 4998.14
2024-12-14 | C | 6.15 | 4946.35
2024-12-14 | C | 6.15 | 5016
2024-12-14 | C | 6.15 | 4908.39
2024-12-14 | C | 6.15 | 4905.89
2024-12-14 | C | 4.1 | 3294.22
2024-12-14 | C | 4.1 | 3343.63
2024-12-14 | C | 4.1 | 3244.99
2024-12-14 | C | 4.1 | 3289.06
2024-12-14 | C | 4.1 | 3285.9
2024-12-14 | C | 4.1 | 3307.37
2024-12-14 | C | 4.1 | 3265.27
2024-12-14 | C | 4.1 | 3340.98
2024-12-14 | C | 4.1 | 3321.86
2024-12-14 | C | 4.1 | 3320.56
2024-12-14 | C | 4.1 | 3331.76
2024-12-14 | C | 4.1 | 3252.75
2024-12-14 | C | 4.1 | 3313.1
2024-12-14 | C | 4.1 | 3330.57
2024-12-14 | C | 4.1 | 3291.53
2024-12-14 | C | 4.1 | 3295.2
2024-12-14 | C | 4.1 | 3291.76
2024-12-14 | C | 4.1 | 3285.19
2024-12-14 | C | 4.1 | 3333.73
2024-12-14 | C | 4.1 | 3317.97
2024-12-14 | C | 2.05 | 1754.36
2024-12-14 | C | 2.05 | 1735.39
2024-12-14 | C | 2.05 | 1722.86
2024-12-14 | C | 2.05 | 1766.98
2024-12-14 | C | 2.05 | 1726.01
2024-12-14 | C | 2.05 | 1714.68
2024-12-14 | C | 2.05 | 1800.97
2024-12-14 | C | 2.05 | 1817.04
2024-12-14 | C | 2.05 | 1732.58
2024-12-14 | C | 2.05 | 1780.36
2024-12-14 | C | 2.05 | 1738.5
2024-12-14 | C | 2.05 | 1737.46
2024-12-14 | C | 2.05 | 1729.02
2024-12-14 | C | 2.05 | 1778.02
2024-12-14 | C | 2.05 | 1723.35
2024-12-14 | C | 2.05 | 1757.38
2024-12-14 | C | 2.05 | 1756.44
2024-12-14 | C | 2.05 | 1743.05
2024-12-14 | C | 2.05 | 1745.67
2024-12-14 | C | 2.05 | 1783.07
2024-12-14 | C | 0 | 114.02
2024-12-14 | C | 0 | 124.98
2024-12-14 | C | 0 | 123.59
2024-12-14 | C | 0 | 107.16
2024-12-14 | C | 0 | 121.53
2024-12-14 | C | 0 | 121.87
2024-12-14 | C | 0 | 123.04
2024-12-14 | C | 0 | 115.05
2024-12-14 | C | 0 | 120.11
2024-12-14 | C | 0 | 120.65
2024-12-14 | C | 0 | 119.78
2024-12-14 | C | 0 | 121.45
2024-12-14 | C | 0 | 121.39
2024-12-14 | C | 0 | 109.79
2024-12-14 | C | 0 | 114.1
2024-12-14 | C | 0 | 106.69
2024-12-14 | C | 0 | 114.85
2024-12-14 | C | 0 | 121.59
2024-12-14 | C | 0 | 116.52
2024-12-14 | C | 0 | 118.53
2024-12-14 | C | NA | 4896.05 | 5630.4575
2024-12-14 | C | NA | 4853.63
2024-12-14 | C | NA | 4931.15
2024-12-14 | C | NA | 4880.8
2024-12-14 | C | NA | 4899.29
2024-12-14 | C | NA | 4835.19
2024-12-14 | C | NA | 4930.01
2024-12-14 | C | NA | 4930.61
2024-12-14 | C | NA | 4862.92
2024-12-14 | C | NA | 4857.65
2024-12-14 | C | NA | 4835.31
2024-12-14 | C | NA | 4798.54
2024-12-14 | C | NA | 4931.6
2024-12-14 | C | NA | 4923.18
2024-12-14 | C | NA | 5013.96
2024-12-14 | C | NA | 4902.61
2024-12-14 | C | NA | 4897.9
2024-12-14 | C | NA | 4814.68
2024-12-14 | C | NA | 4801.06
2024-12-14 | C | NA | 4926.21

## time 1h30
Column1 | Column2
TITLE
DATA TYPE
ORIGIN | JASCO
OWNER
DATE | 24/12/11
TIME | 14:36:00
SPECTROMETER/DATA SYSTEM | JASCO Corp,, FP-8200, Rev, 1,00
RESOLUTION
DELTAX | 1
XUNITS | SEC
YUNITS | INTENSITY
FIRSTX | 0
LASTX | 5400
NPOINTS | 5401
FIRSTY | 7225.85
MAXY | 7252.73
MINY | 7122.72
XYDATA
0 | 7225.85
1 | 7231.44
2 | 7230.43
3 | 7226.53
4 | 7228.91
5 | 7216.61
6 | 7220.57
7 | 7235.04
8 | 7227.83
9 | 7244.57
10 | 7232.91
11 | 7210.18
12 | 7248.79
13 | 7237.02
14 | 7236.52
15 | 7223.18
16 | 7228.71
17 | 7227.41
18 | 7217.83
19 | 7227.72
20 | 7234.8
21 | 7229.82
22 | 7211.97
23 | 7237.99
24 | 7221.6
25 | 7222.4
26 | 7219.42
27 | 7230.77
28 | 7243.43
29 | 7227.31
30 | 7245.77
31 | 7238.58
32 | 7219.39
33 | 7218.94
34 | 7221.86
35 | 7229.73
36 | 7232.78
37 | 7228.26
38 | 7238.65
39 | 7229.48
40 | 7224.91
41 | 7228.96
42 | 7221.91
43 | 7220.31
44 | 7224.97
45 | 7230.98
46 | 7229.03
47 | 7230.36
48 | 7232.7
49 | 7217.61
50 | 7221.32
51 | 7227.16
52 | 7231.55
53 | 7223
54 | 7220.28
55 | 7221.02
56 | 7233.3
57 | 7225.83
58 | 7211.97
59 | 7233.23
60 | 7233.53
61 | 7222.06
62 | 7236.24
63 | 7227.3
64 | 7249.17
65 | 7220.72
66 | 7200.44
67 | 7239.52
68 | 7223.33
69 | 7228.79
70 | 7233.79
71 | 7239.02
72 | 7216.64
73 | 7238.94
74 | 7214.28
75 | 7244.62
76 | 7235.21
77 | 7243.89
78 | 7221.54
79 | 7237.76
80 | 7226.42
81 | 7207.32
82 | 7227.03
83 | 7219.9
84 | 7218.15
85 | 7210.27
86 | 7228.08
87 | 7234.93
88 | 7230.6
89 | 7219.32
90 | 7231.67
91 | 7209.59
92 | 7206.17
93 | 7229.13
94 | 7218.46
95 | 7219.15
96 | 7218.53
97 | 7218.76
98 | 7211.76
99 | 7213.41
100 | 7223.68
101 | 7201.73
102 | 7228.61
103 | 7209.51
104 | 7218.29
105 | 7216.82
106 | 7230.96
107 | 7224.79
108 | 7223.24
109 | 7224.31
110 | 7222.57
111 | 7241.67
112 | 7225.64
113 | 7223.16
114 | 7221.47
115 | 7245.44
116 | 7233.83
117 | 7218.28
118 | 7223.45
119 | 7218.84
120 | 7219.02
121 | 7226.92
122 | 7228.49
123 | 7214.65
124 | 7212.24
125 | 7222.14
126 | 7216.87
127 | 7212.44
128 | 7194.49
129 | 7230.58
130 | 7220.56
131 | 7226.09
132 | 7216.58
133 | 7231.29
134 | 7224.97
135 | 7235.44
136 | 7213.88
137 | 7213.95
138 | 7224.29
139 | 7213.53
140 | 7245.75
141 | 7208.81
142 | 7215.65
143 | 7229.67
144 | 7232.39
145 | 7242.34
146 | 7211.98
147 | 7240.47
148 | 7215.06
149 | 7223.63
150 | 7225.16
151 | 7234.63
152 | 7216.92
153 | 7224.45
154 | 7223.57
155 | 7213.87
156 | 7219.99
157 | 7207.75
158 | 7213.66
159 | 7222.11
160 | 7235.52
161 | 7221.26
162 | 7221.97
163 | 7225.13
164 | 7223.32
165 | 7219.32
166 | 7216.97
167 | 7200.72
168 | 7207.71
169 | 7210.85
170 | 7209.51
171 | 7216.84
172 | 7236.29
173 | 7232.85
174 | 7212.2
175 | 7203.61
176 | 7214.15
177 | 7220.54
178 | 7235.02
179 | 7213.34
180 | 7215.39
181 | 7228.7
182 | 7216.7
183 | 7210.43
184 | 7223.83
185 | 7208.67
186 | 7208.81
187 | 7223.85
188 | 7230.96
189 | 7207.96
190 | 7228.78
191 | 7211.34
192 | 7218.13
193 | 7205.75
194 | 7207.25
195 | 7222.42
196 | 7225.39
197 | 7223.01
198 | 7223.5
199 | 7236.57
200 | 7198.13
201 | 7231.28
202 | 7207.83
203 | 7204.99
204 | 7196.76
205 | 7216.9
206 | 7218.81
207 | 7219.41
208 | 7223.75
209 | 7231.37
210 | 7215.68
211 | 7212.69
212 | 7210.43
213 | 7222.97
214 | 7221.28
215 | 7241.15
216 | 7237.19
217 | 7246.69
218 | 7214.89
219 | 7232.56
220 | 7218.23
221 | 7232.35
222 | 7194.94
223 | 7229.73
224 | 7227.11
225 | 7233.41
226 | 7209.81
227 | 7229.77
228 | 7226.19
229 | 7218.77
230 | 7217.68
231 | 7221.59
232 | 7217.14
233 | 7234.69
234 | 7208.2
235 | 7218.09
236 | 7208
237 | 7216.79
238 | 7230.52
239 | 7241.39
240 | 7225.89
241 | 7220.21
242 | 7224.26
243 | 7206.19
244 | 7222.88
245 | 7222.54
246 | 7218.52
247 | 7231.72
248 | 7225.64
249 | 7228.59
250 | 7218.83
251 | 7219.17
252 | 7210.44
253 | 7221.33
254 | 7214.04
255 | 7221.6
256 | 7218.09
257 | 7238.95
258 | 7220.67
259 | 7228.06
260 | 7218.93
261 | 7222.19
262 | 7210.4
263 | 7217.5
264 | 7228.26
265 | 7215.41
266 | 7231.92
267 | 7218.91
268 | 7245.21
269 | 7216.16
270 | 7207.98
271 | 7217.23
272 | 7240.54
273 | 7223.53
274 | 7219.02
275 | 7220.95
276 | 7225.08
277 | 7204.17
278 | 7236.29
279 | 7234.52
280 | 7221.66
281 | 7241.24
282 | 7234.11
283 | 7226.7
284 | 7231.73
285 | 7223.34
286 | 7227.24
287 | 7219.75
288 | 7222.61
289 | 7209.56
290 | 7228.65
291 | 7217.06
292 | 7230.59
293 | 7214.43
294 | 7220.92
295 | 7233.97
296 | 7227.16
297 | 7207.57
298 | 7239.08
299 | 7226.56
300 | 7216.97
301 | 7225.8
302 | 7221.04
303 | 7239.33
304 | 7233.91
305 | 7220.36
306 | 7222.48
307 | 7232.93
308 | 7218.14
309 | 7211.05
310 | 7233.43
311 | 7207.33
312 | 7209.4
313 | 7225.14
314 | 7228.09
315 | 7233.88
316 | 7204.13
317 | 7195.34
318 | 7223.51
319 | 7213.36
320 | 7235.63
321 | 7204.03
322 | 7228.38
323 | 7228.87
324 | 7207.82
325 | 7231.88
326 | 7221.91
327 | 7210.46
328 | 7216.63
329 | 7229.13
330 | 7209.74
331 | 7207.74
332 | 7219.58
333 | 7208.63
334 | 7231.31
335 | 7206.94
336 | 7220.44
337 | 7220.77
338 | 7219
339 | 7220.96
340 | 7215.95
341 | 7226.31
342 | 7224.09
343 | 7238.23
344 | 7205.29
345 | 7213.43
346 | 7222.51
347 | 7217.34
348 | 7208.44
349 | 7214.22
350 | 7230.85
351 | 7211.15
352 | 7229.81
353 | 7223.18
354 | 7235.1
355 | 7200.67
356 | 7221.31
357 | 7207.71
358 | 7213.67
359 | 7216.76
360 | 7218.7
361 | 7219.54
362 | 7199.97
363 | 7219.26
364 | 7227.45
365 | 7194.74
366 | 7214.15
367 | 7210.69
368 | 7208.62
369 | 7233.8
370 | 7215.11
371 | 7226.56
372 | 7210.41
373 | 7222.27
374 | 7233.92
375 | 7219.53
376 | 7199.33
377 | 7200.28
378 | 7222.42
379 | 7209.79
380 | 7229.65
381 | 7216.93
382 | 7209.85
383 | 7217.18
384 | 7207.9
385 | 7207.55
386 | 7214.15
387 | 7226.92
388 | 7222.46
389 | 7214.59
390 | 7214.88
391 | 7218.47
392 | 7215.73
393 | 7225.09
394 | 7223.38
395 | 7214.42
396 | 7216.78
397 | 7223.58
398 | 7211.22
399 | 7223.19
400 | 7212.47
401 | 7219.95
402 | 7216.51
403 | 7227.64
404 | 7224.06
405 | 7209.24
406 | 7216.28
407 | 7236.65
408 | 7214.69
409 | 7221.36
410 | 7216.14
411 | 7229.51
412 | 7219.45
413 | 7226.14
414 | 7208.61
415 | 7229.62
416 | 7252.73
417 | 7206.94
418 | 7242.47
419 | 7214.49
420 | 7248.9
421 | 7237.75
422 | 7246.14
423 | 7219.14
424 | 7223.82
425 | 7232.36
426 | 7219.65
427 | 7216.66
428 | 7222.88
429 | 7238.17
430 | 7234.79
431 | 7220.69
432 | 7227.04
433 | 7231.2
434 | 7244.18
435 | 7228.34
436 | 7205.19
437 | 7206.54
438 | 7231.95
439 | 7214.74
440 | 7228.71
441 | 7219.55
442 | 7235.78
443 | 7241.6
444 | 7215.35
445 | 7219.16
446 | 7219.9
447 | 7226.06
448 | 7239.23
449 | 7217.17
450 | 7219.2
451 | 7217.06
452 | 7219.7
453 | 7229.16
454 | 7229.15
455 | 7240.05
456 | 7228.3
457 | 7205.57
458 | 7237.63
459 | 7231.43
460 | 7217.68
461 | 7211.77
462 | 7231.33
463 | 7205.04
464 | 7227.73
465 | 7224.3
466 | 7233.81
467 | 7215.85
468 | 7217.48
469 | 7229.11
470 | 7197.28
471 | 7197.77
472 | 7200.43
473 | 7207.61
474 | 7239.89
475 | 7220.06
476 | 7230.73
477 | 7206.24
478 | 7211.39
479 | 7221.07
480 | 7222.37
481 | 7205.52
482 | 7210.38
483 | 7231.29
484 | 7204.42
485 | 7213.02
486 | 7239.81
487 | 7213
488 | 7201.1
489 | 7207.89
490 | 7196.6
491 | 7203.55
492 | 7211.68
493 | 7211.4
494 | 7220.1
495 | 7224.42
496 | 7218.32
497 | 7209.8
498 | 7211.05
499 | 7208.95
500 | 7218.36
501 | 7223.32
502 | 7194.13
503 | 7210.82
504 | 7214.77
505 | 7218.04
506 | 7207.6
507 | 7226.47
508 | 7224.38
509 | 7200.54
510 | 7198.68
511 | 7222.68
512 | 7197.3
513 | 7206.46
514 | 7214.38
515 | 7196.93
516 | 7208.74
517 | 7215.98
518 | 7217.51
519 | 7218.96
520 | 7226.86
521 | 7216.72
522 | 7220.51
523 | 7203.85
524 | 7230.28
525 | 7215.99
526 | 7228.66
527 | 7218.71
528 | 7232.82
529 | 7206.1
530 | 7215.11
531 | 7224.42
532 | 7226.71
533 | 7220.21
534 | 7229.1
535 | 7219.62
536 | 7219.46
537 | 7200.17
538 | 7222.88
539 | 7223.02
540 | 7201.84
541 | 7206.76
542 | 7228.07
543 | 7215.1
544 | 7207.45
545 | 7221.28
546 | 7204.49
547 | 7209.66
548 | 7201.03
549 | 7214.15
550 | 7216.27
551 | 7217.43
552 | 7201.25
553 | 7229.51
554 | 7217.12
555 | 7199.9
556 | 7216.5
557 | 7228.3
558 | 7216.01
559 | 7211.77
560 | 7228.65
561 | 7225.82
562 | 7218.07
563 | 7207.78
564 | 7198.34
565 | 7222.79
566 | 7217.34
567 | 7242.32
568 | 7202.71
569 | 7219.83
570 | 7208.1
571 | 7211.14
572 | 7209.95
573 | 7205.39
574 | 7210.98
575 | 7216.26
576 | 7210.16
577 | 7194.72
578 | 7193.6
579 | 7189.38
580 | 7212.45
581 | 7201.64
582 | 7203.96
583 | 7206.92
584 | 7212.65
585 | 7218.06
586 | 7212.54
587 | 7225.09
588 | 7209.78
589 | 7199.99
590 | 7220.4
591 | 7234.65
592 | 7213.58
593 | 7228.97
594 | 7221.15
595 | 7203.18
596 | 7209.03
597 | 7210.96
598 | 7216.21
599 | 7225.15
600 | 7220.08
601 | 7212.78
602 | 7216.2
603 | 7216.18
604 | 7214.53
605 | 7205.43
606 | 7221.98
607 | 7226.53
608 | 7213.81
609 | 7225.05
610 | 7213.83
611 | 7216.21
612 | 7216.24
613 | 7226.07
614 | 7211.8
615 | 7195.21
616 | 7221.33
617 | 7232.65
618 | 7212.97
619 | 7223.48
620 | 7221.16
621 | 7209.63
622 | 7212.92
623 | 7204.15
624 | 7211.18
625 | 7219.4
626 | 7219.88
627 | 7226.73
628 | 7227.26
629 | 7212.84
630 | 7217.72
631 | 7225.29
632 | 7219.83
633 | 7216.36
634 | 7218.11
635 | 7219.01
636 | 7213.59
637 | 7202.29
638 | 7226.14
639 | 7196.65
640 | 7205.73
641 | 7219.48
642 | 7224.43
643 | 7229.05
644 | 7220.9
645 | 7202.14
646 | 7196.6
647 | 7198.68
648 | 7226.83
649 | 7225.76
650 | 7215.2
651 | 7190.59
652 | 7224.91
653 | 7201.31
654 | 7204.05
655 | 7204.46
656 | 7199.54
657 | 7196.94
658 | 7201.98
659 | 7208.02
660 | 7227.16
661 | 7208.29
662 | 7194.78
663 | 7198.94
664 | 7208.05
665 | 7213.51
666 | 7216.3
667 | 7213.86
668 | 7209.66
669 | 7196.29
670 | 7209.85
671 | 7210.21
672 | 7204.52
673 | 7204.24
674 | 7198.04
675 | 7194.44
676 | 7202.56
677 | 7195.47
678 | 7197.68
679 | 7200.03
680 | 7201.3
681 | 7206.53
682 | 7200.04
683 | 7211
684 | 7215.26
685 | 7217.09
686 | 7207.78
687 | 7190.99
688 | 7189.17
689 | 7202.31
690 | 7187.31
691 | 7204.83
692 | 7180.81
693 | 7188.54
694 | 7217.8
695 | 7198.79
696 | 7204.5
697 | 7198.15
698 | 7200.78
699 | 7186.23
700 | 7198.78
701 | 7196.67
702 | 7199
703 | 7199.23
704 | 7193.81
705 | 7208.95
706 | 7201.73
707 | 7197.45
708 | 7217.15
709 | 7207.31
710 | 7207.21
711 | 7206.08
712 | 7187.98
713 | 7196.27
714 | 7192.11
715 | 7200.65
716 | 7192.13
717 | 7187.83
718 | 7217.47
719 | 7202.7
720 | 7192.9
721 | 7209.19
722 | 7191.68
723 | 7205.06
724 | 7195.89
725 | 7195.32
726 | 7205.68
727 | 7193.77
728 | 7213.9
729 | 7197.76
730 | 7209.18
731 | 7195.23
732 | 7222.9
733 | 7178.43
734 | 7209.52
735 | 7201.96
736 | 7193.43
737 | 7201.19
738 | 7190.05
739 | 7180.92
740 | 7214.19
741 | 7203.45
742 | 7198.5
743 | 7214.41
744 | 7215.25
745 | 7211.9
746 | 7229.29
747 | 7205.05
748 | 7202.99
749 | 7192.49
750 | 7217.11
751 | 7222.96
752 | 7185.59
753 | 7210.72
754 | 7208.96
755 | 7200.17
756 | 7217.4
757 | 7204.65
758 | 7217.04
759 | 7218.8
760 | 7194.33
761 | 7205.76
762 | 7216.33
763 | 7209.64
764 | 7226.25
765 | 7201.59
766 | 7230.78
767 | 7216.05
768 | 7204.69
769 | 7201.82
770 | 7216.75
771 | 7205.31
772 | 7200.29
773 | 7212.12
774 | 7211.22
775 | 7228.13
776 | 7200.26
777 | 7230.27
778 | 7222.8
779 | 7206.11
780 | 7244.05
781 | 7213.53
782 | 7225.65
783 | 7235.32
784 | 7221.16
785 | 7209.41
786 | 7228.68
787 | 7209.31
788 | 7220.5
789 | 7217.97
790 | 7230.98
791 | 7228.83
792 | 7199.84
793 | 7207.99
794 | 7206.89
795 | 7217.19
796 | 7203.28
797 | 7202.37
798 | 7222.1
799 | 7209.86
800 | 7208.32
801 | 7210.36
802 | 7205.04
803 | 7201.68
804 | 7215.91
805 | 7194.68
806 | 7220.56
807 | 7212.32
808 | 7217.98
809 | 7209.88
810 | 7219.4
811 | 7214.91
812 | 7215.12
813 | 7221.53
814 | 7200.28
815 | 7209.86
816 | 7220.28
817 | 7200.06
818 | 7210.94
819 | 7206.06
820 | 7225.05
821 | 7199.87
822 | 7221.8
823 | 7216.14
824 | 7190.61
825 | 7209.8
826 | 7226.89
827 | 7214.11
828 | 7212.48
829 | 7222.77
830 | 7189.61
831 | 7208.53
832 | 7221.7
833 | 7197.89
834 | 7210.12
835 | 7229.07
836 | 7213.18
837 | 7210.38
838 | 7207.04
839 | 7223.36
840 | 7216.57
841 | 7199.07
842 | 7207.77
843 | 7217.48
844 | 7218.12
845 | 7211.96
846 | 7217.98
847 | 7222.59
848 | 7224.9
849 | 7198.74
850 | 7210.41
851 | 7220.35
852 | 7196
853 | 7214.83
854 | 7210.97
855 | 7217.72
856 | 7205.61
857 | 7203.94
858 | 7207.76
859 | 7200.07
860 | 7200.46
861 | 7200.07
862 | 7224.51
863 | 7224.66
864 | 7210.35
865 | 7209.61
866 | 7232.75
867 | 7220.17
868 | 7224.22
869 | 7201.39
870 | 7232.11
871 | 7223.22
872 | 7239.4
873 | 7219.29
874 | 7192.08
875 | 7200.75
876 | 7217.82
877 | 7231.92
878 | 7215.1
879 | 7205.03
880 | 7209.34
881 | 7225.37
882 | 7205.06
883 | 7201.8
884 | 7227.5
885 | 7203.18
886 | 7224.72
887 | 7216.26
888 | 7215.01
889 | 7205.33
890 | 7231.88
891 | 7219.68
892 | 7209.84
893 | 7210.03
894 | 7199.77
895 | 7217.85
896 | 7209.66
897 | 7199.66
898 | 7214.14
899 | 7232.66
900 | 7198.25
901 | 7210.17
902 | 7215.65
903 | 7206.79
904 | 7217.29
905 | 7197.07
906 | 7221.98
907 | 7208.01
908 | 7211.15
909 | 7220.62
910 | 7222.23
911 | 7216.6
912 | 7206.53
913 | 7214
914 | 7199.24
915 | 7199.86
916 | 7216.86
917 | 7222.31
918 | 7221.54
919 | 7197.48
920 | 7209.39
921 | 7214.33
922 | 7233.58
923 | 7213.03
924 | 7213.49
925 | 7215.66
926 | 7232.98
927 | 7212.9
928 | 7211.02
929 | 7222.29
930 | 7205.51
931 | 7231.91
932 | 7202.13
933 | 7197.99
934 | 7227.64
935 | 7196.1
936 | 7190.08
937 | 7217.44
938 | 7209.16
939 | 7217.67
940 | 7215.04
941 | 7227.35
942 | 7204.04
943 | 7214.92
944 | 7214.15
945 | 7214.77
946 | 7206.27
947 | 7217.58
948 | 7210.9
949 | 7223.41
950 | 7226.27
951 | 7219.3
952 | 7216.16
953 | 7196.72
954 | 7211.44
955 | 7202.25
956 | 7221.81
957 | 7240.97
958 | 7210.79
959 | 7207.31
960 | 7222
961 | 7212.99
962 | 7216.95
963 | 7212.38
964 | 7237.55
965 | 7213.68
966 | 7211.97
967 | 7209.4
968 | 7211.29
969 | 7215.26
970 | 7217.04
971 | 7213.3
972 | 7197.88
973 | 7219.24
974 | 7222.57
975 | 7209.59
976 | 7211.75
977 | 7206.91
978 | 7212.82
979 | 7203.93
980 | 7207.35
981 | 7197.45
982 | 7219.88
983 | 7197.07
984 | 7206.2
985 | 7212.61
986 | 7194.46
987 | 7211.1
988 | 7217.62
989 | 7196.23
990 | 7208.7
991 | 7210.15
992 | 7200.94
993 | 7223.18
994 | 7186.25
995 | 7214.28
996 | 7219.62
997 | 7216.66
998 | 7212.83
999 | 7215.68
1000 | 7217.44
1001 | 7214.22
1002 | 7207.81
1003 | 7220.87
1004 | 7207.06
1005 | 7231.98
1006 | 7228.28
1007 | 7215.25
1008 | 7205.17
1009 | 7208.99
1010 | 7218.61
1011 | 7202.8
1012 | 7211.54
1013 | 7223.93
1014 | 7201.35
1015 | 7207.66
1016 | 7210.33
1017 | 7217.06
1018 | 7213.64
1019 | 7221.79
1020 | 7208.35
1021 | 7216.7
1022 | 7225.84
1023 | 7216.72
1024 | 7231.24
1025 | 7214.48
1026 | 7204.51
1027 | 7223.63
1028 | 7214.55
1029 | 7225.44
1030 | 7206.37
1031 | 7229.17
1032 | 7224.02
1033 | 7179.7
1034 | 7201.15
1035 | 7226.11
1036 | 7203.53
1037 | 7221.79
1038 | 7220.47
1039 | 7207.56
1040 | 7212.33
1041 | 7209.57
1042 | 7210.28
1043 | 7218.09
1044 | 7205.74
1045 | 7200.88
1046 | 7214.58
1047 | 7219.04
1048 | 7223.54
1049 | 7208.38
1050 | 7194.54
1051 | 7223.23
1052 | 7211.88
1053 | 7205.6
1054 | 7195.02
1055 | 7220.1
1056 | 7210.91
1057 | 7225.98
1058 | 7214.41
1059 | 7220.76
1060 | 7211.58
1061 | 7228.16
1062 | 7222.23
1063 | 7219.34
1064 | 7222.61
1065 | 7210.42
1066 | 7204.13
1067 | 7217.74
1068 | 7197.93
1069 | 7228.89
1070 | 7205.33
1071 | 7231.18
1072 | 7212.79
1073 | 7226.78
1074 | 7216.32
1075 | 7208.63
1076 | 7211.85
1077 | 7230.32
1078 | 7210.18
1079 | 7218.03
1080 | 7218.27
1081 | 7210.82
1082 | 7202.34
1083 | 7226.92
1084 | 7206.08
1085 | 7207.95
1086 | 7216.87
1087 | 7213.89
1088 | 7217.66
1089 | 7239.31
1090 | 7224.16
1091 | 7215.44
1092 | 7210.27
1093 | 7227.1
1094 | 7214.07
1095 | 7223.5
1096 | 7217.43
1097 | 7205.61
1098 | 7206.3
1099 | 7218.66
1100 | 7210.33
1101 | 7204.93
1102 | 7206.98
1103 | 7217
1104 | 7221.01
1105 | 7219.47
1106 | 7224.56
1107 | 7206.03
1108 | 7216.68
1109 | 7224.89
1110 | 7209.46
1111 | 7211.64
1112 | 7221
1113 | 7207.76
1114 | 7210.32
1115 | 7212.99
1116 | 7212.08
1117 | 7211.25
1118 | 7216.32
1119 | 7230
1120 | 7208.92
1121 | 7215.23
1122 | 7201.15
1123 | 7226.19
1124 | 7200.51
1125 | 7220.03
1126 | 7222.05
1127 | 7223.16
1128 | 7231.57
1129 | 7217.06
1130 | 7233.36
1131 | 7205.21
1132 | 7204.34
1133 | 7230.81
1134 | 7202.42
1135 | 7206.81
1136 | 7227.88
1137 | 7203.9
1138 | 7219.19
1139 | 7190.74
1140 | 7213.05
1141 | 7229.62
1142 | 7185.95
1143 | 7218.82
1144 | 7208
1145 | 7215.07
1146 | 7215.89
1147 | 7225.58
1148 | 7205.12
1149 | 7212.93
1150 | 7227.8
1151 | 7216.06
1152 | 7209.69
1153 | 7211.56
1154 | 7206.84
1155 | 7208.87
1156 | 7212.62
1157 | 7206.77
1158 | 7220.47
1159 | 7230.26
1160 | 7232.91
1161 | 7203.21
1162 | 7192.81
1163 | 7221.72
1164 | 7193.78
1165 | 7216.48
1166 | 7216.81
1167 | 7203.95
1168 | 7200.41
1169 | 7220.14
1170 | 7208.95
1171 | 7222.75
1172 | 7205.36
1173 | 7213.92
1174 | 7220.28
1175 | 7221.11
1176 | 7219.22
1177 | 7195.02
1178 | 7204.21
1179 | 7206.03
1180 | 7199.9
1181 | 7206.08
1182 | 7222.28
1183 | 7221.3
1184 | 7219.79
1185 | 7215.05
1186 | 7200.92
1187 | 7231.26
1188 | 7203.48
1189 | 7213.84
1190 | 7207.01
1191 | 7213.42
1192 | 7212.93
1193 | 7218.22
1194 | 7209.43
1195 | 7201.5
1196 | 7219.61
1197 | 7204.69
1198 | 7218.79
1199 | 7210.88
1200 | 7207.82
1201 | 7205.76
1202 | 7207.92
1203 | 7208.32
1204 | 7210.77
1205 | 7208.5
1206 | 7226.22
1207 | 7201.25
1208 | 7218.03
1209 | 7200.41
1210 | 7226.71
1211 | 7206.67
1212 | 7212.11
1213 | 7226.84
1214 | 7188.79
1215 | 7215.91
1216 | 7192.97
1217 | 7200.36
1218 | 7201.72
1219 | 7207.89
1220 | 7204.14
1221 | 7196.67
1222 | 7214.85
1223 | 7219.89
1224 | 7195.83
1225 | 7207.53
1226 | 7218.05
1227 | 7207.52
1228 | 7200.54
1229 | 7194.39
1230 | 7204.47
1231 | 7200.61
1232 | 7200.33
1233 | 7223.23
1234 | 7209.08
1235 | 7209.34
1236 | 7221.66
1237 | 7206.69
1238 | 7209.68
1239 | 7203.86
1240 | 7211.42
1241 | 7209.26
1242 | 7187.11
1243 | 7214.77
1244 | 7209
1245 | 7203.49
1246 | 7211.56
1247 | 7220.93
1248 | 7191.77
1249 | 7217.48
1250 | 7220.6
1251 | 7205.61
1252 | 7198.18
1253 | 7203.13
1254 | 7212.81
1255 | 7210.74
1256 | 7225.82
1257 | 7231.45
1258 | 7206.92
1259 | 7221.7
1260 | 7209.94
1261 | 7193.83
1262 | 7215.82
1263 | 7202.65
1264 | 7211.78
1265 | 7221.48
1266 | 7222.92
1267 | 7231.83
1268 | 7221.26
1269 | 7209.76
1270 | 7209.63
1271 | 7215.56
1272 | 7209.92
1273 | 7200.87
1274 | 7214.81
1275 | 7214.75
1276 | 7226.66
1277 | 7209.57
1278 | 7205.58
1279 | 7216.08
1280 | 7201.59
1281 | 7198.83
1282 | 7201.36
1283 | 7202.1
1284 | 7209.02
1285 | 7209.58
1286 | 7205.98
1287 | 7229.58
1288 | 7219.01
1289 | 7201.44
1290 | 7182.83
1291 | 7207.54
1292 | 7196.32
1293 | 7203.96
1294 | 7208.98
1295 | 7232.9
1296 | 7194.29
1297 | 7209.17
1298 | 7209.82
1299 | 7203.11
1300 | 7194.04
1301 | 7191.32
1302 | 7228.95
1303 | 7228.16
1304 | 7198.09
1305 | 7217.47
1306 | 7209.47
1307 | 7226.66
1308 | 7203.39
1309 | 7205.53
1310 | 7205.49
1311 | 7209.61
1312 | 7224.02
1313 | 7202.47
1314 | 7196.08
1315 | 7205.16
1316 | 7211.02
1317 | 7211.88
1318 | 7210.38
1319 | 7214.74
1320 | 7210.9
1321 | 7218.74
1322 | 7212.89
1323 | 7190.95
1324 | 7210.16
1325 | 7215.57
1326 | 7208.03
1327 | 7205.59
1328 | 7187.65
1329 | 7211.87
1330 | 7218.58
1331 | 7204.86
1332 | 7204.2
1333 | 7217.2
1334 | 7207.41
1335 | 7207.36
1336 | 7204.79
1337 | 7194.55
1338 | 7207.25
1339 | 7208.56
1340 | 7213.21
1341 | 7218.88
1342 | 7206.84
1343 | 7201.65
1344 | 7194.41
1345 | 7193.53
1346 | 7213.48
1347 | 7200.61
1348 | 7215.57
1349 | 7218.01
1350 | 7193.77
1351 | 7204.26
1352 | 7207.44
1353 | 7205.94
1354 | 7201.36
1355 | 7212.52
1356 | 7208.19
1357 | 7221.24
1358 | 7205.05
1359 | 7213.73
1360 | 7201.42
1361 | 7220.88
1362 | 7189.94
1363 | 7191.83
1364 | 7204.44
1365 | 7213.48
1366 | 7222.87
1367 | 7213.35
1368 | 7201.23
1369 | 7212.75
1370 | 7203.03
1371 | 7194.98
1372 | 7215.23
1373 | 7209.87
1374 | 7215.06
1375 | 7205.77
1376 | 7198.09
1377 | 7208.18
1378 | 7208.77
1379 | 7203.83
1380 | 7214.46
1381 | 7207.38
1382 | 7204.15
1383 | 7206.19
1384 | 7211.93
1385 | 7210.54
1386 | 7210.55
1387 | 7200.31
1388 | 7199.85
1389 | 7203.12
1390 | 7196.64
1391 | 7204.97
1392 | 7210.11
1393 | 7210.74
1394 | 7223.12
1395 | 7217.62
1396 | 7192.05
1397 | 7210.58
1398 | 7199.76
1399 | 7219.39
1400 | 7185.4
1401 | 7190.63
1402 | 7199.21
1403 | 7195.73
1404 | 7210.57
1405 | 7214.23
1406 | 7203.03
1407 | 7200
1408 | 7199.4
1409 | 7196.07
1410 | 7194.75
1411 | 7194.21
1412 | 7210.44
1413 | 7223.33
1414 | 7186.12
1415 | 7214.89
1416 | 7201.73
1417 | 7209.06
1418 | 7213.7
1419 | 7201.97
1420 | 7209.68
1421 | 7220.6
1422 | 7194.06
1423 | 7212.07
1424 | 7208.2
1425 | 7218.39
1426 | 7211.61
1427 | 7195.69
1428 | 7201.97
1429 | 7204.39
1430 | 7201.28
1431 | 7211.09
1432 | 7204.77
1433 | 7206.28
1434 | 7215.11
1435 | 7215.19
1436 | 7212.27
1437 | 7217.02
1438 | 7195.12
1439 | 7203.57
1440 | 7195.47
1441 | 7207.6
1442 | 7195.41
1443 | 7202.76
1444 | 7218.91
1445 | 7193.1
1446 | 7208.52
1447 | 7196.16
1448 | 7199.1
1449 | 7209.89
1450 | 7200.04
1451 | 7201.53
1452 | 7203.15
1453 | 7209.77
1454 | 7213.9
1455 | 7227.99
1456 | 7208.04
1457 | 7192.51
1458 | 7197.03
1459 | 7204.92
1460 | 7196.51
1461 | 7204.79
1462 | 7202.75
1463 | 7212.38
1464 | 7200.48
1465 | 7209.26
1466 | 7199.47
1467 | 7194.77
1468 | 7202.68
1469 | 7201.27
1470 | 7198.33
1471 | 7209.24
1472 | 7205.58
1473 | 7202.15
1474 | 7205.44
1475 | 7193.25
1476 | 7189.31
1477 | 7213.33
1478 | 7195.52
1479 | 7213.38
1480 | 7207.93
1481 | 7200.6
1482 | 7200.91
1483 | 7215.25
1484 | 7196.43
1485 | 7188.58
1486 | 7207.88
1487 | 7199.92
1488 | 7209.06
1489 | 7217.8
1490 | 7218.94
1491 | 7202.57
1492 | 7214.86
1493 | 7207.82
1494 | 7199.23
1495 | 7202.66
1496 | 7208.45
1497 | 7199.14
1498 | 7206.88
1499 | 7215.45
1500 | 7204.96
1501 | 7195.36
1502 | 7212.72
1503 | 7232.04
1504 | 7195.36
1505 | 7193.7
1506 | 7188.11
1507 | 7200.17
1508 | 7224.86
1509 | 7205.67
1510 | 7200.19
1511 | 7214.28
1512 | 7202.54
1513 | 7200.46
1514 | 7203.95
1515 | 7198.38
1516 | 7201.02
1517 | 7199.62
1518 | 7201.37
1519 | 7184.84
1520 | 7220.79
1521 | 7212.4
1522 | 7212.01
1523 | 7202.53
1524 | 7208.1
1525 | 7212.06
1526 | 7217.89
1527 | 7195.89
1528 | 7201.85
1529 | 7223.77
1530 | 7198.83
1531 | 7198.96
1532 | 7195.69
1533 | 7215.35
1534 | 7201.63
1535 | 7200.67
1536 | 7200.83
1537 | 7208.26
1538 | 7215.91
1539 | 7192.11
1540 | 7203.16
1541 | 7208.72
1542 | 7206.49
1543 | 7192.02
1544 | 7214.06
1545 | 7197.19
1546 | 7209.11
1547 | 7198.2
1548 | 7195.32
1549 | 7211.37
1550 | 7203.81
1551 | 7219.42
1552 | 7212.76
1553 | 7200.87
1554 | 7201.2
1555 | 7210.22
1556 | 7207.65
1557 | 7188.52
1558 | 7207.26
1559 | 7190.16
1560 | 7182.12
1561 | 7212.98
1562 | 7184.95
1563 | 7217
1564 | 7208.02
1565 | 7219.21
1566 | 7202.57
1567 | 7207.93
1568 | 7198.73
1569 | 7194.78
1570 | 7196.84
1571 | 7198.9
1572 | 7211.62
1573 | 7217.4
1574 | 7202.13
1575 | 7170.52
1576 | 7186.48
1577 | 7186.74
1578 | 7195.35
1579 | 7200.01
1580 | 7207.26
1581 | 7185.13
1582 | 7210.15
1583 | 7196.46
1584 | 7224.96
1585 | 7208.58
1586 | 7205.67
1587 | 7188.72
1588 | 7203.61
1589 | 7213.79
1590 | 7188.34
1591 | 7191.45
1592 | 7197.91
1593 | 7196.64
1594 | 7193.81
1595 | 7182.75
1596 | 7217.07
1597 | 7213.88
1598 | 7194.98
1599 | 7192.56
1600 | 7192.62
1601 | 7209.75
1602 | 7196.1
1603 | 7195.28
1604 | 7208.78
1605 | 7211.45
1606 | 7198.72
1607 | 7196.35
1608 | 7222.1
1609 | 7199.58
1610 | 7202.63
1611 | 7210.11
1612 | 7191.49
1613 | 7213.38
1614 | 7201.43
1615 | 7213.13
1616 | 7200.19
1617 | 7210.37
1618 | 7189
1619 | 7192.32
1620 | 7221
1621 | 7203.72
1622 | 7198.42
1623 | 7205.39
1624 | 7195.06
1625 | 7210.01
1626 | 7210.61
1627 | 7202.14
1628 | 7185.95
1629 | 7204.22
1630 | 7200.13
1631 | 7201.61
1632 | 7202.11
1633 | 7190.76
1634 | 7189.18
1635 | 7215.63
1636 | 7190.75
1637 | 7175.01
1638 | 7193.72
1639 | 7200.99
1640 | 7205.84
1641 | 7206.85
1642 | 7202.23
1643 | 7196.99
1644 | 7194.38
1645 | 7195.62
1646 | 7193.22
1647 | 7210
1648 | 7218.04
1649 | 7190.17
1650 | 7211.22
1651 | 7209.34
1652 | 7192.38
1653 | 7194.48
1654 | 7190.35
1655 | 7222.55
1656 | 7206.52
1657 | 7199.09
1658 | 7205.55
1659 | 7175.47
1660 | 7215.29
1661 | 7195.99
1662 | 7204.26
1663 | 7203.67
1664 | 7217.31
1665 | 7202.15
1666 | 7190.7
1667 | 7189.27
1668 | 7211.44
1669 | 7188.51
1670 | 7197.27
1671 | 7191.88
1672 | 7208.18
1673 | 7195.81
1674 | 7214.06
1675 | 7206.4
1676 | 7200.11
1677 | 7212.39
1678 | 7199.02
1679 | 7207.97
1680 | 7215.55
1681 | 7205.35
1682 | 7207.96
1683 | 7210.33
1684 | 7191.21
1685 | 7196.7
1686 | 7211.31
1687 | 7200.64
1688 | 7204.42
1689 | 7200.17
1690 | 7202.25
1691 | 7201.02
1692 | 7191.63
1693 | 7197.64
1694 | 7193.67
1695 | 7190.88
1696 | 7209.17
1697 | 7189.39
1698 | 7189.88
1699 | 7206.23
1700 | 7208.98
1701 | 7212.79
1702 | 7202.67
1703 | 7200.57
1704 | 7204.58
1705 | 7201.37
1706 | 7197.54
1707 | 7186.84
1708 | 7181.78
1709 | 7175.73
1710 | 7204.03
1711 | 7193.71
1712 | 7194.02
1713 | 7187.69
1714 | 7190.51
1715 | 7190.71
1716 | 7219.32
1717 | 7183.07
1718 | 7213.86
1719 | 7189.95
1720 | 7184.1
1721 | 7204.5
1722 | 7196.92
1723 | 7213.02
1724 | 7199.67
1725 | 7197.54
1726 | 7213.55
1727 | 7197.92
1728 | 7179.11
1729 | 7201.82
1730 | 7191.8
1731 | 7185.72
1732 | 7204.59
1733 | 7182.74
1734 | 7183.2
1735 | 7195.26
1736 | 7201.43
1737 | 7196.41
1738 | 7205.78
1739 | 7195.78
1740 | 7185.24
1741 | 7187.84
1742 | 7188.23
1743 | 7206.81
1744 | 7191.12
1745 | 7196.06
1746 | 7180.5
1747 | 7194.35
1748 | 7191.24
1749 | 7207.89
1750 | 7205.44
1751 | 7187.52
1752 | 7177.46
1753 | 7188.52
1754 | 7173.25
1755 | 7190.5
1756 | 7171.83
1757 | 7175.96
1758 | 7196.15
1759 | 7188.75
1760 | 7197.49
1761 | 7191.82
1762 | 7187.13
1763 | 7215.91
1764 | 7191.31
1765 | 7186.08
1766 | 7204.4
1767 | 7198.01
1768 | 7200.66
1769 | 7202.21
1770 | 7190.45
1771 | 7195.97
1772 | 7193.44
1773 | 7190.07
1774 | 7186.39
1775 | 7180.57
1776 | 7175.59
1777 | 7191.84
1778 | 7176.92
1779 | 7199.39
1780 | 7196.64
1781 | 7198.18
1782 | 7194.26
1783 | 7185.16
1784 | 7200.92
1785 | 7194
1786 | 7185.4
1787 | 7181.89
1788 | 7185.43
1789 | 7177.22
1790 | 7187.26
1791 | 7184.36
1792 | 7197.84
1793 | 7191.6
1794 | 7188.82
1795 | 7199.29
1796 | 7209.96
1797 | 7195.41
1798 | 7185.94
1799 | 7195.07
1800 | 7199.88
1801 | 7202.8
1802 | 7199.43
1803 | 7191.62
1804 | 7186.12
1805 | 7164.75
1806 | 7189.11
1807 | 7192.14
1808 | 7203.23
1809 | 7185.32
1810 | 7182.11
1811 | 7206.38
1812 | 7205.49
1813 | 7191.86
1814 | 7200.64
1815 | 7200.64
1816 | 7207.73
1817 | 7205.44
1818 | 7200.23
1819 | 7181.37
1820 | 7190.2
1821 | 7212.07
1822 | 7194.67
1823 | 7203.71
1824 | 7181.95
1825 | 7217.07
1826 | 7196.27
1827 | 7208.33
1828 | 7193.44
1829 | 7206.94
1830 | 7196.63
1831 | 7198.98
1832 | 7181.75
1833 | 7189.54
1834 | 7187.79
1835 | 7201.84
1836 | 7194.99
1837 | 7207.07
1838 | 7186.32
1839 | 7196.36
1840 | 7204.32
1841 | 7189.61
1842 | 7208.27
1843 | 7193.91
1844 | 7202.77
1845 | 7188.19
1846 | 7212.88
1847 | 7191.36
1848 | 7206.22
1849 | 7201.35
1850 | 7209.79
1851 | 7216.75
1852 | 7219.67
1853 | 7200.73
1854 | 7207.51
1855 | 7205.89
1856 | 7204.78
1857 | 7197.84
1858 | 7192.44
1859 | 7201.36
1860 | 7213.18
1861 | 7223.08
1862 | 7213.89
1863 | 7202.45
1864 | 7208.99
1865 | 7211.65
1866 | 7201.11
1867 | 7212.92
1868 | 7202.68
1869 | 7214.86
1870 | 7190.88
1871 | 7184.3
1872 | 7214.33
1873 | 7206.15
1874 | 7211.18
1875 | 7205.23
1876 | 7195.67
1877 | 7195.24
1878 | 7211.54
1879 | 7188.82
1880 | 7212.85
1881 | 7211.09
1882 | 7211.2
1883 | 7207.45
1884 | 7200.93
1885 | 7200.39
1886 | 7204.04
1887 | 7199.81
1888 | 7211.76
1889 | 7207.94
1890 | 7202.49
1891 | 7195
1892 | 7217.13
1893 | 7191.45
1894 | 7229.43
1895 | 7207.55
1896 | 7212.11
1897 | 7216.22
1898 | 7209.11
1899 | 7203.99
1900 | 7231.27
1901 | 7211.3
1902 | 7221.92
1903 | 7196.5
1904 | 7219.48
1905 | 7201.59
1906 | 7201.62
1907 | 7219.42
1908 | 7219.96
1909 | 7204.02
1910 | 7204.55
1911 | 7200.12
1912 | 7218
1913 | 7206.06
1914 | 7226.95
1915 | 7211.82
1916 | 7211.3
1917 | 7190.05
1918 | 7206.55
1919 | 7215.92
1920 | 7195.41
1921 | 7215.61
1922 | 7205.33
1923 | 7211.77
1924 | 7212.44
1925 | 7203.89
1926 | 7216.19
1927 | 7199.46
1928 | 7204.71
1929 | 7201.63
1930 | 7203.98
1931 | 7202.86
1932 | 7190.67
1933 | 7198.02
1934 | 7203.42
1935 | 7198.62
1936 | 7211.51
1937 | 7219.38
1938 | 7180.77
1939 | 7196.07
1940 | 7211.29
1941 | 7220.36
1942 | 7207.13
1943 | 7205.9
1944 | 7204.56
1945 | 7193.86
1946 | 7214.71
1947 | 7187.12
1948 | 7214
1949 | 7178.13
1950 | 7204.98
1951 | 7204.32
1952 | 7213.03
1953 | 7207.89
1954 | 7206.13
1955 | 7211.76
1956 | 7186.29
1957 | 7214.38
1958 | 7204.13
1959 | 7215.17
1960 | 7215.92
1961 | 7209.18
1962 | 7191.06
1963 | 7197.78
1964 | 7199.58
1965 | 7206.55
1966 | 7197.41
1967 | 7207.01
1968 | 7197.34
1969 | 7205.34
1970 | 7194.82
1971 | 7199.06
1972 | 7195.96
1973 | 7202.31
1974 | 7186.4
1975 | 7203.01
1976 | 7191.64
1977 | 7208.99
1978 | 7206.87
1979 | 7186.63
1980 | 7200.23
1981 | 7213.68
1982 | 7204.75
1983 | 7209.75
1984 | 7189.21
1985 | 7213.34
1986 | 7213.48
1987 | 7201.11
1988 | 7196.85
1989 | 7217.91
1990 | 7190.45
1991 | 7220.44
1992 | 7222.12
1993 | 7210.51
1994 | 7214.57
1995 | 7215.54
1996 | 7220.34
1997 | 7190.64
1998 | 7237.92
1999 | 7202.4
2000 | 7189.85
2001 | 7192.24
2002 | 7198.45
2003 | 7211.28
2004 | 7200.13
2005 | 7208.85
2006 | 7221.53
2007 | 7200.6
2008 | 7187.5
2009 | 7185.38
2010 | 7190.04
2011 | 7212.55
2012 | 7200.48
2013 | 7196.1
2014 | 7195.44
2015 | 7201.72
2016 | 7204.64
2017 | 7204.89
2018 | 7213.42
2019 | 7194.99
2020 | 7206.75
2021 | 7191.62
2022 | 7212.39
2023 | 7194.65
2024 | 7215.11
2025 | 7201.21
2026 | 7208.4
2027 | 7210.71
2028 | 7200.32
2029 | 7193.81
2030 | 7186.83
2031 | 7184.77
2032 | 7197.68
2033 | 7194.17
2034 | 7200.45
2035 | 7207.1
2036 | 7192.52
2037 | 7194.53
2038 | 7205.43
2039 | 7225.91
2040 | 7191.64
2041 | 7204.75
2042 | 7192.41
2043 | 7202.64
2044 | 7211.1
2045 | 7212.85
2046 | 7217.66
2047 | 7206.85
2048 | 7201.88
2049 | 7219.84
2050 | 7194.1
2051 | 7193.1
2052 | 7209.05
2053 | 7197.49
2054 | 7207.15
2055 | 7190.49
2056 | 7204.72
2057 | 7207.86
2058 | 7183.65
2059 | 7186.79
2060 | 7197.77
2061 | 7197.23
2062 | 7193.06
2063 | 7207.57
2064 | 7179.2
2065 | 7197.08
2066 | 7176.17
2067 | 7200.34
2068 | 7194.88
2069 | 7207.65
2070 | 7184.02
2071 | 7198.82
2072 | 7190.66
2073 | 7192.95
2074 | 7188.1
2075 | 7197.03
2076 | 7208.03
2077 | 7204.28
2078 | 7197.28
2079 | 7181.98
2080 | 7197.44
2081 | 7228.89
2082 | 7197.09
2083 | 7198.72
2084 | 7204.43
2085 | 7190.88
2086 | 7191.82
2087 | 7187.25
2088 | 7197.66
2089 | 7192.75
2090 | 7201.64
2091 | 7188.54
2092 | 7186.06
2093 | 7192.99
2094 | 7199.7
2095 | 7199.08
2096 | 7195.26
2097 | 7190.34
2098 | 7209.13
2099 | 7198.19
2100 | 7221.91
2101 | 7201.86
2102 | 7179.29
2103 | 7198.74
2104 | 7208.46
2105 | 7210.99
2106 | 7207.14
2107 | 7196.67
2108 | 7191.99
2109 | 7198.14
2110 | 7187.7
2111 | 7174.67
2112 | 7207.43
2113 | 7189.3
2114 | 7204.87
2115 | 7186.52
2116 | 7185.23
2117 | 7198.54
2118 | 7187.95
2119 | 7196.44
2120 | 7200.33
2121 | 7208.8
2122 | 7200.64
2123 | 7207.37
2124 | 7180.42
2125 | 7219.16
2126 | 7194.68
2127 | 7194.05
2128 | 7193.66
2129 | 7201.23
2130 | 7193.59
2131 | 7202.77
2132 | 7190.01
2133 | 7203.28
2134 | 7197.4
2135 | 7185.86
2136 | 7181.75
2137 | 7186.89
2138 | 7204.76
2139 | 7208.26
2140 | 7211.5
2141 | 7202.46
2142 | 7189.94
2143 | 7210.99
2144 | 7211.12
2145 | 7185.28
2146 | 7189.42
2147 | 7182.26
2148 | 7191.13
2149 | 7192.39
2150 | 7209.71
2151 | 7192.17
2152 | 7180.36
2153 | 7218.44
2154 | 7204.82
2155 | 7193.57
2156 | 7191.84
2157 | 7207.65
2158 | 7191.94
2159 | 7200.49
2160 | 7198.5
2161 | 7193.97
2162 | 7201.63
2163 | 7188.77
2164 | 7200.4
2165 | 7191.45
2166 | 7203.3
2167 | 7185.75
2168 | 7184.91
2169 | 7195.31
2170 | 7202.5
2171 | 7197.55
2172 | 7184.54
2173 | 7213.82
2174 | 7198.77
2175 | 7191.04
2176 | 7210.75
2177 | 7202.85
2178 | 7186.4
2179 | 7209.53
2180 | 7216.6
2181 | 7184.32
2182 | 7194.68
2183 | 7202.35
2184 | 7203.83
2185 | 7203.1
2186 | 7219.47
2187 | 7198.3
2188 | 7200.03
2189 | 7193.36
2190 | 7188.5
2191 | 7184.63
2192 | 7193.44
2193 | 7191.38
2194 | 7190.14
2195 | 7194
2196 | 7206.01
2197 | 7192.35
2198 | 7208.7
2199 | 7201.38
2200 | 7184.2
2201 | 7209.72
2202 | 7190.91
2203 | 7196.98
2204 | 7195.41
2205 | 7191.2
2206 | 7190.92
2207 | 7183.61
2208 | 7183.2
2209 | 7182.16
2210 | 7203.63
2211 | 7180.18
2212 | 7199.62
2213 | 7208.31
2214 | 7180.88
2215 | 7170.91
2216 | 7185.96
2217 | 7186.47
2218 | 7195.07
2219 | 7186.13
2220 | 7215.59
2221 | 7193.34
2222 | 7187.74
2223 | 7206.94
2224 | 7197.69
2225 | 7185.13
2226 | 7198.02
2227 | 7198.01
2228 | 7181.39
2229 | 7201.94
2230 | 7176.54
2231 | 7198.97
2232 | 7189.57
2233 | 7175.88
2234 | 7194.06
2235 | 7198.57
2236 | 7201.34
2237 | 7197.74
2238 | 7202.02
2239 | 7197.81
2240 | 7182.8
2241 | 7207.32
2242 | 7210.02
2243 | 7192
2244 | 7195.83
2245 | 7167.95
2246 | 7189.08
2247 | 7211.34
2248 | 7183.05
2249 | 7189.27
2250 | 7204.21
2251 | 7182.56
2252 | 7210.72
2253 | 7199.03
2254 | 7185.85
2255 | 7198.51
2256 | 7184
2257 | 7208.25
2258 | 7200.35
2259 | 7199.15
2260 | 7168.16
2261 | 7218.6
2262 | 7206.17
2263 | 7178.31
2264 | 7187.07
2265 | 7191.74
2266 | 7192.46
2267 | 7213.2
2268 | 7199.67
2269 | 7204.33
2270 | 7182.64
2271 | 7206.66
2272 | 7199.04
2273 | 7205.29
2274 | 7195.32
2275 | 7214.22
2276 | 7203.13
2277 | 7215.03
2278 | 7187.64
2279 | 7190.44
2280 | 7189.47
2281 | 7188.59
2282 | 7178.56
2283 | 7183.19
2284 | 7189.07
2285 | 7182.7
2286 | 7175.74
2287 | 7194.05
2288 | 7210.53
2289 | 7198.09
2290 | 7222.11
2291 | 7194
2292 | 7200.24
2293 | 7214.72
2294 | 7201.38
2295 | 7208.25
2296 | 7205.77
2297 | 7182.1
2298 | 7195.46
2299 | 7200.96
2300 | 7176.94
2301 | 7187.09
2302 | 7218.56
2303 | 7205.07
2304 | 7213.86
2305 | 7186.64
2306 | 7194.62
2307 | 7193.88
2308 | 7216.96
2309 | 7180.18
2310 | 7189.83
2311 | 7195.95
2312 | 7185.64
2313 | 7206.6
2314 | 7198.71
2315 | 7199.89
2316 | 7202.94
2317 | 7198.83
2318 | 7197.98
2319 | 7199.52
2320 | 7180.49
2321 | 7193.38
2322 | 7207.82
2323 | 7195.38
2324 | 7207.75
2325 | 7199.49
2326 | 7195.75
2327 | 7188.4
2328 | 7182.94
2329 | 7197.62
2330 | 7195.57
2331 | 7179.52
2332 | 7189.76
2333 | 7186.31
2334 | 7201.99
2335 | 7185.19
2336 | 7196.31
2337 | 7201.13
2338 | 7172.48
2339 | 7193.33
2340 | 7187.87
2341 | 7186.89
2342 | 7190.32
2343 | 7198.66
2344 | 7197.74
2345 | 7182.87
2346 | 7208.62
2347 | 7197.14
2348 | 7177.63
2349 | 7185.27
2350 | 7174.61
2351 | 7191.05
2352 | 7195.85
2353 | 7195.7
2354 | 7201.83
2355 | 7190.61
2356 | 7207.17
2357 | 7201.39
2358 | 7186.02
2359 | 7193.86
2360 | 7188.6
2361 | 7196.87
2362 | 7187.21
2363 | 7202.18
2364 | 7193.72
2365 | 7191.51
2366 | 7202.91
2367 | 7191.71
2368 | 7184.96
2369 | 7192.06
2370 | 7180.24
2371 | 7203.47
2372 | 7190.69
2373 | 7206.12
2374 | 7200.04
2375 | 7192.99
2376 | 7184.65
2377 | 7205.3
2378 | 7195.04
2379 | 7179.96
2380 | 7193.26
2381 | 7202.95
2382 | 7176.48
2383 | 7196.55
2384 | 7211.68
2385 | 7183.4
2386 | 7183.36
2387 | 7186.28
2388 | 7205.83
2389 | 7208.02
2390 | 7213.75
2391 | 7196.05
2392 | 7204.14
2393 | 7193.21
2394 | 7201.29
2395 | 7180.8
2396 | 7192.7
2397 | 7215.36
2398 | 7191.75
2399 | 7202.09
2400 | 7175.85
2401 | 7188.32
2402 | 7168.42
2403 | 7202
2404 | 7200.07
2405 | 7192.32
2406 | 7184.9
2407 | 7191.98
2408 | 7190.74
2409 | 7197.81
2410 | 7195.72
2411 | 7185.11
2412 | 7202.95
2413 | 7183.74
2414 | 7191.47
2415 | 7186.11
2416 | 7198.29
2417 | 7195.36
2418 | 7201.48
2419 | 7182.34
2420 | 7212.13
2421 | 7196.48
2422 | 7193.97
2423 | 7184.67
2424 | 7192.74
2425 | 7201.24
2426 | 7189.93
2427 | 7194.24
2428 | 7189.13
2429 | 7178.57
2430 | 7192.1
2431 | 7203.11
2432 | 7198.31
2433 | 7209.08
2434 | 7184.51
2435 | 7199.63
2436 | 7182.2
2437 | 7192.4
2438 | 7197.35
2439 | 7197.64
2440 | 7195.1
2441 | 7189.64
2442 | 7190.32
2443 | 7181.94
2444 | 7194.66
2445 | 7201.77
2446 | 7191.39
2447 | 7203.24
2448 | 7199.43
2449 | 7189.42
2450 | 7187.53
2451 | 7175.34
2452 | 7182.82
2453 | 7205.19
2454 | 7201.06
2455 | 7197.87
2456 | 7218.99
2457 | 7198.55
2458 | 7201.6
2459 | 7200.61
2460 | 7187.4
2461 | 7199.85
2462 | 7218.49
2463 | 7195.57
2464 | 7206.13
2465 | 7198.97
2466 | 7207.92
2467 | 7202.62
2468 | 7182.82
2469 | 7191.36
2470 | 7194.37
2471 | 7180.67
2472 | 7202.3
2473 | 7183.81
2474 | 7199.13
2475 | 7174.87
2476 | 7203.78
2477 | 7209.99
2478 | 7181.28
2479 | 7197.72
2480 | 7202.26
2481 | 7193.86
2482 | 7196.76
2483 | 7188.61
2484 | 7187.44
2485 | 7190.46
2486 | 7195.57
2487 | 7194.81
2488 | 7201.91
2489 | 7171.36
2490 | 7186.27
2491 | 7205.55
2492 | 7186.01
2493 | 7210.29
2494 | 7195.48
2495 | 7197.2
2496 | 7196.02
2497 | 7182.47
2498 | 7196.15
2499 | 7196.66
2500 | 7196.63
2501 | 7189.06
2502 | 7182.7
2503 | 7198.23
2504 | 7190.88
2505 | 7185.95
2506 | 7189.76
2507 | 7210.05
2508 | 7201.18
2509 | 7186.74
2510 | 7176.07
2511 | 7203.18
2512 | 7190.01
2513 | 7186.03
2514 | 7208.08
2515 | 7183.68
2516 | 7176.75
2517 | 7185.52
2518 | 7185.66
2519 | 7173.49
2520 | 7193.79
2521 | 7200.96
2522 | 7208.37
2523 | 7178.36
2524 | 7188.75
2525 | 7188.42
2526 | 7189.51
2527 | 7178.95
2528 | 7179.39
2529 | 7198.93
2530 | 7191.25
2531 | 7198.41
2532 | 7180.74
2533 | 7187.7
2534 | 7198.19
2535 | 7197.76
2536 | 7195.05
2537 | 7194.96
2538 | 7195.18
2539 | 7188.98
2540 | 7180.76
2541 | 7182.39
2542 | 7199.36
2543 | 7197.39
2544 | 7202.34
2545 | 7199.1
2546 | 7188.73
2547 | 7188.37
2548 | 7190.06
2549 | 7184.09
2550 | 7196.75
2551 | 7200.53
2552 | 7202.34
2553 | 7204.68
2554 | 7191.27
2555 | 7177.3
2556 | 7200.12
2557 | 7190.31
2558 | 7193.88
2559 | 7202.23
2560 | 7203.62
2561 | 7192.32
2562 | 7200.17
2563 | 7181.58
2564 | 7207.45
2565 | 7197.33
2566 | 7192.3
2567 | 7212.98
2568 | 7197.71
2569 | 7212.22
2570 | 7181.38
2571 | 7180.41
2572 | 7197.29
2573 | 7198.3
2574 | 7196.45
2575 | 7182.12
2576 | 7206.48
2577 | 7185.55
2578 | 7192.46
2579 | 7192.38
2580 | 7185.35
2581 | 7183.39
2582 | 7212.79
2583 | 7186.5
2584 | 7194.51
2585 | 7192.22
2586 | 7197.07
2587 | 7194.39
2588 | 7196.54
2589 | 7163.77
2590 | 7200.1
2591 | 7200.16
2592 | 7185.67
2593 | 7180.28
2594 | 7181.82
2595 | 7188.44
2596 | 7184.6
2597 | 7182.34
2598 | 7210.61
2599 | 7194.57
2600 | 7177.65
2601 | 7198.81
2602 | 7166.28
2603 | 7182.47
2604 | 7189.18
2605 | 7191.99
2606 | 7195.45
2607 | 7179.62
2608 | 7177.36
2609 | 7185.81
2610 | 7189.19
2611 | 7199.29
2612 | 7181.07
2613 | 7185.08
2614 | 7164.77
2615 | 7179.65
2616 | 7192.82
2617 | 7176.3
2618 | 7185.22
2619 | 7204.74
2620 | 7182.81
2621 | 7167.92
2622 | 7187.49
2623 | 7184.31
2624 | 7198.87
2625 | 7175.94
2626 | 7196.73
2627 | 7176.64
2628 | 7202.7
2629 | 7182.85
2630 | 7162.38
2631 | 7181.45
2632 | 7189.28
2633 | 7183.16
2634 | 7197.29
2635 | 7173.33
2636 | 7170.85
2637 | 7186.51
2638 | 7189.22
2639 | 7190.57
2640 | 7190.45
2641 | 7182.91
2642 | 7190.85
2643 | 7184.49
2644 | 7196.82
2645 | 7193.89
2646 | 7180.08
2647 | 7192.25
2648 | 7167.73
2649 | 7192.25
2650 | 7193.74
2651 | 7179.82
2652 | 7180.29
2653 | 7184.9
2654 | 7200.79
2655 | 7187.97
2656 | 7184.94
2657 | 7172.18
2658 | 7180.94
2659 | 7187.56
2660 | 7194.08
2661 | 7187.92
2662 | 7193.73
2663 | 7171.21
2664 | 7169.97
2665 | 7194.81
2666 | 7179.56
2667 | 7178.56
2668 | 7183.75
2669 | 7181.4
2670 | 7171.63
2671 | 7175.67
2672 | 7184.01
2673 | 7180.38
2674 | 7183.35
2675 | 7195.5
2676 | 7186.19
2677 | 7163.01
2678 | 7182.8
2679 | 7182
2680 | 7171
2681 | 7184.95
2682 | 7179.88
2683 | 7198.88
2684 | 7173.29
2685 | 7188.38
2686 | 7184.07
2687 | 7191.93
2688 | 7173.66
2689 | 7162.67
2690 | 7183.31
2691 | 7181.54
2692 | 7176.86
2693 | 7187.47
2694 | 7171.19
2695 | 7182.61
2696 | 7176.1
2697 | 7195.86
2698 | 7199.51
2699 | 7195.62
2700 | 7189.88
2701 | 7191.95
2702 | 7188.08
2703 | 7190.39
2704 | 7183.77
2705 | 7183.28
2706 | 7188.21
2707 | 7179.85
2708 | 7173.99
2709 | 7178.7
2710 | 7191.93
2711 | 7181.75
2712 | 7191.23
2713 | 7188.13
2714 | 7173.49
2715 | 7183.46
2716 | 7182.5
2717 | 7185.64
2718 | 7188.83
2719 | 7201.21
2720 | 7191.43
2721 | 7178.9
2722 | 7192.89
2723 | 7197.94
2724 | 7192.58
2725 | 7198.96
2726 | 7169.71
2727 | 7185.1
2728 | 7170.51
2729 | 7177.11
2730 | 7185.52
2731 | 7193.14
2732 | 7195.96
2733 | 7202.06
2734 | 7180.36
2735 | 7165.5
2736 | 7192.03
2737 | 7208.74
2738 | 7174.53
2739 | 7206.77
2740 | 7167.44
2741 | 7194.44
2742 | 7166.71
2743 | 7180.7
2744 | 7192.53
2745 | 7179.55
2746 | 7177.98
2747 | 7184.57
2748 | 7197.64
2749 | 7197.32
2750 | 7179.17
2751 | 7167.7
2752 | 7159.43
2753 | 7180.63
2754 | 7204.81
2755 | 7200.02
2756 | 7203.29
2757 | 7174.45
2758 | 7186.44
2759 | 7197.89
2760 | 7202.72
2761 | 7209.56
2762 | 7209.74
2763 | 7192.13
2764 | 7176.41
2765 | 7175.96
2766 | 7204.63
2767 | 7187.16
2768 | 7206.72
2769 | 7189.51
2770 | 7186.21
2771 | 7203.53
2772 | 7192.96
2773 | 7185.09
2774 | 7187.92
2775 | 7179.69
2776 | 7183.65
2777 | 7181.47
2778 | 7188.65
2779 | 7183.37
2780 | 7197.6
2781 | 7181.04
2782 | 7191.02
2783 | 7167.58
2784 | 7177.32
2785 | 7195.38
2786 | 7202.32
2787 | 7176.12
2788 | 7185.95
2789 | 7185.86
2790 | 7180.12
2791 | 7175.39
2792 | 7180.38
2793 | 7191.62
2794 | 7197.02
2795 | 7189.48
2796 | 7195.76
2797 | 7190.23
2798 | 7193.12
2799 | 7198.88
2800 | 7180.46
2801 | 7195.31
2802 | 7197.71
2803 | 7210.61
2804 | 7189.01
2805 | 7194.12
2806 | 7205.04
2807 | 7185.51
2808 | 7198.95
2809 | 7182.53
2810 | 7185.01
2811 | 7187.45
2812 | 7189.45
2813 | 7185.07
2814 | 7198.72
2815 | 7211.66
2816 | 7180.22
2817 | 7198.26
2818 | 7196.28
2819 | 7190.73
2820 | 7177
2821 | 7194.18
2822 | 7181.56
2823 | 7189.09
2824 | 7190.05
2825 | 7175.67
2826 | 7193.88
2827 | 7195.07
2828 | 7209.39
2829 | 7178.57
2830 | 7178.2
2831 | 7182.65
2832 | 7181.75
2833 | 7200.87
2834 | 7186.65
2835 | 7189.48
2836 | 7201.48
2837 | 7197.24
2838 | 7199.77
2839 | 7171.22
2840 | 7191.16
2841 | 7191.26
2842 | 7185.01
2843 | 7191.99
2844 | 7169.67
2845 | 7188.76
2846 | 7181.7
2847 | 7200.4
2848 | 7175.83
2849 | 7191.25
2850 | 7191.5
2851 | 7172.08
2852 | 7196.21
2853 | 7177.47
2854 | 7187.79
2855 | 7183.46
2856 | 7189.27
2857 | 7171.55
2858 | 7150.19
2859 | 7187.5
2860 | 7172.33
2861 | 7197.26
2862 | 7159.02
2863 | 7177.09
2864 | 7171.54
2865 | 7167.76
2866 | 7185.4
2867 | 7177.9
2868 | 7190.71
2869 | 7167.72
2870 | 7167.73
2871 | 7182.13
2872 | 7164.59
2873 | 7179.2
2874 | 7173.38
2875 | 7187.61
2876 | 7178.22
2877 | 7166.75
2878 | 7198.78
2879 | 7185.39
2880 | 7162.67
2881 | 7193.99
2882 | 7187.37
2883 | 7167.01
2884 | 7179.17
2885 | 7165.79
2886 | 7170.57
2887 | 7176.71
2888 | 7170.48
2889 | 7165.36
2890 | 7187.32
2891 | 7162.94
2892 | 7174.64
2893 | 7171.48
2894 | 7169.61
2895 | 7178.84
2896 | 7172.55
2897 | 7169.73
2898 | 7157.95
2899 | 7193.85
2900 | 7193.13
2901 | 7166.67
2902 | 7179.55
2903 | 7185.21
2904 | 7169.32
2905 | 7174.03
2906 | 7194.45
2907 | 7183.86
2908 | 7178.93
2909 | 7181.29
2910 | 7189.79
2911 | 7177.44
2912 | 7178.69
2913 | 7194.15
2914 | 7192.44
2915 | 7173.58
2916 | 7186.31
2917 | 7174.47
2918 | 7185.54
2919 | 7184.34
2920 | 7183.42
2921 | 7193.15
2922 | 7179.69
2923 | 7221.01
2924 | 7206.16
2925 | 7186.79
2926 | 7183.5
2927 | 7186.7
2928 | 7192.55
2929 | 7197.51
2930 | 7180.78
2931 | 7177.97
2932 | 7187.05
2933 | 7200.76
2934 | 7188.46
2935 | 7186.52
2936 | 7189.11
2937 | 7196.77
2938 | 7187.87
2939 | 7202.97
2940 | 7203.04
2941 | 7211.5
2942 | 7184.41
2943 | 7190.19
2944 | 7186.59
2945 | 7192.73
2946 | 7176.02
2947 | 7194.72
2948 | 7186.81
2949 | 7174.65
2950 | 7189.8
2951 | 7186.64
2952 | 7201.27
2953 | 7204.12
2954 | 7199.97
2955 | 7188.04
2956 | 7187.13
2957 | 7176.97
2958 | 7176.56
2959 | 7195.65
2960 | 7187.72
2961 | 7182.23
2962 | 7192.52
2963 | 7188.29
2964 | 7183.53
2965 | 7168.68
2966 | 7178.84
2967 | 7197.14
2968 | 7175.47
2969 | 7169.77
2970 | 7206.56
2971 | 7190.24
2972 | 7176.24
2973 | 7179.28
2974 | 7199.8
2975 | 7189.53
2976 | 7173.17
2977 | 7185.31
2978 | 7183.15
2979 | 7175.46
2980 | 7178.28
2981 | 7178.04
2982 | 7180.87
2983 | 7168.31
2984 | 7193.28
2985 | 7192.38
2986 | 7188.86
2987 | 7187.77
2988 | 7178.35
2989 | 7197.49
2990 | 7182.78
2991 | 7172.22
2992 | 7180.41
2993 | 7186.36
2994 | 7184.93
2995 | 7191.43
2996 | 7187.23
2997 | 7186.34
2998 | 7189.08
2999 | 7198.9
3000 | 7198.17
3001 | 7183.93
3002 | 7193.37
3003 | 7190.79
3004 | 7187.63
3005 | 7186.1
3006 | 7199.04
3007 | 7177.83
3008 | 7183.6
3009 | 7178.88
3010 | 7190.35
3011 | 7192.8
3012 | 7181.02
3013 | 7181.17
3014 | 7196.18
3015 | 7166.32
3016 | 7198.26
3017 | 7168.21
3018 | 7189.94
3019 | 7186.43
3020 | 7198.12
3021 | 7197
3022 | 7194.87
3023 | 7177.94
3024 | 7168.22
3025 | 7181.81
3026 | 7190.55
3027 | 7177.59
3028 | 7174.47
3029 | 7180.93
3030 | 7188.13
3031 | 7181.56
3032 | 7151.81
3033 | 7188.78
3034 | 7158.45
3035 | 7165.05
3036 | 7184.3
3037 | 7181.42
3038 | 7179.28
3039 | 7180.8
3040 | 7176.83
3041 | 7171.3
3042 | 7174.48
3043 | 7175.88
3044 | 7183.1
3045 | 7195.27
3046 | 7182.47
3047 | 7176.65
3048 | 7181.29
3049 | 7180.03
3050 | 7170.34
3051 | 7184.5
3052 | 7200.21
3053 | 7169.08
3054 | 7183.74
3055 | 7183.39
3056 | 7171.92
3057 | 7170.85
3058 | 7158.31
3059 | 7171.3
3060 | 7178.76
3061 | 7182.48
3062 | 7165.41
3063 | 7201.01
3064 | 7182.33
3065 | 7171.74
3066 | 7184.43
3067 | 7184.57
3068 | 7180.07
3069 | 7181.81
3070 | 7177.85
3071 | 7182.18
3072 | 7174.19
3073 | 7174.61
3074 | 7200.11
3075 | 7189.69
3076 | 7178.03
3077 | 7186.61
3078 | 7172.57
3079 | 7181.2
3080 | 7163.02
3081 | 7179.27
3082 | 7167.71
3083 | 7192.54
3084 | 7185.56
3085 | 7195.76
3086 | 7207.29
3087 | 7167.78
3088 | 7188.83
3089 | 7174.32
3090 | 7185.46
3091 | 7182.54
3092 | 7196.59
3093 | 7188.51
3094 | 7175.11
3095 | 7198.25
3096 | 7171.78
3097 | 7155.73
3098 | 7185.97
3099 | 7196.58
3100 | 7198.64
3101 | 7189.63
3102 | 7159.81
3103 | 7182.83
3104 | 7186.56
3105 | 7168.25
3106 | 7185.59
3107 | 7178.61
3108 | 7177.3
3109 | 7200.41
3110 | 7193.52
3111 | 7185.66
3112 | 7190.31
3113 | 7207.38
3114 | 7185.67
3115 | 7190.46
3116 | 7166.86
3117 | 7185.99
3118 | 7186.45
3119 | 7185.18
3120 | 7190.96
3121 | 7176.63
3122 | 7191.97
3123 | 7185.08
3124 | 7174.88
3125 | 7189.67
3126 | 7198.78
3127 | 7180.07
3128 | 7184.46
3129 | 7181.67
3130 | 7189.59
3131 | 7184.58
3132 | 7191.79
3133 | 7183.71
3134 | 7183.56
3135 | 7192.43
3136 | 7182.34
3137 | 7163.56
3138 | 7191.16
3139 | 7178.2
3140 | 7200.4
3141 | 7188.69
3142 | 7187.18
3143 | 7173.8
3144 | 7195.91
3145 | 7177.76
3146 | 7194.13
3147 | 7172.68
3148 | 7194.78
3149 | 7204.74
3150 | 7169.05
3151 | 7194.94
3152 | 7179.07
3153 | 7187.82
3154 | 7197.34
3155 | 7198.91
3156 | 7168.24
3157 | 7176.78
3158 | 7196.54
3159 | 7191.07
3160 | 7201.82
3161 | 7181.72
3162 | 7184.21
3163 | 7187.76
3164 | 7177.19
3165 | 7205.48
3166 | 7186.73
3167 | 7190.07
3168 | 7177.31
3169 | 7171.32
3170 | 7170.51
3171 | 7174.25
3172 | 7171.51
3173 | 7169.81
3174 | 7183.98
3175 | 7180.11
3176 | 7193.81
3177 | 7187.09
3178 | 7161.6
3179 | 7183.13
3180 | 7173.4
3181 | 7180.29
3182 | 7179.2
3183 | 7174.84
3184 | 7171.05
3185 | 7192.35
3186 | 7191.86
3187 | 7183.24
3188 | 7190.52
3189 | 7173.8
3190 | 7178.87
3191 | 7181.08
3192 | 7170.35
3193 | 7207.27
3194 | 7173.81
3195 | 7187.22
3196 | 7180.77
3197 | 7181.5
3198 | 7170.79
3199 | 7179.09
3200 | 7167.55
3201 | 7180.36
3202 | 7170.33
3203 | 7193.14
3204 | 7193.99
3205 | 7187.16
3206 | 7187.75
3207 | 7205.63
3208 | 7184.02
3209 | 7185.27
3210 | 7178.29
3211 | 7177.07
3212 | 7205.26
3213 | 7178.51
3214 | 7186.45
3215 | 7192.62
3216 | 7182.31
3217 | 7178.86
3218 | 7175.94
3219 | 7186.75
3220 | 7198.97
3221 | 7171.37
3222 | 7197.13
3223 | 7169.77
3224 | 7193.51
3225 | 7194.63
3226 | 7167.96
3227 | 7174.15
3228 | 7168.2
3229 | 7172.5
3230 | 7191.92
3231 | 7188.92
3232 | 7167.15
3233 | 7177.93
3234 | 7186.04
3235 | 7212.39
3236 | 7172.91
3237 | 7177.78
3238 | 7169.35
3239 | 7196.93
3240 | 7159.4
3241 | 7196.05
3242 | 7182.63
3243 | 7171.19
3244 | 7187.42
3245 | 7188.34
3246 | 7175.08
3247 | 7171.94
3248 | 7189.65
3249 | 7199.3
3250 | 7189.04
3251 | 7182.72
3252 | 7182.67
3253 | 7173.59
3254 | 7175.82
3255 | 7174.73
3256 | 7186.79
3257 | 7183.23
3258 | 7197.68
3259 | 7191.17
3260 | 7185.33
3261 | 7173.5
3262 | 7169.14
3263 | 7179.88
3264 | 7184.83
3265 | 7203.54
3266 | 7169.92
3267 | 7173.21
3268 | 7177.98
3269 | 7186.08
3270 | 7182.87
3271 | 7195.13
3272 | 7177.71
3273 | 7202.16
3274 | 7186.2
3275 | 7177.88
3276 | 7200.28
3277 | 7193.19
3278 | 7188.48
3279 | 7191.46
3280 | 7169.53
3281 | 7189.62
3282 | 7195.23
3283 | 7182.52
3284 | 7189.62
3285 | 7161.61
3286 | 7187.3
3287 | 7199.59
3288 | 7160.76
3289 | 7184.07
3290 | 7174.68
3291 | 7188.65
3292 | 7194.35
3293 | 7192.02
3294 | 7188.65
3295 | 7171.42
3296 | 7194.47
3297 | 7188.29
3298 | 7175.14
3299 | 7185.06
3300 | 7175.83
3301 | 7174.71
3302 | 7181.26
3303 | 7167.57
3304 | 7166.8
3305 | 7184.02
3306 | 7183.62
3307 | 7180.31
3308 | 7179.35
3309 | 7170.75
3310 | 7184.36
3311 | 7185.33
3312 | 7193.94
3313 | 7180.35
3314 | 7191.39
3315 | 7185.03
3316 | 7192.23
3317 | 7178.41
3318 | 7178.58
3319 | 7169.25
3320 | 7173.35
3321 | 7194.33
3322 | 7170.49
3323 | 7180.15
3324 | 7184.83
3325 | 7167.03
3326 | 7170.61
3327 | 7189.46
3328 | 7182.7
3329 | 7170.65
3330 | 7170.15
3331 | 7185.27
3332 | 7185.91
3333 | 7189.57
3334 | 7185.93
3335 | 7185.79
3336 | 7195.97
3337 | 7178.98
3338 | 7192.16
3339 | 7187.38
3340 | 7177.96
3341 | 7173.13
3342 | 7196.62
3343 | 7193.78
3344 | 7186.89
3345 | 7179.1
3346 | 7181.99
3347 | 7182.58
3348 | 7177.74
3349 | 7196.35
3350 | 7178.92
3351 | 7184.07
3352 | 7192.43
3353 | 7179.49
3354 | 7166.28
3355 | 7182.23
3356 | 7186.25
3357 | 7182.8
3358 | 7197.42
3359 | 7188.64
3360 | 7182.48
3361 | 7183.84
3362 | 7194.2
3363 | 7185.45
3364 | 7186.18
3365 | 7181.43
3366 | 7184.21
3367 | 7158.78
3368 | 7190.95
3369 | 7184.1
3370 | 7183.36
3371 | 7196.08
3372 | 7171.21
3373 | 7194.1
3374 | 7198.6
3375 | 7167.41
3376 | 7192.41
3377 | 7186.5
3378 | 7193.33
3379 | 7174.76
3380 | 7182.11
3381 | 7166.98
3382 | 7201.84
3383 | 7170.76
3384 | 7184.16
3385 | 7192.62
3386 | 7181.4
3387 | 7172.87
3388 | 7191.71
3389 | 7170.5
3390 | 7188.67
3391 | 7171.59
3392 | 7196.17
3393 | 7178.19
3394 | 7178.66
3395 | 7175.38
3396 | 7167.31
3397 | 7175.04
3398 | 7174.09
3399 | 7199.29
3400 | 7181.02
3401 | 7176.58
3402 | 7173.77
3403 | 7176.44
3404 | 7183.61
3405 | 7189.61
3406 | 7193.77
3407 | 7181.58
3408 | 7186.16
3409 | 7180.79
3410 | 7178.76
3411 | 7171.43
3412 | 7177.38
3413 | 7168.77
3414 | 7189.68
3415 | 7188.82
3416 | 7183.93
3417 | 7205.95
3418 | 7201.7
3419 | 7192.62
3420 | 7175.81
3421 | 7192.36
3422 | 7182.49
3423 | 7168.75
3424 | 7157.48
3425 | 7175.57
3426 | 7182.04
3427 | 7181.47
3428 | 7167.25
3429 | 7174.93
3430 | 7179.65
3431 | 7185.67
3432 | 7163.79
3433 | 7183.38
3434 | 7179.74
3435 | 7182.23
3436 | 7196.6
3437 | 7173.61
3438 | 7186.68
3439 | 7177.25
3440 | 7186.4
3441 | 7184.84
3442 | 7181.93
3443 | 7184.33
3444 | 7183.57
3445 | 7180.55
3446 | 7191.19
3447 | 7201.08
3448 | 7171.84
3449 | 7198.08
3450 | 7173.76
3451 | 7186.7
3452 | 7165.28
3453 | 7174.13
3454 | 7177.19
3455 | 7164.3
3456 | 7185.31
3457 | 7171.23
3458 | 7181.49
3459 | 7171.3
3460 | 7177.65
3461 | 7201.02
3462 | 7181.21
3463 | 7193.43
3464 | 7175.34
3465 | 7171.32
3466 | 7171.21
3467 | 7188.47
3468 | 7190.76
3469 | 7169
3470 | 7180.13
3471 | 7180.03
3472 | 7189.59
3473 | 7192.83
3474 | 7183.15
3475 | 7189.85
3476 | 7190.73
3477 | 7198.06
3478 | 7192.02
3479 | 7192.3
3480 | 7182.77
3481 | 7196.47
3482 | 7180.56
3483 | 7162.65
3484 | 7175.35
3485 | 7191.18
3486 | 7184.09
3487 | 7190.81
3488 | 7175.22
3489 | 7188.78
3490 | 7181.16
3491 | 7193.79
3492 | 7191.77
3493 | 7185.97
3494 | 7170.11
3495 | 7183.84
3496 | 7176.95
3497 | 7169.02
3498 | 7174.41
3499 | 7176.31
3500 | 7194.2
3501 | 7196.23
3502 | 7175.64
3503 | 7182.04
3504 | 7180.72
3505 | 7186.68
3506 | 7181.17
3507 | 7193.8
3508 | 7177.1
3509 | 7188.81
3510 | 7173.02
3511 | 7162.34
3512 | 7184.41
3513 | 7174.63
3514 | 7190.95
3515 | 7177.76
3516 | 7170.47
3517 | 7199.91
3518 | 7176.28
3519 | 7198.92
3520 | 7174.89
3521 | 7187.36
3522 | 7197.08
3523 | 7179.28
3524 | 7171.82
3525 | 7193.47
3526 | 7170.38
3527 | 7191.73
3528 | 7186.79
3529 | 7196.87
3530 | 7176.9
3531 | 7193.16
3532 | 7175.98
3533 | 7185.91
3534 | 7182.6
3535 | 7172.83
3536 | 7183
3537 | 7204.08
3538 | 7189.39
3539 | 7172.62
3540 | 7184.25
3541 | 7189.04
3542 | 7176.9
3543 | 7188.29
3544 | 7208.44
3545 | 7175
3546 | 7203.55
3547 | 7169.71
3548 | 7187.08
3549 | 7191.61
3550 | 7195.21
3551 | 7191.45
3552 | 7158.66
3553 | 7189.95
3554 | 7165.41
3555 | 7182
3556 | 7189.23
3557 | 7178.06
3558 | 7191.55
3559 | 7180.56
3560 | 7180.08
3561 | 7192.69
3562 | 7181.62
3563 | 7177.72
3564 | 7174.46
3565 | 7179.36
3566 | 7190.58
3567 | 7168.72
3568 | 7187.77
3569 | 7184.36
3570 | 7184.03
3571 | 7164.94
3572 | 7174.74
3573 | 7169.34
3574 | 7172
3575 | 7186.03
3576 | 7181.25
3577 | 7182.85
3578 | 7185.09
3579 | 7189.81
3580 | 7167.62
3581 | 7182.65
3582 | 7175.62
3583 | 7179.62
3584 | 7167.47
3585 | 7182.32
3586 | 7188.67
3587 | 7177.86
3588 | 7179.95
3589 | 7173.68
3590 | 7168.27
3591 | 7185.29
3592 | 7170.89
3593 | 7167.3
3594 | 7178.67
3595 | 7173.58
3596 | 7207.38
3597 | 7176.41
3598 | 7191.05
3599 | 7184.76
3600 | 7186.25
3601 | 7192.42
3602 | 7187.98
3603 | 7163.92
3604 | 7184.36
3605 | 7174.25
3606 | 7192.74
3607 | 7179.65
3608 | 7163.84
3609 | 7190.61
3610 | 7181.13
3611 | 7206.32
3612 | 7172.82
3613 | 7172.76
3614 | 7158.16
3615 | 7172.79
3616 | 7177.61
3617 | 7183.57
3618 | 7186.29
3619 | 7197.42
3620 | 7193.08
3621 | 7185.13
3622 | 7183.43
3623 | 7197.06
3624 | 7201.03
3625 | 7167.24
3626 | 7183.88
3627 | 7189.89
3628 | 7182.6
3629 | 7186.04
3630 | 7153.89
3631 | 7177.73
3632 | 7178.7
3633 | 7162.14
3634 | 7173.83
3635 | 7187.5
3636 | 7189.79
3637 | 7170.12
3638 | 7181.1
3639 | 7179.8
3640 | 7177.4
3641 | 7181.19
3642 | 7163.24
3643 | 7201.74
3644 | 7181.21
3645 | 7180.78
3646 | 7168.49
3647 | 7194.91
3648 | 7176.82
3649 | 7185.52
3650 | 7182.77
3651 | 7194.71
3652 | 7194.38
3653 | 7176.87
3654 | 7170.04
3655 | 7178.63
3656 | 7180.88
3657 | 7173.67
3658 | 7182.68
3659 | 7176.6
3660 | 7178.9
3661 | 7186.86
3662 | 7165.14
3663 | 7183.09
3664 | 7182.58
3665 | 7196.39
3666 | 7195.08
3667 | 7189.36
3668 | 7161.8
3669 | 7176.44
3670 | 7183.45
3671 | 7180.07
3672 | 7182.57
3673 | 7170.87
3674 | 7177.45
3675 | 7193.74
3676 | 7172.81
3677 | 7170.59
3678 | 7172.69
3679 | 7178.85
3680 | 7174.44
3681 | 7181.55
3682 | 7182.08
3683 | 7167.81
3684 | 7174.1
3685 | 7173.55
3686 | 7174.29
3687 | 7180.42
3688 | 7163.13
3689 | 7192.25
3690 | 7186.22
3691 | 7175.7
3692 | 7164.31
3693 | 7167.89
3694 | 7178.35
3695 | 7179.36
3696 | 7169.46
3697 | 7172.37
3698 | 7174.97
3699 | 7178.22
3700 | 7172.61
3701 | 7157.2
3702 | 7167.72
3703 | 7166.15
3704 | 7169.86
3705 | 7164.38
3706 | 7155.01
3707 | 7186.62
3708 | 7175.39
3709 | 7183.48
3710 | 7177.64
3711 | 7168.11
3712 | 7182.53
3713 | 7194.09
3714 | 7171.91
3715 | 7174.48
3716 | 7180.19
3717 | 7178.31
3718 | 7162.39
3719 | 7186.07
3720 | 7177.41
3721 | 7162
3722 | 7157.05
3723 | 7167.79
3724 | 7182.96
3725 | 7198.22
3726 | 7176.79
3727 | 7181.7
3728 | 7192.79
3729 | 7182.19
3730 | 7172.02
3731 | 7172.25
3732 | 7182.64
3733 | 7179.97
3734 | 7192.95
3735 | 7152.33
3736 | 7163.44
3737 | 7181.21
3738 | 7181.68
3739 | 7176.61
3740 | 7184.3
3741 | 7192.36
3742 | 7187.07
3743 | 7180.84
3744 | 7160.68
3745 | 7194.24
3746 | 7182.58
3747 | 7181.44
3748 | 7168.29
3749 | 7167.87
3750 | 7174.44
3751 | 7181.59
3752 | 7169.49
3753 | 7172.8
3754 | 7178.02
3755 | 7161.15
3756 | 7185.4
3757 | 7173.54
3758 | 7180.1
3759 | 7201.54
3760 | 7170.73
3761 | 7157.26
3762 | 7170.34
3763 | 7178.76
3764 | 7173.89
3765 | 7163.59
3766 | 7170.28
3767 | 7170.75
3768 | 7151.59
3769 | 7170.42
3770 | 7169.19
3771 | 7180.36
3772 | 7177.9
3773 | 7157.84
3774 | 7183.33
3775 | 7181.95
3776 | 7172.49
3777 | 7170.81
3778 | 7170.6
3779 | 7171.86
3780 | 7149.73
3781 | 7147.21
3782 | 7185.61
3783 | 7157.33
3784 | 7157.05
3785 | 7166.84
3786 | 7181.06
3787 | 7166.54
3788 | 7159.55
3789 | 7169.39
3790 | 7165.68
3791 | 7176.81
3792 | 7185.12
3793 | 7160.77
3794 | 7157.05
3795 | 7168.33
3796 | 7168.73
3797 | 7160.52
3798 | 7163.24
3799 | 7160.56
3800 | 7165.36
3801 | 7171.7
3802 | 7166.15
3803 | 7171.12
3804 | 7170.93
3805 | 7151.69
3806 | 7158.45
3807 | 7152.93
3808 | 7168.25
3809 | 7165.93
3810 | 7159.86
3811 | 7158.68
3812 | 7157
3813 | 7133.12
3814 | 7144.34
3815 | 7152.8
3816 | 7150.43
3817 | 7151.99
3818 | 7151.86
3819 | 7144.8
3820 | 7147.03
3821 | 7158.81
3822 | 7141.71
3823 | 7173.13
3824 | 7149.6
3825 | 7160.8
3826 | 7176.37
3827 | 7137.85
3828 | 7158.38
3829 | 7150.92
3830 | 7135.36
3831 | 7152.68
3832 | 7141.45
3833 | 7163.28
3834 | 7165.29
3835 | 7159.97
3836 | 7148
3837 | 7165.35
3838 | 7172.36
3839 | 7148.65
3840 | 7159.14
3841 | 7146.4
3842 | 7162.88
3843 | 7166.73
3844 | 7169.95
3845 | 7166.64
3846 | 7168.28
3847 | 7176.54
3848 | 7170.92
3849 | 7166.7
3850 | 7167.15
3851 | 7182.23
3852 | 7176.49
3853 | 7171.24
3854 | 7169.05
3855 | 7164.87
3856 | 7175.44
3857 | 7168.08
3858 | 7159.83
3859 | 7166.25
3860 | 7160.15
3861 | 7166.85
3862 | 7172.28
3863 | 7187.01
3864 | 7164.82
3865 | 7181.12
3866 | 7179.67
3867 | 7166.89
3868 | 7155.15
3869 | 7161.66
3870 | 7175.56
3871 | 7158.33
3872 | 7161.38
3873 | 7182.02
3874 | 7162.46
3875 | 7181.63
3876 | 7177.59
3877 | 7162.21
3878 | 7177.09
3879 | 7170.02
3880 | 7166.66
3881 | 7171.52
3882 | 7158.49
3883 | 7153.11
3884 | 7178.83
3885 | 7166.75
3886 | 7180.24
3887 | 7148.27
3888 | 7193.04
3889 | 7169.81
3890 | 7189.24
3891 | 7161.72
3892 | 7163.67
3893 | 7172.51
3894 | 7166.39
3895 | 7163.94
3896 | 7176.18
3897 | 7163.48
3898 | 7167.35
3899 | 7182.91
3900 | 7169.67
3901 | 7181.33
3902 | 7163.23
3903 | 7167.28
3904 | 7180.65
3905 | 7166.96
3906 | 7165.08
3907 | 7173.38
3908 | 7174.93
3909 | 7173.83
3910 | 7155.63
3911 | 7167.21
3912 | 7170.28
3913 | 7183.65
3914 | 7150.96
3915 | 7176.07
3916 | 7170.65
3917 | 7154.67
3918 | 7183.13
3919 | 7152.53
3920 | 7155.79
3921 | 7165.9
3922 | 7171.63
3923 | 7164.81
3924 | 7167.31
3925 | 7160.25
3926 | 7183.25
3927 | 7170.13
3928 | 7164.61
3929 | 7160.44
3930 | 7160.58
3931 | 7164.01
3932 | 7136.6
3933 | 7168.22
3934 | 7171.35
3935 | 7170.83
3936 | 7176.52
3937 | 7179.91
3938 | 7162.46
3939 | 7176.76
3940 | 7159.73
3941 | 7181.1
3942 | 7154.63
3943 | 7174.86
3944 | 7181.77
3945 | 7183.84
3946 | 7165.36
3947 | 7155.5
3948 | 7185.26
3949 | 7179.19
3950 | 7157.72
3951 | 7159.39
3952 | 7167.11
3953 | 7172.11
3954 | 7163.42
3955 | 7179.03
3956 | 7183.04
3957 | 7164.65
3958 | 7173.85
3959 | 7176.51
3960 | 7169.67
3961 | 7191.37
3962 | 7172.6
3963 | 7169.88
3964 | 7166.02
3965 | 7171.81
3966 | 7170.7
3967 | 7197.44
3968 | 7182.11
3969 | 7171.99
3970 | 7181.38
3971 | 7187.16
3972 | 7177.25
3973 | 7179.97
3974 | 7180.14
3975 | 7173.17
3976 | 7195.55
3977 | 7174.69
3978 | 7169.83
3979 | 7194.27
3980 | 7180.94
3981 | 7178.3
3982 | 7183.76
3983 | 7181.78
3984 | 7179.06
3985 | 7185.8
3986 | 7171.31
3987 | 7179.18
3988 | 7175.1
3989 | 7194.21
3990 | 7165.87
3991 | 7197.6
3992 | 7170.49
3993 | 7181.88
3994 | 7179.03
3995 | 7179.29
3996 | 7155.99
3997 | 7174.13
3998 | 7182.74
3999 | 7183.78
4000 | 7172.82
4001 | 7191.54
4002 | 7161.54
4003 | 7188.46
4004 | 7200.4
4005 | 7186.02
4006 | 7175.05
4007 | 7177.53
4008 | 7165.88
4009 | 7179.51
4010 | 7181.12
4011 | 7179.68
4012 | 7148.37
4013 | 7174.4
4014 | 7174.51
4015 | 7192.51
4016 | 7173.52
4017 | 7187.52
4018 | 7185.52
4019 | 7156.76
4020 | 7154.68
4021 | 7180.34
4022 | 7195.31
4023 | 7167
4024 | 7182.52
4025 | 7185.12
4026 | 7177.61
4027 | 7183.59
4028 | 7173.84
4029 | 7164.98
4030 | 7195.03
4031 | 7182.68
4032 | 7189.86
4033 | 7173.08
4034 | 7170.95
4035 | 7165.81
4036 | 7186.66
4037 | 7156.5
4038 | 7165.43
4039 | 7172.39
4040 | 7184.51
4041 | 7171.63
4042 | 7177.25
4043 | 7174.06
4044 | 7195.34
4045 | 7169.06
4046 | 7171.17
4047 | 7167.25
4048 | 7163.45
4049 | 7175.73
4050 | 7185.52
4051 | 7190.8
4052 | 7164.63
4053 | 7184.66
4054 | 7174.08
4055 | 7192.12
4056 | 7180.36
4057 | 7184.59
4058 | 7177.13
4059 | 7181.5
4060 | 7180.3
4061 | 7159.27
4062 | 7163.58
4063 | 7170.7
4064 | 7183.3
4065 | 7184.92
4066 | 7172.21
4067 | 7185.33
4068 | 7179.3
4069 | 7174.79
4070 | 7164.54
4071 | 7161.05
4072 | 7171.15
4073 | 7175.61
4074 | 7155.83
4075 | 7176.61
4076 | 7167.45
4077 | 7159.94
4078 | 7172.35
4079 | 7169.4
4080 | 7191.58
4081 | 7147.91
4082 | 7171.2
4083 | 7160
4084 | 7161.36
4085 | 7162.43
4086 | 7180.3
4087 | 7173.04
4088 | 7153.65
4089 | 7155.92
4090 | 7167.92
4091 | 7165.21
4092 | 7160.29
4093 | 7159.35
4094 | 7158.22
4095 | 7159.16
4096 | 7169.63
4097 | 7148.16
4098 | 7174.7
4099 | 7171.63
4100 | 7165.33
4101 | 7165.21
4102 | 7154.5
4103 | 7178.46
4104 | 7161.89
4105 | 7164.66
4106 | 7163.9
4107 | 7171.34
4108 | 7152.51
4109 | 7166.25
4110 | 7157.65
4111 | 7168.46
4112 | 7183.47
4113 | 7170.59
4114 | 7189.15
4115 | 7175.91
4116 | 7162.28
4117 | 7188.67
4118 | 7171.67
4119 | 7170.94
4120 | 7179.06
4121 | 7177.08
4122 | 7182.83
4123 | 7185.81
4124 | 7166.68
4125 | 7175.44
4126 | 7176.04
4127 | 7173.37
4128 | 7170.19
4129 | 7175.27
4130 | 7161.37
4131 | 7167.67
4132 | 7178.15
4133 | 7176.57
4134 | 7168.4
4135 | 7175.12
4136 | 7172.82
4137 | 7165.98
4138 | 7182.34
4139 | 7160.11
4140 | 7190.32
4141 | 7189.3
4142 | 7181.59
4143 | 7173.61
4144 | 7176.87
4145 | 7167.03
4146 | 7172.31
4147 | 7175.16
4148 | 7184.19
4149 | 7151.06
4150 | 7190.05
4151 | 7165.49
4152 | 7176.49
4153 | 7159.93
4154 | 7158.22
4155 | 7169.59
4156 | 7152.22
4157 | 7172.78
4158 | 7181.35
4159 | 7173.4
4160 | 7190.17
4161 | 7184.6
4162 | 7179.69
4163 | 7176.98
4164 | 7171.54
4165 | 7179.79
4166 | 7162.28
4167 | 7179.48
4168 | 7169.23
4169 | 7165.63
4170 | 7167.44
4171 | 7162.8
4172 | 7180.15
4173 | 7198.76
4174 | 7180.49
4175 | 7163.14
4176 | 7173.5
4177 | 7194.12
4178 | 7169.69
4179 | 7165.86
4180 | 7180.24
4181 | 7178.63
4182 | 7178.44
4183 | 7179.68
4184 | 7157.92
4185 | 7182.92
4186 | 7186.49
4187 | 7165.45
4188 | 7174.56
4189 | 7184.49
4190 | 7159.38
4191 | 7175.37
4192 | 7171.09
4193 | 7174.19
4194 | 7186.73
4195 | 7162.63
4196 | 7176.79
4197 | 7177.15
4198 | 7175.34
4199 | 7165.4
4200 | 7187.85
4201 | 7166.38
4202 | 7160.65
4203 | 7164.65
4204 | 7179.06
4205 | 7162.4
4206 | 7162.42
4207 | 7183.21
4208 | 7183.58
4209 | 7172.07
4210 | 7181.47
4211 | 7164.4
4212 | 7185.83
4213 | 7193.04
4214 | 7181.26
4215 | 7191.77
4216 | 7177.9
4217 | 7180.83
4218 | 7178.33
4219 | 7175.11
4220 | 7168.85
4221 | 7159.5
4222 | 7165.45
4223 | 7163.25
4224 | 7177.23
4225 | 7170.46
4226 | 7179.15
4227 | 7168.75
4228 | 7151.42
4229 | 7159.65
4230 | 7182.08
4231 | 7176.66
4232 | 7168.39
4233 | 7178.72
4234 | 7177.34
4235 | 7178.81
4236 | 7167.45
4237 | 7175.43
4238 | 7175.66
4239 | 7167.64
4240 | 7163.81
4241 | 7171.14
4242 | 7169.04
4243 | 7185.07
4244 | 7173.32
4245 | 7168.21
4246 | 7160.33
4247 | 7182.54
4248 | 7159.19
4249 | 7186.73
4250 | 7180.55
4251 | 7174.67
4252 | 7155.56
4253 | 7166.65
4254 | 7154.9
4255 | 7143.67
4256 | 7196.58
4257 | 7165.96
4258 | 7166.38
4259 | 7178.88
4260 | 7173.25
4261 | 7179.25
4262 | 7167.42
4263 | 7179.82
4264 | 7166.33
4265 | 7168
4266 | 7163.68
4267 | 7157.8
4268 | 7182.72
4269 | 7177.66
4270 | 7162.48
4271 | 7158.56
4272 | 7159.1
4273 | 7174.71
4274 | 7172.42
4275 | 7167.27
4276 | 7177.29
4277 | 7164.61
4278 | 7180.11
4279 | 7188.54
4280 | 7163.56
4281 | 7164.76
4282 | 7176.6
4283 | 7162.31
4284 | 7167.56
4285 | 7177.43
4286 | 7178.71
4287 | 7164.38
4288 | 7160.97
4289 | 7183.83
4290 | 7158.29
4291 | 7176.84
4292 | 7173.55
4293 | 7163.41
4294 | 7165.16
4295 | 7183.87
4296 | 7178.34
4297 | 7171.75
4298 | 7174.05
4299 | 7177.54
4300 | 7180.46
4301 | 7170.76
4302 | 7175.93
4303 | 7156.64
4304 | 7187.52
4305 | 7162.94
4306 | 7178.94
4307 | 7184.38
4308 | 7167.8
4309 | 7166.81
4310 | 7182.32
4311 | 7172.73
4312 | 7167.17
4313 | 7165.49
4314 | 7190.55
4315 | 7178.11
4316 | 7152.13
4317 | 7169.57
4318 | 7171.62
4319 | 7159.71
4320 | 7172.73
4321 | 7159.97
4322 | 7178.01
4323 | 7181.71
4324 | 7179.07
4325 | 7191.31
4326 | 7168.48
4327 | 7179.89
4328 | 7164.77
4329 | 7178.19
4330 | 7185.77
4331 | 7183.31
4332 | 7167.23
4333 | 7164.75
4334 | 7173.5
4335 | 7200.87
4336 | 7183.75
4337 | 7184.39
4338 | 7178.72
4339 | 7195.59
4340 | 7177.7
4341 | 7174.92
4342 | 7176.85
4343 | 7172.4
4344 | 7164.26
4345 | 7173.91
4346 | 7166.12
4347 | 7179.05
4348 | 7158.8
4349 | 7175.32
4350 | 7165.7
4351 | 7175.34
4352 | 7179.48
4353 | 7187.34
4354 | 7167.77
4355 | 7185.25
4356 | 7178.98
4357 | 7177.22
4358 | 7162.34
4359 | 7161.31
4360 | 7162.27
4361 | 7168.84
4362 | 7184.07
4363 | 7158.16
4364 | 7194.63
4365 | 7179.03
4366 | 7172.93
4367 | 7173.09
4368 | 7168.62
4369 | 7168.14
4370 | 7157
4371 | 7169.15
4372 | 7168.22
4373 | 7194.2
4374 | 7188.54
4375 | 7162.34
4376 | 7181.26
4377 | 7180.7
4378 | 7169.56
4379 | 7158.08
4380 | 7170.6
4381 | 7186.24
4382 | 7173.55
4383 | 7176.79
4384 | 7183.52
4385 | 7167.68
4386 | 7159.81
4387 | 7170.21
4388 | 7161.46
4389 | 7154.4
4390 | 7178.84
4391 | 7159.42
4392 | 7162.16
4393 | 7161.87
4394 | 7171.67
4395 | 7182.84
4396 | 7161.66
4397 | 7175.41
4398 | 7172.08
4399 | 7163.29
4400 | 7176.61
4401 | 7173.38
4402 | 7164.86
4403 | 7197.21
4404 | 7159.7
4405 | 7163.51
4406 | 7170.33
4407 | 7155.93
4408 | 7167.08
4409 | 7170.45
4410 | 7172.95
4411 | 7162.4
4412 | 7172.06
4413 | 7174.8
4414 | 7181.48
4415 | 7181.76
4416 | 7176.74
4417 | 7155.07
4418 | 7166.32
4419 | 7175.14
4420 | 7171.64
4421 | 7161.62
4422 | 7171.24
4423 | 7155.86
4424 | 7164.13
4425 | 7150.5
4426 | 7161.14
4427 | 7159.12
4428 | 7186.9
4429 | 7166.12
4430 | 7162.17
4431 | 7158.07
4432 | 7182.21
4433 | 7188.65
4434 | 7176.31
4435 | 7180.9
4436 | 7174.05
4437 | 7163.22
4438 | 7165.19
4439 | 7156.4
4440 | 7166.94
4441 | 7165.92
4442 | 7183.06
4443 | 7172.06
4444 | 7157.59
4445 | 7174.6
4446 | 7178.76
4447 | 7167.69
4448 | 7162.48
4449 | 7170.45
4450 | 7181.94
4451 | 7157.63
4452 | 7175.27
4453 | 7157.18
4454 | 7159.36
4455 | 7179.48
4456 | 7155.96
4457 | 7176.22
4458 | 7175.61
4459 | 7180.93
4460 | 7160.89
4461 | 7172.64
4462 | 7171.67
4463 | 7169.53
4464 | 7182.64
4465 | 7176.39
4466 | 7175.14
4467 | 7177.59
4468 | 7158.67
4469 | 7159.16
4470 | 7161.23
4471 | 7170.11
4472 | 7175.78
4473 | 7177.18
4474 | 7169.57
4475 | 7191.43
4476 | 7156.11
4477 | 7164.58
4478 | 7162.85
4479 | 7172.2
4480 | 7167.56
4481 | 7163.73
4482 | 7173.33
4483 | 7173.13
4484 | 7183.57
4485 | 7172.41
4486 | 7190.29
4487 | 7180.68
4488 | 7162.1
4489 | 7154.95
4490 | 7188.07
4491 | 7161.19
4492 | 7162.84
4493 | 7183.79
4494 | 7152.41
4495 | 7158.34
4496 | 7161.59
4497 | 7165.97
4498 | 7146.55
4499 | 7180.09
4500 | 7170.37
4501 | 7157.43
4502 | 7179.56
4503 | 7152.47
4504 | 7173.26
4505 | 7166.73
4506 | 7172.18
4507 | 7158.64
4508 | 7165.31
4509 | 7169.75
4510 | 7158.76
4511 | 7181.44
4512 | 7175.5
4513 | 7179.03
4514 | 7172.77
4515 | 7179.17
4516 | 7172.65
4517 | 7166.19
4518 | 7157.76
4519 | 7175.86
4520 | 7163.23
4521 | 7178.58
4522 | 7176.48
4523 | 7158.09
4524 | 7154.65
4525 | 7183.72
4526 | 7160.9
4527 | 7166.43
4528 | 7162.76
4529 | 7162.01
4530 | 7163.06
4531 | 7174.19
4532 | 7182.52
4533 | 7164.43
4534 | 7153.39
4535 | 7182.54
4536 | 7174.75
4537 | 7170.92
4538 | 7196.64
4539 | 7167.58
4540 | 7183.74
4541 | 7176.12
4542 | 7173.01
4543 | 7182.68
4544 | 7169.27
4545 | 7173.9
4546 | 7162.6
4547 | 7178.98
4548 | 7167.74
4549 | 7168.87
4550 | 7155.1
4551 | 7180.23
4552 | 7188.68
4553 | 7177.86
4554 | 7165.48
4555 | 7167.84
4556 | 7170.94
4557 | 7160.49
4558 | 7164.45
4559 | 7172.85
4560 | 7165.48
4561 | 7184.27
4562 | 7179.94
4563 | 7177.9
4564 | 7165.75
4565 | 7163.14
4566 | 7169.14
4567 | 7186.22
4568 | 7181.04
4569 | 7170.3
4570 | 7170.82
4571 | 7176.4
4572 | 7184.72
4573 | 7158.92
4574 | 7160.12
4575 | 7161.81
4576 | 7165.91
4577 | 7182.34
4578 | 7169.03
4579 | 7168.09
4580 | 7165.7
4581 | 7174.91
4582 | 7163.8
4583 | 7180.42
4584 | 7161.17
4585 | 7161.75
4586 | 7174.74
4587 | 7175.72
4588 | 7192.99
4589 | 7171.97
4590 | 7185.24
4591 | 7178.9
4592 | 7170.67
4593 | 7171.22
4594 | 7169.01
4595 | 7161.74
4596 | 7191.27
4597 | 7176.14
4598 | 7164.46
4599 | 7178.81
4600 | 7171.23
4601 | 7167.55
4602 | 7163.55
4603 | 7183.2
4604 | 7174.01
4605 | 7151.38
4606 | 7179.12
4607 | 7173.36
4608 | 7186.78
4609 | 7186.16
4610 | 7169.54
4611 | 7163.19
4612 | 7167.73
4613 | 7176.21
4614 | 7142.37
4615 | 7168.62
4616 | 7162.11
4617 | 7168.75
4618 | 7155.95
4619 | 7165.41
4620 | 7147.61
4621 | 7165.35
4622 | 7172.4
4623 | 7162.47
4624 | 7158.84
4625 | 7169.75
4626 | 7169.9
4627 | 7162.53
4628 | 7174.47
4629 | 7165.02
4630 | 7170.06
4631 | 7150.31
4632 | 7163.01
4633 | 7162.32
4634 | 7185.96
4635 | 7163.66
4636 | 7172.45
4637 | 7158.04
4638 | 7170.02
4639 | 7157.06
4640 | 7176.21
4641 | 7167.77
4642 | 7163.12
4643 | 7165.68
4644 | 7177.51
4645 | 7156.54
4646 | 7169.23
4647 | 7150.45
4648 | 7172.65
4649 | 7180.86
4650 | 7154.46
4651 | 7153.46
4652 | 7166.31
4653 | 7167
4654 | 7157.31
4655 | 7182.63
4656 | 7147.35
4657 | 7148.56
4658 | 7166.1
4659 | 7161.97
4660 | 7158.4
4661 | 7157.36
4662 | 7163.92
4663 | 7169.21
4664 | 7162.37
4665 | 7153.56
4666 | 7163.39
4667 | 7161.53
4668 | 7160.56
4669 | 7167.23
4670 | 7143.75
4671 | 7154.22
4672 | 7167.5
4673 | 7169.2
4674 | 7164.15
4675 | 7167.39
4676 | 7163.45
4677 | 7156.18
4678 | 7167.08
4679 | 7174.66
4680 | 7162
4681 | 7146.7
4682 | 7159.98
4683 | 7155.64
4684 | 7164.88
4685 | 7181.65
4686 | 7167.23
4687 | 7162.1
4688 | 7147.46
4689 | 7152.17
4690 | 7161.98
4691 | 7174.17
4692 | 7168.17
4693 | 7157.45
4694 | 7155.56
4695 | 7167.89
4696 | 7172.86
4697 | 7161.99
4698 | 7178.23
4699 | 7168.98
4700 | 7151.52
4701 | 7166.81
4702 | 7161.16
4703 | 7136.47
4704 | 7159.25
4705 | 7161.89
4706 | 7176.51
4707 | 7154.06
4708 | 7150.22
4709 | 7164.11
4710 | 7147.86
4711 | 7139.95
4712 | 7146.52
4713 | 7148.11
4714 | 7165.77
4715 | 7160.54
4716 | 7148.99
4717 | 7154.91
4718 | 7156.99
4719 | 7150.3
4720 | 7176.03
4721 | 7142.33
4722 | 7154.6
4723 | 7166.56
4724 | 7155.32
4725 | 7167.55
4726 | 7149.83
4727 | 7163.44
4728 | 7161.28
4729 | 7146.71
4730 | 7151.44
4731 | 7148.81
4732 | 7138.39
4733 | 7153.23
4734 | 7149.1
4735 | 7142.36
4736 | 7172.89
4737 | 7144.24
4738 | 7161.68
4739 | 7124.27
4740 | 7166.49
4741 | 7141.58
4742 | 7156.52
4743 | 7153.75
4744 | 7157.24
4745 | 7141.58
4746 | 7157.2
4747 | 7169.85
4748 | 7154.68
4749 | 7163.6
4750 | 7162.35
4751 | 7164.83
4752 | 7161.57
4753 | 7155.37
4754 | 7153.93
4755 | 7175.87
4756 | 7156.67
4757 | 7157.82
4758 | 7168.83
4759 | 7166.91
4760 | 7161.01
4761 | 7171.4
4762 | 7156.82
4763 | 7142.2
4764 | 7164.5
4765 | 7173.07
4766 | 7166.38
4767 | 7145.73
4768 | 7169.72
4769 | 7162.74
4770 | 7159.1
4771 | 7172.42
4772 | 7146.81
4773 | 7147.76
4774 | 7158.28
4775 | 7169.51
4776 | 7163.17
4777 | 7162.49
4778 | 7156.38
4779 | 7166.35
4780 | 7152.96
4781 | 7162.29
4782 | 7158.07
4783 | 7156.02
4784 | 7166.42
4785 | 7172.97
4786 | 7157.28
4787 | 7165.75
4788 | 7162.85
4789 | 7176.98
4790 | 7153.25
4791 | 7159.52
4792 | 7145.32
4793 | 7168.17
4794 | 7176.89
4795 | 7170.93
4796 | 7184.16
4797 | 7186.1
4798 | 7168.73
4799 | 7174.33
4800 | 7158.3
4801 | 7185.16
4802 | 7170.77
4803 | 7193.16
4804 | 7157.14
4805 | 7153.36
4806 | 7156.36
4807 | 7174.18
4808 | 7154.96
4809 | 7173.37
4810 | 7147.82
4811 | 7162.47
4812 | 7163.86
4813 | 7158
4814 | 7152.61
4815 | 7171.22
4816 | 7186.55
4817 | 7171.29
4818 | 7158.25
4819 | 7148.03
4820 | 7177.98
4821 | 7161.02
4822 | 7189.73
4823 | 7161.09
4824 | 7178.58
4825 | 7169.12
4826 | 7157.54
4827 | 7173.61
4828 | 7186.06
4829 | 7163.96
4830 | 7160.39
4831 | 7169.91
4832 | 7192.28
4833 | 7175.12
4834 | 7172.52
4835 | 7173.83
4836 | 7175.13
4837 | 7189.14
4838 | 7179.32
4839 | 7186.11
4840 | 7182.77
4841 | 7179.25
4842 | 7169.18
4843 | 7182.64
4844 | 7178.83
4845 | 7162.26
4846 | 7184.19
4847 | 7175.56
4848 | 7162
4849 | 7196.05
4850 | 7168.64
4851 | 7154.38
4852 | 7164.47
4853 | 7178.59
4854 | 7158.23
4855 | 7173.05
4856 | 7184.11
4857 | 7169.42
4858 | 7180.63
4859 | 7179.39
4860 | 7180
4861 | 7181.08
4862 | 7172.99
4863 | 7165.41
4864 | 7182.78
4865 | 7172.64
4866 | 7168.45
4867 | 7174.92
4868 | 7186.48
4869 | 7177.73
4870 | 7171.38
4871 | 7176.26
4872 | 7188.03
4873 | 7171.2
4874 | 7173.15
4875 | 7162.06
4876 | 7164.49
4877 | 7150.91
4878 | 7153.52
4879 | 7166.25
4880 | 7168.83
4881 | 7173.89
4882 | 7169.05
4883 | 7160.98
4884 | 7164.38
4885 | 7186.78
4886 | 7164.53
4887 | 7181.34
4888 | 7156.85
4889 | 7161.32
4890 | 7172.35
4891 | 7161.94
4892 | 7176.29
4893 | 7168.32
4894 | 7162.98
4895 | 7181.73
4896 | 7175.39
4897 | 7170.48
4898 | 7159.67
4899 | 7154.8
4900 | 7163.36
4901 | 7173.64
4902 | 7167.11
4903 | 7170.92
4904 | 7168.34
4905 | 7181.89
4906 | 7182.22
4907 | 7154.67
4908 | 7168.94
4909 | 7161.92
4910 | 7153.18
4911 | 7145.7
4912 | 7149.49
4913 | 7163.15
4914 | 7176.77
4915 | 7176.09
4916 | 7181.76
4917 | 7158.59
4918 | 7171.47
4919 | 7182.82
4920 | 7156.6
4921 | 7159.52
4922 | 7157.06
4923 | 7192.88
4924 | 7168.22
4925 | 7163.47
4926 | 7178.39
4927 | 7180.67
4928 | 7164.08
4929 | 7166.12
4930 | 7155.21
4931 | 7169.33
4932 | 7153.35
4933 | 7178.28
4934 | 7169.61
4935 | 7162.87
4936 | 7172.63
4937 | 7153.04
4938 | 7166.09
4939 | 7155.55
4940 | 7155.79
4941 | 7150.96
4942 | 7144.64
4943 | 7161.04
4944 | 7166.71
4945 | 7185.4
4946 | 7166.79
4947 | 7176.2
4948 | 7178.49
4949 | 7171.19
4950 | 7156.78
4951 | 7164.58
4952 | 7184.38
4953 | 7163.46
4954 | 7167.2
4955 | 7166
4956 | 7160.03
4957 | 7183.18
4958 | 7176.06
4959 | 7176.64
4960 | 7178.52
4961 | 7158.59
4962 | 7174.51
4963 | 7170.92
4964 | 7164.62
4965 | 7170.33
4966 | 7165.61
4967 | 7166.99
4968 | 7164.62
4969 | 7168.66
4970 | 7178.5
4971 | 7168.49
4972 | 7181.26
4973 | 7182.13
4974 | 7157
4975 | 7151.39
4976 | 7168.3
4977 | 7164.97
4978 | 7153.54
4979 | 7181.63
4980 | 7171.81
4981 | 7175.73
4982 | 7149.16
4983 | 7165.88
4984 | 7167.87
4985 | 7151.49
4986 | 7152.74
4987 | 7153.7
4988 | 7166.74
4989 | 7168.11
4990 | 7178.28
4991 | 7196.01
4992 | 7170.17
4993 | 7155.18
4994 | 7154.4
4995 | 7164.52
4996 | 7161.92
4997 | 7157.41
4998 | 7180.03
4999 | 7157.55
5000 | 7174.48
5001 | 7166.94
5002 | 7151.33
5003 | 7150.04
5004 | 7158.4
5005 | 7170.39
5006 | 7162.33
5007 | 7160.7
5008 | 7174.08
5009 | 7153.82
5010 | 7159.71
5011 | 7173.07
5012 | 7163.55
5013 | 7169.77
5014 | 7155.91
5015 | 7172.41
5016 | 7159.96
5017 | 7179.27
5018 | 7153.24
5019 | 7168.28
5020 | 7162.13
5021 | 7164.06
5022 | 7175.54
5023 | 7175.02
5024 | 7171.31
5025 | 7151.37
5026 | 7186.89
5027 | 7182.9
5028 | 7157.85
5029 | 7176.54
5030 | 7162.91
5031 | 7169.01
5032 | 7160.65
5033 | 7170.24
5034 | 7162.97
5035 | 7166.78
5036 | 7172.85
5037 | 7170.3
5038 | 7159.46
5039 | 7163.45
5040 | 7163.86
5041 | 7158.22
5042 | 7175.41
5043 | 7182.77
5044 | 7170.42
5045 | 7173.53
5046 | 7176.06
5047 | 7167.29
5048 | 7174.24
5049 | 7152.92
5050 | 7159.93
5051 | 7177.53
5052 | 7161.33
5053 | 7169.48
5054 | 7153.33
5055 | 7154.06
5056 | 7167.82
5057 | 7186.55
5058 | 7163.43
5059 | 7160.36
5060 | 7169.56
5061 | 7170.19
5062 | 7159.85
5063 | 7169.17
5064 | 7165.26
5065 | 7168.38
5066 | 7170.33
5067 | 7161.71
5068 | 7184.26
5069 | 7167.29
5070 | 7181.76
5071 | 7166.85
5072 | 7171.4
5073 | 7172.47
5074 | 7180.04
5075 | 7173.68
5076 | 7186.61
5077 | 7153.81
5078 | 7172.7
5079 | 7166.32
5080 | 7169.19
5081 | 7168.96
5082 | 7162.44
5083 | 7171.93
5084 | 7154.44
5085 | 7161.64
5086 | 7177.78
5087 | 7165.63
5088 | 7165.3
5089 | 7156.32
5090 | 7162.78
5091 | 7169.37
5092 | 7173.42
5093 | 7172.96
5094 | 7170.13
5095 | 7164.01
5096 | 7163.89
5097 | 7178.47
5098 | 7163.82
5099 | 7158.37
5100 | 7183.2
5101 | 7181.76
5102 | 7182.09
5103 | 7153.62
5104 | 7157.34
5105 | 7175.88
5106 | 7166.08
5107 | 7164.06
5108 | 7163.92
5109 | 7161.15
5110 | 7178.68
5111 | 7143.47
5112 | 7180.64
5113 | 7183.87
5114 | 7163.36
5115 | 7178.96
5116 | 7173.49
5117 | 7178.87
5118 | 7155.11
5119 | 7148.83
5120 | 7164.46
5121 | 7175.07
5122 | 7171.06
5123 | 7158.38
5124 | 7151.33
5125 | 7170.36
5126 | 7188.72
5127 | 7158.16
5128 | 7169.36
5129 | 7181.93
5130 | 7179.41
5131 | 7175.44
5132 | 7152.71
5133 | 7158.57
5134 | 7156.66
5135 | 7176.56
5136 | 7158.81
5137 | 7187.52
5138 | 7171.99
5139 | 7160.56
5140 | 7168.16
5141 | 7161.07
5142 | 7167.88
5143 | 7168.8
5144 | 7160.59
5145 | 7158.99
5146 | 7173.18
5147 | 7154.34
5148 | 7182.69
5149 | 7176.45
5150 | 7160.95
5151 | 7166.66
5152 | 7156.91
5153 | 7157.64
5154 | 7164.91
5155 | 7174.2
5156 | 7157.58
5157 | 7164.11
5158 | 7161.21
5159 | 7174.84
5160 | 7162.95
5161 | 7181.24
5162 | 7165.77
5163 | 7189.21
5164 | 7185.99
5165 | 7156.25
5166 | 7159
5167 | 7161.89
5168 | 7168.94
5169 | 7177.25
5170 | 7183.14
5171 | 7170.67
5172 | 7166.69
5173 | 7159.74
5174 | 7164.97
5175 | 7191.62
5176 | 7147.03
5177 | 7173.27
5178 | 7160.84
5179 | 7155.11
5180 | 7169.25
5181 | 7171.41
5182 | 7181.38
5183 | 7183.08
5184 | 7171.21
5185 | 7174.45
5186 | 7170.69
5187 | 7160.65
5188 | 7157.61
5189 | 7164.22
5190 | 7164.43
5191 | 7177.34
5192 | 7160.23
5193 | 7166.53
5194 | 7157.28
5195 | 7179.18
5196 | 7167.09
5197 | 7150.97
5198 | 7158.72
5199 | 7176.2
5200 | 7169.82
5201 | 7152.54
5202 | 7152.05
5203 | 7157.91
5204 | 7167.9
5205 | 7170.75
5206 | 7163.07
5207 | 7169.65
5208 | 7161.46
5209 | 7174.37
5210 | 7146.53
5211 | 7181.04
5212 | 7178.15
5213 | 7160.27
5214 | 7169.71
5215 | 7159.73
5216 | 7159.54
5217 | 7157.14
5218 | 7160.81
5219 | 7152.01
5220 | 7165.94
5221 | 7154.84
5222 | 7165.26
5223 | 7169.84
5224 | 7166.7
5225 | 7171.98
5226 | 7185.04
5227 | 7174.13
5228 | 7163.44
5229 | 7176.48
5230 | 7191.63
5231 | 7176.61
5232 | 7161.07
5233 | 7187.62
5234 | 7170.85
5235 | 7167.3
5236 | 7181.72
5237 | 7170.41
5238 | 7153.63
5239 | 7158.07
5240 | 7169.11
5241 | 7172.68
5242 | 7157.87
5243 | 7157.45
5244 | 7177.1
5245 | 7174.87
5246 | 7164.7
5247 | 7177.07
5248 | 7166.04
5249 | 7175.11
5250 | 7149.99
5251 | 7154.95
5252 | 7166.99
5253 | 7153.41
5254 | 7157.87
5255 | 7163.27
5256 | 7177.46
5257 | 7149.91
5258 | 7166.63
5259 | 7174.79
5260 | 7167.08
5261 | 7173.9
5262 | 7177.24
5263 | 7162.98
5264 | 7163.56
5265 | 7164.61
5266 | 7152.81
5267 | 7147.1
5268 | 7166.78
5269 | 7163.8
5270 | 7168.65
5271 | 7178.7
5272 | 7158.94
5273 | 7150.1
5274 | 7161.49
5275 | 7166.88
5276 | 7166.59
5277 | 7165.1
5278 | 7166.29
5279 | 7162.48
5280 | 7151.71
5281 | 7184.35
5282 | 7148.5
5283 | 7165.46
5284 | 7144.54
5285 | 7164.57
5286 | 7156.96
5287 | 7168.12
5288 | 7165.76
5289 | 7160.57
5290 | 7160.28
5291 | 7167.21
5292 | 7169.81
5293 | 7159.17
5294 | 7157.86
5295 | 7153.08
5296 | 7156.25
5297 | 7170.44
5298 | 7167.71
5299 | 7146.94
5300 | 7180.53
5301 | 7164.82
5302 | 7170.72
5303 | 7145.07
5304 | 7158.87
5305 | 7154.09
5306 | 7162.49
5307 | 7166.67
5308 | 7190.48
5309 | 7164.66
5310 | 7152.18
5311 | 7180.23
5312 | 7164.99
5313 | 7163.32
5314 | 7160.84
5315 | 7160.63
5316 | 7167.27
5317 | 7159.88
5318 | 7154.83
5319 | 7160.24
5320 | 7169.64
5321 | 7148.55
5322 | 7171.47
5323 | 7159.49
5324 | 7172.66
5325 | 7163.33
5326 | 7177.69
5327 | 7176.25
5328 | 7162.71
5329 | 7172.36
5330 | 7167.98
5331 | 7145.91
5332 | 7174.81
5333 | 7166.03
5334 | 7182.99
5335 | 7157.4
5336 | 7160.2
5337 | 7167.46
5338 | 7171.53
5339 | 7163.54
5340 | 7151.51
5341 | 7182.82
5342 | 7160.51
5343 | 7166.75
5344 | 7151.59
5345 | 7174.77
5346 | 7156.6
5347 | 7166.48
5348 | 7151
5349 | 7154.7
5350 | 7154.84
5351 | 7149.94
5352 | 7165.97
5353 | 7155.69
5354 | 7168.52
5355 | 7151.48
5356 | 7145.29
5357 | 7162.76
5358 | 7154.56
5359 | 7169.24
5360 | 7153.25
5361 | 7155.95
5362 | 7154.56
5363 | 7146.02
5364 | 7166.67
5365 | 7159.68
5366 | 7149.35
5367 | 7166.56
5368 | 7122.72
5369 | 7160.89
5370 | 7157.81
5371 | 7150.08
5372 | 7172.77
5373 | 7159.09
5374 | 7172.2
5375 | 7142.3
5376 | 7155.22
5377 | 7151.78
5378 | 7151.27
5379 | 7166.49
5380 | 7170.35
5381 | 7159.09
5382 | 7162.25
5383 | 7176.98
5384 | 7148.7
5385 | 7161.22
5386 | 7157.45
5387 | 7143.28
5388 | 7154.61
5389 | 7161.55
5390 | 7166.81
5391 | 7164.23
5392 | 7148.16
5393 | 7156.55
5394 | 7146.45
5395 | 7154.71
5396 | 7164.29
5397 | 7151.76
5398 | 7146.37
5399 | 7140.32
5400 | 7166.01
[Comments]
Sample name
Operator
Division
Comment
[Detailed Information]
Creation date | 2024/12/11 14:36
Start | 0 sec
End | 5400 sec
Data interval | 1 sec
Data points | 5401
[Measurement Information]
Model name | FP-8200
Serial No, | D030361448
Photometric mode | Em intensity
Ex bandwidth | 5 nm
Em bandwidth | 2,5 nm
Response | 0,02 sec
Sensitivity | Medium
Ex wavelength | 350 nm
Em wavelength | 450 nm
Auto gain | OFF
Light source | Xe
Filter | Do not use
[Accessory Information]
Accessory | SCE-845
Serial No, | D030361448